## Tracuu (hidden)
Thông số quả cân chuẩn (k=2) |  |  |  | Cân chuẩn (k=1) |  |  | Gamma (γ) |  |  | Đường kính trong cổ bình |  |  | Đồng hồ nhiệt độ, độ ẩm |  |  |  | SSCP Bình định mức
GTDN (g) | KLQU (g) | ĐKĐBĐ (mg) |  | Mức cân (g) | ĐKĐBĐ (mg) |  | Vật liệu | Giá trị γ (oC-1) |  | GTDĐ (mL) |  | Đường kính trong (mm) | U (THERMO.04) | 0.14 | oC |  | DTDĐ (mL) |  | SSCP (mL)
0.001 | 0.001002 | 0.002 |  | 0.001 | 0.000112 |  | Duran | 9.9e-06 |  | Cổ hẹp | 1 | 7 | U (THERMO_H.01) | 0.4 | oC |  |  |  | Cấp A | Cấp B
0.002 | 0.002 | 0.002 |  | 0.002 | 0.000106 |  | Pyrex | 9.9e-06 |  |  | 2 | 7 |  | 2 | %RH |  | Cổ hẹp | 1 | 0.025 | 0.05
0.002* | 0.002001 | 0.002 |  | 0.005 | 0.000106 |  | Rasotherm | 9.9e-06 |  |  | 5 | 7 | [5 empty] | 2 | 0.025 | 0.05
0.005 | 0.005 | 0.002 |  | 0.01 | 0.000101 |  | Durobax | 1.5e-05 |  |  | 10 | 7 | Đồng hồ áp suất tuyệt đối |  |  |  |  | 5 | 0.025 | 0.05
0.01 | 0.01 | 0.003 |  | 0.02 | 0.000106 |  | Fiolax | 1.5e-05 |  |  | 20 | 9 | U (KT12-09) | 5 | hPa |  |  | 10 | 0.025 | 0.05
0.02 | 0.020006 | 0.003 |  | 0.05 | 0.000101 |  | AR | 2.7e-05 |  |  | 25 | 9 | [5 empty] | 20 | 0.04 | 0.08
0.02* | 0.020002 | 0.003 |  | 0.1 | 0.000141 |  | AW | 2.7e-05 |  |  | 50 | 11 | [5 empty] | 25 | 0.04 | 0.08
0.05 | 0.050002 | 0.004 |  | 0.2 | 0.000112 |  | GW | 2.7e-05 |  |  | 100 | 13 | [5 empty] | 50 | 0.06 | 0.12
0.1 | 0.1 | 0.005 |  | 0.5 | 0.000123 |  | Thủy tinh thường | 2.7e-05 |  |  | 200 | 15.5 | [5 empty] | 100 | 0.1 | 0.2
0.2 | 0.2 | 0.006 |  | 1 | 0.000123 |  | Nhựa | 0.00024 |  |  | 250 | 15.5 | [5 empty] | 200 | 0.15 | 0.3
0.2* | 0.2 | 0.006 |  | 2 | 0.000107 |  | Borosilicate 3.3 | 9.9e-06 |  |  | 500 | 19 | [5 empty] | 250 | 0.15 | 0.3
0.5 | 0.50002 | 0.008 |  | 5 | 0.000106 |  | Borosilicate 5.0 | 1.5e-05 |  |  | 1000 | 23 | [5 empty] | 500 | 0.25 | 0.5
1 | 1 | 0.01 |  | 10 | 0.000129 | [5 empty] | 2000 | 27.5 | [5 empty] | 1000 | 0.4 | 0.8
2 | 2 | 0.013 |  | 20 | 0.000107 | [5 empty] | 5000 | 38 | [5 empty] | 2000 | 0.6 | 1.2
2* | 2 | 0.013 |  | 50 | 0.000107 |  |  |  |  | Cổ rộng | 5 | 9 | [5 empty] | 5000 | 1.2 | 2.4
5 | 5 | 0.017 |  | 100 | 0.000119 | [5 empty] | 10 | 9 |  |  |  |  | Cổ rộng | 5 | 0.04 | 0.08
10 | 10 | 0.02 |  | 200 | 0.041 | [5 empty] | 20 | 11 | [5 empty] | 10 | 0.04 | 0.08
20 | 20 | 0.027 |  | 500 | 0.041 | [5 empty] | 25 | 11 | [5 empty] | 20 | 0.06 | 0.12
20* | 20 | 0.027 |  | 1000 | 0.041 | [5 empty] | 50 | 13 | [5 empty] | 25 | 0.06 | 0.12
50 | 50 | 0.033 |  | 2000 | 0.051 | [5 empty] | 1000 | 27.5 | [5 empty] | 50 | 0.1 | 0.2
100 | 100.0001 | 0.053 |  | 5000 | 14.515 | 6200 | [11 empty] | 1000 | 0.6 | 1.2
200 | 200 | 0.1 |  | 25 | 0.000107
200* | 200 | 0.1 |  | 150 | 0.000119
500 | 500.0002 | 0.27 |  | 250 | 0.041
1000 | 1000.0005 | 0.53 |  | 300 | 0.041
2000 | 2000.0002 | 1.5 |  | 400 | 0.041
5000 | 5000.0003 | 2.7 |  | 600 | 0.041
25 | 25 | 0.044 |  | 800 | 0.041
150 | 150.0001 | 0.086 |  | 3000 | 15.085 | 6200
250 | 250.0001 | 0.133
300 | 300.0001 | 0.153
400 | 400.0001 | 0.2
600 | 600.0003 | 0.323
800 | 800.0003 | 0.423
3000 | 3000.0006 | 2.03
Ống đong chia độ | [6 empty] | Bình định mức | [5 empty] | Pipet một mức |  |  |  |  | Cốc đong có miệng rót (Beaker) | [5 empty] | Pipet CHIA VẠCH
GTDĐ (mL) |  | Gía trị độ chia (mL) | Dung tích tới vạch thấp nhất | SSCP (mL) |  |  | GTDĐ (mL) |  | Đường kính trong (mm) | SSCP (mL) |  |  | Dung tích danh định (mL) | Đường kính trong của pipet (mm) | Sai số cho phép |  |  | Dung tích danh định (mL) |  | Đường kính ngoài (mm) | Độ dày thành cốc (mm) | Sai số cho phép (mL) | Chiều cao (mm) | Dung tích danh định (mL) | Gía trị vạch chia (mL) | Sai số cho phép
|  |  |  | Cấp A | Cấp B |  |  |  |  | Cấp A | Cấp B |  |  |  | A và AS (mL) | B (mL) | [9 empty] | A và AS (mL) | B (mL)
1a, 1b | 5 | 0.1 | 1 | 0.05 | 0.1 |  | Cổ hẹp | 1 | 7 | 0.025 | 0.05 |  | 0.5 | 2.5 | 0.005 | 0.01 |  | Dạng thấp | 5 | 22 | 0.7 | 1 | 32 | 0.1 | 0.01 | 0.006 | 0.01
| 10 | 0.2 | 1.4 | 0.1 | 0.2 |  |  | 2 | 7 | 0.025 | 0.05 |  | 1 | 3 | 0.008 | 0.015 |  |  | 10 | 26 | 0.7 | 1 | 37 | 0.2 | 0.01 | 0.006 | 0.01
| 25 | 0.5 | 2.5 | 0.25 | 0.5 |  |  | 5 | 7 | 0.025 | 0.05 |  | 2 | 3.5 | 0.01 | 0.02 |  |  | 25 | 34 | 0.7 | 1 | 52 | 0.5 | 0.01 | 0.006 | 0.01
| 50 | 1 | 5 | 0.5 | 1 |  |  | 10 | 7 | 0.025 | 0.05 |  | 5 | 4 | 0.015 | 0.03 |  |  | 50 | 42 | 0.8 | 2.5 | 62 | 1 | 0.01 | 0.007 | 0.01
| 100 | 1 | 10 | 0.5 | 1 |  |  | 20 | 9 | 0.04 | 0.08 |  | 10 | 4.5 | 0.02 | 0.04 |  |  | 100 | 50 | 0.9 | 5 | 72 | 1 | 0.1 | 0.007 | 0.01
| 250 | 2 | 26 | 1 | 2 |  |  | 25 | 9 | 0.04 | 0.08 |  | 20 | 5.5 | 0.03 | 0.06 |  |  | 150 | 60 | 1 | 5 | 82 | 2 | 0.02 | 0.01 | 0.02
| 500 | 5 | 50 | 2.5 | 5 |  |  | 50 | 11 | 0.06 | 0.12 |  | 25 | 5.5 | 0.03 | 0.06 |  |  | 250 | 70 | 1.1 | 10 | 97 | 2 | 0.1 | 0.01 | 0.02
| 1000 | 10 | 100 | 5 | 10 |  |  | 100 | 13 | 0.1 | 0.2 |  | 50 | 6 | 0.05 | 0.1 |  |  | 400 | 80 | 1.2 | 10 | 113 | 5 | 0.05 | 0.03 | 0.05
| 2000 | 20 | 200 | 10 | 20 |  |  | 200 | 15.5 | 0.15 | 0.3 |  | 100 | 7.5 | 0.08 | 0.15 |  |  | 500 | 85 | 1.3 | 15 | 118 | 5 | 0.1 | 0.03 | 0.05
2 | 5 | 0.5 | 1 |  | 0.2 |  |  | 250 | 15.5 | 0.15 | 0.3 | [7 empty] | 600 | 90 | 1.3 | 20 | 128 | 10 | 0.1 | 0.05 | 0.1
| 10 | 1 | 2 |  | 0.3 |  |  | 500 | 19 | 0.25 | 0.5 | [7 empty] | 800 | 100 | 1.3 | 20 | 138 | 20 | 0.1 | 0.1 | 0.2
| 25 | 1 | 5 |  | 0.5 |  |  | 1000 | 23 | 0.4 | 0.8 | [7 empty] | 1000 | 105 | 1.3 | 20 | 148 | 25 | 0.1 | 0.1 | 0.2
| 50 | 2 | 10 |  | 1 |  |  | 2000 | 27.5 | 0.6 | 1.2 | [7 empty] | 2000 | 130 | 1.4 | 50 | 188 | 25 | 0.2 | 0.1 | 0.2
| 100 | 2 | 12 |  | 1 |  |  | 5000 | 38 | 1.2 | 2.4 | [7 empty] | 3000 | 150 | 1.7 | 50 | 214
| 250 | 5 | 30 |  | 2 |  | Cổ rộng | 5 | 9 | 0.04 | 0.08 | [7 empty] | 5000 | 170 | 2 | 100 | 274
| 500 | 10 | 60 |  | 5 |  |  | 10 | 9 | 0.04 | 0.08 | [6 empty] | Dạng cao | 50 | 38 | 0.8 | 2.5 | 72
| 1000 | 20 | 100 |  | 10 |  |  | 20 | 11 | 0.06 | 0.12 | [7 empty] | 100 | 48 | 0.9 | 5 | 82
| 2000 | 50 | 200 |  | 20 |  |  | 25 | 11 | 0.06 | 0.12 | [7 empty] | 150 | 54 | 1 | 5 | 97
[8 empty] | 50 | 13 | 0.1 | 0.2 | [7 empty] | 250 | 60 | 1.1 | 10 | 123
[8 empty] | 1000 | 27.5 | 0.6 | 1.2 | [7 empty] | 400 | 70 | 1.2 | 10 | 133
[19 empty] | 500 | 79 | 1.3 | 15 | 140
[19 empty] | 600 | 80 | 1.3 | 20 | 153
[19 empty] | 800 | 90 | 1.3 | 20 | 178
[19 empty] | 1000 | 95 | 1.3 | 20 | 183
[19 empty] | 2000 | 120 | 1.4 | 50 | 244
[19 empty] | 3000 | 135 | 1.7 | 50 | 284
[19 empty] | 5000 | 160 | 2 | 100 | 324
BẢNG THÔNG SỐ DUNG TÍCH HIỆU CHUẨN
Mức dung tích danh định (mL) | Mức 2 (mL) | Mức 1 (mL)
0.5 | 0.2 | 0.1
1 | 0.5 | 0.2
2 | 1 | 0.5
5 | 2 | 1
10 | 5 | 1
20 | 10 | 5
25 | 10 | 5
50 | 20 | 10
100 | 50 | 20
150 | 100 | 50
200 | 100 | 50
250 | 100 | 50
400 | 200 | 100
500 | 200 | 100
600 | 200 | 100
800 | 400 | 200
1000 | 500 | 200
2000 | 1000 | 500
3000 | 2000 | 500
5000 | 2000 | 500

## TBC
STT | KEY | CODE | NAME | NSX | MODEL | SERIAL | GCN | LIÊN KẾT CHUẨN | NGÀY HIỆU CHUẨN | Next cal date
1 | baro | LAB-BARO.01 | Thermo-hygro-barometer | LUTRON | MHB-382SD | AM.29323 | AC-1868 | TECHMASTER | 29/08/2025 | 2026-08-29
2 | volume | LAB-WATERFLOW.01 | Water Flow Meter | N/A | N/A |  | KT3-00042BDT5 | QUATEST 3 | 27/08/2025 | 2026-08-27
3 | timer | LAB-TIMER.01 | Stopwatch | CASIO | HS-70W | J003Q02 | KT3-01161BDE5/1 | QUATEST 3 | 2025-08-09 | 2026-08-09 | 0.47 S/D
4 | tacho2 | LAB-TACHO.03 | Digital Tachometer | FLUKE | 931 | N/A | KT3-01161BDE5/2 | QUATEST 3 | 2025-08-09 | 2026-08-09
5 | tacho3 | LAB-TACHO.02 | Digital Tachometer | TENMARS | TM-4100N | 230500607 | KT3-01161BDE5/3 | QUATEST 3 | 2025-08-09 | 2026-08-09
6 | timer2 | LAB-TIMER.02 | STOPWATCH | Extech | 365510 | N/A | TSG-0-554097 | QUATEST 3 | 17/9/2025 | 2026-09-17
7 | 10kg | LAB-10M1WEIGHT.01 | Weight M1-10 | CHINA |  | -- | KT3-0122BKL5 | QUATEST 3 | 2025-03-07 | 2026-03-07
8 | 20kg1 | LAB-20M1WEIGHT.02 | Weight M1-20 | CHINA |  | -- | KT3-0122BKL5 | QUATEST 3 | 2025-03-07 | 2026-03-07
9 | 20kg |  | Weight M1-20 | CHINA |  | -- | KT3-0122BKL5 | QUATEST 3 | 2025-03-07 | 2026-03-07
10 | 20kg2 | LAB-20M1WEIGHT.03 | Weight M1-20 | CHINA |  | -- | KT3-0122BKL5 | QUATEST 3 | 2025-03-07 | 2026-03-07
11 | 20kg3 | LAB-20M1WEIGHT.04 | Weight M1-20 | CHINA |  | -- | KT3-0122BKL5 | QUATEST 3 | 2025-03-07 | 2026-03-07
12 | 20kg4 | LAB-20M1WEIGHT.05 | Weight M1-20 | CHINA |  | -- | KT3-0122BKL5 | QUATEST 3 | 2025-03-07 | 2026-03-07
13 | caliper2 | LAB-CALIPER.02 | Digital caliper | Mitutoyo | CD-20APX | A17096618 | KT3-01293BDD5/1 | QUATEST 3 | 2025-08-09 | 2026-08-09
14 | panme | LAB-PANME.01 | Venier micrometer | MITUTOYO | 293-240-30 | 74363125 | KT3-01293BDD5/2 | QUATEST 3 |  | 1900-12-31
15 | feeler | LAB-FEELER.01 | Feeler gauge | NIGATA SEIKI | NO.150ML | 01293BDD5/3 | KT3-01293BDD5/3 | QUATEST 3 | 2025-08-09 | 2026-08-09
16 | needle | LAB-CAL_TOOL_16 | Needle detector card | HASHIMA | 0.8; 1.0; 1.2 |  | KT3-01293BDD5/4 KT3-01293BDD5/5 KT3-01293BDD5/5 | QUATEST 3 | 28/08/2025 | 2026-08-28
17 | thickness2 | LAB-GBLOCK.01 | Gauge block grade 0 | CHINA |  | N/A | KT3-01293BDD5/7 | QUATEST 3 | 2025-03-09 | 2026-03-09
18 | hum | LAB-THER_HUM.02 | Thermo-hygrometer | TESTO | 174H | 37341315 | KT3-01499BHL5/1 | QUATEST 3 | 2025-04-09 | 2026-04-09 | shc -5%rh
19 | flowmeter | LAB-AIRFLOW.02 | Air Flow Meter | TENMARS | TM-4001 | 230500116 | KT3-01499BHL5/2 | QUATEST 3 | 2025-04-09 | 2026-04-09
20 | load210g | LAB-BALANCE.03 | Digital Balance | OHAUS | PA214 | B523014198 | KT3-01514BKL5/1 | QUATEST 3 | 22/08/2025 | 2026-08-22
21 | load620g | LAB-BALANCE.02 | Digital Balance | VIBRA SHINKO | AB623 | 240318362 | KT3-01514BKL5/2 | QUATEST 3 | 22/08/2025 | 2026-08-22
22 | balance2 | LAB-WEIGHTSET.02 | F1 Standard Weight Set | METTLER TOLEDO | F1 | 30402730 | KT3-01514BKL5/3 | QUATEST 3 | 29/08/2025 | 2026-08-29
23 | pressure2 | LAB-PRESSURE.02 | Pressure Gauge | LUTRON | PS-9302 | I.407815 Q620822 | KT3-01858BCO5/1 | QUATEST 3 | 2025-04-09 | 2026-04-09
24 | force3 | LAB-MINIFORCE.02 | Pressure Gauge | CHINA | L400C | LZ0209312K Loadcell MODEL: LFT-25-20KG SN: LZ0209312K | KT3-01858BCO5/2 | QUATEST 3 | 25/08/2025 | 2026-08-25
25 | infrared | LAB-CAL_TOOL_26 | Infrared thermomter | FLUKE 572-2 | 572-2 | 36230624WS | KT3-02187BNH5/1 | QUATEST 3 | 2025-05-09 | 2026-05-09
26 | Temp1 | LAB-THERMO.03 | Digital Thermometer | Extech |  | H330198 | KT3-02187BNH5/2 | QUATEST 3 | 2025-10-09 | 2026-10-09
27 | pressure | LAB-PRESSURE.01 | Pressure Gauge | BAROLI | RAISOVA 547 | 2257189 | KT3-02205BCO4/2 | QUATEST 3 | 2025-05-12 | 2026-05-12
28 | force2 | LAB-FORCE.01 | Loadcell and indicator | VMC/ Laumas |  | 15081469 201184903 | KT3-02386BCO4 | QUATEST 3 |  | 1900-12-31
29 | Temp2 | LAB-THERMO.04 | Digital Thermometer | Extech |  | H399380 | KT3-02443BNH4/2 | QUATEST 3 | 2025-05-12 | 2026-05-12
30 | thickness | LAB-THICKNESS.01 | Digital Indicator | INSIZE | 2103-05 | M1807248075 | TSG-0-554096 | TECHMASTER | 17/09/2025 | 2026-09-17
31 | 1kg | LAB-1M1WEIGHT.01 | Weight M1-1 | CHINA |  | 1875 |  |  |  | 1900-12-31
32 | 2kg | LAB-2M1WEIGHT.01 | Weight M1-2 | CHINA |  | 1873 |  |  |  | 1900-12-31
33 | 2kg1 | LAB-2M1WEIGHT.01 | Weight M1-2 | CHINA |  | 1874 |  |  |  | 1900-12-31
34 | chamber | LAB-9HSENSORS.01 | Temp-Humidity 9CH | China |  | -- |  |  |  | 1900-12-31
35 | angle | LAB-ANGLE.01 | Angle Meter | LEVEL BOX |  | 1659BDD9/1 |  |  |  | 1900-12-31
36 | load10kg | LAB-BALANCE.01 | Digital Balance | LABEX | HC-B100001 | 170630023 |  |  |  | 1900-12-31
37 | load20kg | LAB-BALANCE.04 | Digital Balance | TPS |  | CTPS57157100124 |  |  |  | 1900-12-31
38 | load3000g | LAB-BALANCE.05 | Digital Balance | [6 empty] | 1900-12-31
39 | load5000g | LAB-BALANCE.06 | Digital Balance | [6 empty] | 1900-12-31
40 | caliper1 | LAB-CALIPER.01 | Digital caliper | Mitutoyo | CD-15APX | B21022878 |  |  |  | 1900-12-31
41 | hardness | LAB-DUROKIT.01 | durometer kit | [6 empty] | 1900-12-31
42 | force1 | LAB-FORCE.02 | Loadcell and indicator | VMC/ Laumas |  | 33073261 |  |  |  | 1900-12-31
43 | force20kg | LAB-FORCE.03 | Loadcell and indicator | CHINA | L400C | LZ0209312K Loadcell MODEL: LFT-25-20KG SN: LZ0209312K |  |  |  | 1900-12-31
44 | thickness | LAB-GBLOCK.02 | Gauge block grade 0 | CHINA | GRADE 0 | 200303 |  |  |  | 1900-12-31
45 | heaterblock | LAB-HEATER-01 | Temperature Heater | CHINA |  | 2003323 |  |  |  | 1900-12-31
46 | infrared2 | LAB-INFRARED.01 | Infrared thermomter | FLUKE | 62 MAX | 29940181WS |  |  |  | 1900-12-31
47 | length | LAB-LENGTH.01 | Tensile Length Counter | MITUTOYO | [5 empty] | 1900-12-31
48 | light | LAB-LIGHT.01 | Light Meter | Sekonic | C-700R | JT22-000383 |  |  |  | 1900-12-31
49 | light | LAB-LIGHT.02 | Light Meter | KONICA MILNOTA | CL-200A | 80411030 |  | VMI | 2025-04-09 | 2026-04-09
50 | force4 | LAB-FORCE.04 | Compressor load - indicator | CAS | 50kgf |  |  |  |  | 1900-12-31
51 | force5 | LAB-FORCE.05 | Compressor load - indicator | BONGSHIN | BS-32 | Y30167 BS-32 X23165 |  |  |  | 1900-12-31
52 | force5 | LAB-MINIFORCE.05 | Miniloadcell - indicator 200kg | BONGSHIN | BS-32 | S30056 BS-32 R45167 | TSG-0-554093 | TECHMASTER | 17/09/2025 | 2026-09-17
53 | needle2 | LAB-NEEDLE.02 | Needle detector card | HASHIMA | 0.8; 1.0; 1.2 |  |  |  |  | 1900-12-31
54 | ph | LAB-PH-01 | pH Solution (4; 7; 10) | USA |  | Manufactuter's |  |  |  | 1900-12-31
55 | pingauge | LAB-PINSET.01 | Pin gauge set | CHINA | 4.98 - 5.02 | 5 PCS | TSG-0-554094 | TECHMASTER | 17/09/2025 | 2026-09-17
56 | ruler | LAB-RULER.01 | Steel Ruler | SHINWA |  | 1659BDD9/2 | TSG-0-554095 | TECHMASTER | 17/09/2025 | 2026-09-17
57 | solar | LAB-SOLAR.01 | Solar Power Meter | LUTRON | SPM-1116SD | Q968431 |  |  |  | 1900-12-31
58 | spectro | LAB-SPECTRO.01 | Spectro Cal Kit | JAPAN | JCRM110 JCRM140 | 81128 14240 |  |  |  | 1900-12-31
59 | tacho1 | LAB-TACHO.01 | Digital Tachometer | THB |  | S1049626 |  |  |  | 1900-12-31
60 | hum2 | LAB-THER_H.02 | Thermo-hygrometer | Deree |  | 251865 |  |  |  | 1900-12-31
61 | hum3 | LAB-THER_H.03 | Thermo-hygrometer | Deree |  | 251868 |  |  |  | 1900-12-31
62 | Temp5 | LAB-THERMO.01 | Digital Thermometer | ETI | REFERENCE MK2 | D25216457 |  |  |  | 1900-12-31
63 | Temp4 | LAB-THERMO.02 | Digital Thermometer | ETI |  | D24270697 |  |  |  | 1900-12-31
64 | Temp3 | LAB-THERMO.05 | Digital Thermometer | Extech |  | 1904306 |  |  |  | 1900-12-31
65 | temp11ch | LAB-THERMO.06 | Temperature 11CH | China |  | TK6071IP |  |  |  | 1900-12-31
66 | temp10ch | LAB-THERMO.07 | Temperature logger | GRAPHTEC |  | GL220 |  |  |  | 1900-12-31
67 | timer3 | LAB-TIMER.03 | Stopwatch | Extech | 365510 | N/A |  |  |  | 1900-12-31
68 | balance1 | LAB-WEIGHTSET.01 | F1 Standard Weight Set | METTLER TOLEDO |  | 18072601 |  |  |  | 1900-12-31
69 | balance3 | LAB-WEIGHTSET.03 | F1 Standard Weight Set | [6 empty] | 1900-12-31
70 | wilesstemp | LAB-WIFI_TEPM.01 | Wiless Temperature | MADGETECH | HiTemp140 | 45335 |  |  |  | 1900-12-31
71 | wilesstemp2 | LAB-WIFI_TEPM.02 | Hight temperature data logger | YOWEXA | YPL-190-B | 24230HTC0000004 |  |  |  | 1900-12-31
72 | linear | LAB-LINEAR.01 | LINEAR SCALE |  |  |  | KT3-01596BDD5 | QUATEST 3 | 2025-10-28 | 2026-10-28
73 | [9 empty] | 1900-12-31
74 | [9 empty] | 1900-12-31
75 | [9 empty] | 1900-12-31
76 | [9 empty] | 1900-12-31
77 | [9 empty] | 1900-12-31
78 | [9 empty] | 1900-12-31
79 | [9 empty] | 1900-12-31
80 | [9 empty] | 1900-12-31
81 | [9 empty] | 1900-12-31
82 | [9 empty] | 1900-12-31
83 | [9 empty] | 1900-12-31
84 | [9 empty] | 1900-12-31
85 | [9 empty] | 1900-12-31
86 | [9 empty] | 1900-12-31
87 | [9 empty] | 1900-12-31
88 | [9 empty] | 1900-12-31
89 | [9 empty] | 1900-12-31
90 | [9 empty] | 1900-12-31
91 | [9 empty] | 1900-12-31
92 | [9 empty] | 1900-12-31
93 | [9 empty] | 1900-12-31
94 | [9 empty] | 1900-12-31
95 | [9 empty] | 1900-12-31
96 | [9 empty] | 1900-12-31
97 | [9 empty] | 1900-12-31
98 | [9 empty] | 1900-12-31
99 | [9 empty] | 1900-12-31
100 | [9 empty] | 1900-12-31
101 | [9 empty] | 1900-12-31
102 | [9 empty] | 1900-12-31
103 | [9 empty] | 1900-12-31
104 | [9 empty] | 1900-12-31
105 | [9 empty] | 1900-12-31
106 | [9 empty] | 1900-12-31
107 | [9 empty] | 1900-12-31
108 | [9 empty] | 1900-12-31
109 | [9 empty] | 1900-12-31
110 | [9 empty] | 1900-12-31
111 | [9 empty] | 1900-12-31
112 | [9 empty] | 1900-12-31
113 | [9 empty] | 1900-12-31
114 | [9 empty] | 1900-12-31
115 | [9 empty] | 1900-12-31
116 | [9 empty] | 1900-12-31
117 | [9 empty] | 1900-12-31
118 | [9 empty] | 1900-12-31
119 | [9 empty] | 1900-12-31
120 | [9 empty] | 1900-12-31
121 | [9 empty] | 1900-12-31
122 | [9 empty] | 1900-12-31
123 | [9 empty] | 1900-12-31
124 | [9 empty] | 1900-12-31
125 | [9 empty] | 1900-12-31
126 | [9 empty] | 1900-12-31
127 | [9 empty] | 1900-12-31
128 | [9 empty] | 1900-12-31
129 | [9 empty] | 1900-12-31
130 | [9 empty] | 1900-12-31
131 | [9 empty] | 1900-12-31
132 | [9 empty] | 1900-12-31
133 | [9 empty] | 1900-12-31
134 | [9 empty] | 1900-12-31
135 | [9 empty] | 1900-12-31
136 | [9 empty] | 1900-12-31
137 | [9 empty] | 1900-12-31
138 | [9 empty] | 1900-12-31
139 | [9 empty] | 1900-12-31
140 | [9 empty] | 1900-12-31
141 | [9 empty] | 1900-12-31
142 | [9 empty] | 1900-12-31
143
144
145
146
147
148
149
Formulas:
Next cal date = EDATE([NGÀY HIỆU CHUẨN],12)

## master (hidden)
1 | 2 |  |  | 5 | 6 | 7 | 8 | 9 | 10 | 12 | 11 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21 | 22 |  | TBHC
Calib Area Lĩnh vực HC | Test Method/ Standard | Application | Quy trình HC ISO ISO Calib Procedure | Instrument Thiết bị | Description - EN | Description | PP đo | Measured Parameters Thông số đo | Calibration points Điểm hiệu chuẩn | Tolerance Dung sai cho phép | Parameters - EN | Thông số đo - VN | Unit Đơn vị | Target Giá trị mong muốn | Dung sai (giá trị) | Lower | Upper | Đại lượng đo | Độ phân giải của thiết bị được hiệu chuẩn: | Độ KĐBĐ (Độ hiệu chính đối Oven) | PRICE | Scope of Service | Độ phân giải của thiết bị được hiệu chuẩn: | Độ KĐBĐ | KEY-DL | Thể hiện số thập phân từ kết quả đo | Unit để lấy Độ KĐBĐ
KEY | COC | REF | BB, COC | REF | Quote, BB, COC | COC | Quote | BB, COC | BB, COC | BB, COC | COC | COC |  | COC | [6 empty] | Quote | Quote
Dry-6330 | ISO 6330 | Tumble Dryer | LAB-T-01 LAB-TIME-02 | Accudry, Precision Dryer | Dryer | Máy sấy | Exhaust Temperature Time Cool down Period | Temperature (*) - ℃ | 40 | ±3 | Temperature (*) - ℃ | Nhiệt độ | ℃ | 40 | 3 | 37 | 43 | DEG |  | #N/A | 1500000 | Service & Calibration | #N/A | #N/A | Dry-6330-DEG |  | Std
Dry-6330 |  |  |  | 10: 50; 12: ±3; 14: ℃; 15: 50; 16: 3; 17: 47; 18: 53; 19: DEG; 21: #N/A |  | #N/A | #N/A | Dry-6330-DEG |  | Std | 50
Dry-6330 |  |  |  | 10: 55; 12: ±3; 14: ℃; 15: 55; 16: 3; 17: 52; 18: 58; 19: DEG; 21: #N/A |  | #N/A | #N/A | Dry-6330-DEG |  | Std | 55
Dry-6330 |  |  |  | 10: 60; 12: ±3; 14: ℃; 15: 60; 16: 3; 17: 57; 18: 63; 19: DEG; 21: #N/A |  | #N/A | #N/A | Dry-6330-DEG |  | Std
Dry-6330 |  |  |  | 10: 70; 12: ±3; 14: ℃; 15: 70; 16: 3; 17: 67; 18: 73; 19: DEG; 21: #N/A |  | #N/A | #N/A | Dry-6330-DEG |  | Std
Dry-6330 | [7 empty] | Time - sec | 1800 | N/A | Time - sec | Thời gian | sec | 1800 |  | 1800 | 1800 | TIME |  | #N/A |  |  | #N/A | #N/A | Dry-6330-TIME |  | Y | 1800 - Time - sec
Dry-6330 |  |  |  | 10: 2700; 12: N/A; 14: sec; 15: 2700; 17: 2700; 18: 2700; 19: TIME; 21: #N/A |  | #N/A | #N/A | Dry-6330-TIME |  | Y
Dry-6330 |  |  |  | 10: 3000; 12: N/A; 14: sec; 15: 3000; 17: 3000; 18: 3000; 19: TIME; 21: #N/A |  | #N/A | #N/A | Dry-6330-TIME |  | Y | 3000
Dry-6330 | [7 empty] | Cool down Period - sec | 300 | N/A | Cool down Period - sec | Thời gian làm mát | sec | 300 | N/A | #VALUE! | #VALUE! | TIME |  | #N/A |  |  | #N/A | #N/A | Dry-6330-TIME |  | Y
Dry-6330 |  |  |  | 10: 600; 12: N/A; 14: sec; 15: 600; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: #N/A |  | #N/A | #N/A | Dry-6330-TIME |  | Y
Dry-6330 |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Dry-6330- |  | Y
Dry-6330 | bar: 9; 11
Dry-6330 | bar: 9; 11
Dry-6330 | bar: 9; 11
Dry-6330 | bar: 9; 11
Dry-Puma | Puma | Tumble Dryer | LAB-T-01 LAB-TIME-02 | Accudry, Precision Dryer | Dryer | Máy sấy | Exhaust Temperature Time Cool down Period | Temperature (*) - ℃ | 40 | ±3 | Temperature (*) - ℃ | Nhiệt độ | ℃ | 40 | 3 | 37 | 43 | DEG |  | 0.2 | 1500000 | Service & Calibration | 0.1 | 0.2 | Dry-Puma-DEG |  | Std
Dry-Puma |  |  |  | 10: 50; 12: ±3; 14: ℃; 15: 50; 16: 3; 17: 47; 18: 53; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Dry-Puma-DEG |  | Std | 50
Dry-Puma |  |  |  | 10: 55; 12: ±3; 14: ℃; 15: 55; 16: 3; 17: 52; 18: 58; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Dry-Puma-DEG |  | Std | 55
Dry-Puma |  |  |  | 10: 60; 12: ±3; 14: ℃; 15: 60; 16: 3; 17: 57; 18: 63; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Dry-Puma-DEG |  | Std
Dry-Puma |  |  |  | 10: 70; 12: ±3; 14: ℃; 15: 70; 16: 3; 17: 67; 18: 73; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Dry-Puma-DEG |  | Std
Dry-Puma |  |  |  | 10: 80; 12: ±3; 14: ℃; 15: 80; 16: 3; 17: 77; 18: 83; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Dry-Puma-DEG
Dry-Puma | [7 empty] | Time - sec | 1800 | N/A | Time - sec | Thời gian | sec | 1800 |  | 1800 | 1800 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Dry-Puma-TIME |  | Y | 1800 - Time - sec
Dry-Puma |  |  |  | 10: 2700; 12: N/A; 14: sec; 15: 2700; 17: 2700; 18: 2700; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Dry-Puma-TIME |  | Y
Dry-Puma |  |  |  | 10: 3000; 12: N/A; 14: sec; 15: 3000; 17: 3000; 18: 3000; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Dry-Puma-TIME |  | Y | 3000
Dry-Puma | [7 empty] | Cool down Period - sec | 300 | N/A | Cool down Period - sec | Thời gian làm mát | sec | 300 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Dry-Puma-TIME |  | Y
Dry-Puma |  |  |  | 10: 600; 12: N/A; 14: sec; 15: 600; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Dry-Puma-TIME |  | Y
Dry-Puma |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Dry-Puma- |  | Y
Dry-Puma | bar: 9; 11
Dry-Puma | bar: 9; 11
Dry-Puma | bar: 9; 11
Dry-Puma | bar: 9; 11
Dry-other | On request |  | LAB-T-01 LAB-TIME-02 | 5: Other dryer; 6: Dryer; 7: Máy sấy; 8: Exhaust Temperature Time Cool down Period; 9: Theo chương trình máy; 21: #N/A; 22: 1500000 | Service & Calibration | #N/A | #N/A | Dry-other- |  |  | 0
Dry-other |  |  |  | 9: Time - sec; 12: N/A; 11: Time - sec; 13: Thời gian; 14: sec; 17: 0; 18: 0; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Dry-other-TIME
Dry-other |  |  |  | 9: Cool down - sec; 12: N/A; 11: Cool down - sec; 13: Thời gian làm mát; 14: sec; 17: 0; 18: 0; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Dry-other-TIME |  |  | Cool down - sec
Dry-other |  |  |  | 9: Temp (*) - min - °C; 12: ±; 11: Temp (*) - min - °C; 13: Nhiệt độ tối thiểu; 14: ℃; 17: 0; 18: 0; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Dry-other-DEG |  |  | Temp (*) - min - °C
Dry-other |  |  |  | 9: Temp (*) - max - °C; 12: ±; 11: Temp (*) - max - °C; 13: Nhiệt độ tối đa; 14: ℃; 19: DEG | [6 empty] | Temp (*) - max - °C
Dry-other | [27 empty] | 0
Dry-other |  |  |  | 9: Others; 11: *
Dry-other | [27 empty] | * | bar: 9; 11
Dry-other | [27 empty] | * | bar: 9; 11
Dry-other | bar: 9; 11
Dry-other | bar: 9; 11
Dry-other | bar: 9; 11
Dry-nike | Puma, Nike |  | LAB-T-01 LAB-TIME-02 | Accudry | Dryer | Máy sấy |  | Time - sec | 3000 | ±600 | Time - sec | Thời gian | sec | 3000 | 600 | 2400 | 3600 | TIME |  |  |  |  | 0.001 | 0.02 | Dry-nike-TIME
Dry-nike | [7 empty] | Cool down - sec | 600 | N/A | Cool down - sec | Thời gian làm mát | sec | 600 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Dry-nike-TIME
Dry-nike | [7 empty] | Temp (*) - max 73 - °C | 50 | N/A | Temp (*) - max 73 - °C | Nhiệt độ | ℃ | 50 | N/A | #VALUE! | #VALUE! | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Dry-nike-DEG |  | Y
Dry-nike |  |  |  | 10: 60; 12: N/A; 14: ℃; 15: 60; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG |  | 0.1 | 0.2 | Dry-nike-DEG
Dry-nike |  |  |  | 9: Others; 11: * |  | #N/A | #N/A | Dry-nike- |  | Y
Dry-nike |  |  |  | 9: 1…; 11: *; 21: #N/A |  | #N/A | #N/A | Dry-nike- |  | Y
Dry-nike | [22 empty] | #N/A | #N/A | Dry-nike- |  | Y | bar: 9; 11
Dry-nike |  |  |  | 9: 3…; 11: *; 21: #N/A |  | #N/A | #N/A | Dry-nike-
Dry-nike |  |  |  | 9: 4…; 11: *; 21: #N/A |  | #N/A | #N/A | Dry-nike- |  |  | *
Dry-nike |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Dry-nike- |  |  | *
Dry-hm | H&M TM LA:01 |  | LAB-T-01 LAB-TIME-02 | 6: Dryer; 7: Máy sấy; 8: Exhaust Temperature Time Cool down Period; 9: Low; 11: Low; 21: #N/A; 22: 1500000 | Service & Verification | #N/A | #N/A | Dry-hm- |  | Y | Low
Dry-hm | [7 empty] | Temp (*) - max - °C | 50 | N/A | Temp (*) - max - °C | Nhiệt độ tối đa | ℃ | 50 |  | 0 | 50 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Dry-hm-DEG |  | Std | 50 - Temp (*) - max - °C
Dry-hm |  |  |  | 9: Temp (*) - min - °C; 12: N/A; 11: Temp (*) - min - °C; 13: Nhiệt độ tối thiểu; 14: ℃; 15: N/A; 16: N/A; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Dry-hm-DEG | 1 | Std | Temp (*) - min - °C
Dry-hm | [27 empty] | 0
Dry-hm |  |  |  | 9: Medium; 11: Medium | [6 empty] | Medium
Dry-hm | [7 empty] | Temp (*) - max - °C | 60 | N/A | Temp (*) - max - °C | Nhiệt độ tối đa | ℃ | 60 |  | 0 | 60 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Dry-hm-DEG | 1 | Std
Dry-hm |  |  |  | 9: Temp (*) - min - °C; 12: N/A; 11: Temp (*) - min - °C; 13: Nhiệt độ tối thiểu; 14: ℃; 15: N/A; 16: N/A; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Dry-hm-DEG | 1 | Std
Dry-hm
Dry-hm |  |  |  | 9: High; 11: High | [6 empty] | High
Dry-hm | [7 empty] | Temp (*) - High - max - °C | 65 | N/A | Temp (*) - High - max - °C | Nhiệt độ tối đa | ℃ | 65 |  | 0 | 65 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Dry-hm-DEG | 1 | Std | 65 - Temp (*) - High - max - °C
Dry-hm |  |  |  | 9: Temp (*) - High - min - °C; 12: N/A; 11: Temp (*) - High - min - °C; 13: Nhiệt độ tối thiểu; 14: ℃; 15: N/A; 16: N/A; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Dry-hm-DEG | 1 | Std
Dry-hm
Dry-hm |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Dry-hm- |  | Y
Dry-hm |  |  |  | 9: 1…; 11: *; 21: #N/A |  | #N/A | #N/A | Dry-hm- |  | Y
Dry-hm | bar: 9; 11
Dry-hm | bar: 9; 11
Dry-hm | bar: 9; 11
Dry-hm | bar: 9; 11
Dry-decathlon | DS 0034 V3 |  | LAB-DRY-01 | Accudry | Dryer | Máy sấy | Exhaust Temperature Time Cool down Period | Temp (*) - °C | 70 | ±10 | Temp (*) - °C | Nhiệt độ | ℃ | 70 | 10 | 60 | 80 | DEG |  | 0.2 |  | Service & Calibration | 0.1 | 0.2 | Dry-decathlon-DEG |  | Y
Dry-decathlon
Dry-decathlon | [7 empty] | Time spent at temperature - sec | 5400 | ±600 | Time - sec | Thời gian kéo dài ở mốc nhiệt độ | sec | 5400 | 600 | 60 | 80 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Dry-decathlon-TIME
Dry-decathlon
Dry-decathlon |  |  |  | 9: Total time - sec; 12: N/A; 11: Total time - sec; 13: Tổng thời gian; 14: sec; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Dry-decathlon-TIME
Dry-decathlon |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Dry-decathlon- |  |  | *
Dry-decathlon | [22 empty] | #N/A | #N/A | Dry-decathlon- |  | Y | bar: 9; 11
Dry-decathlon |  |  |  | 9: 2…; 11: *; 21: #N/A |  | #N/A | #N/A | Dry-decathlon- |  | Y | *
Dry-decathlon | [22 empty] | #N/A | #N/A | Dry-decathlon- | bar: 9; 11
Dry-decathlon | [22 empty] | #N/A | #N/A | Dry-decathlon- | bar: 9; 11
Dry-decathlon | [22 empty] | #N/A | #N/A | Dry-decathlon- | bar: 9; 11
Dry-adidas | Adidas PHX-AP0701 |  | LAB-T-01 LAB-TIME-02 | Miele Dryer, Accudry | Dryer | Máy sấy |  | Temp (*) - Low - °C | 55 | ±5 | Temp (*) - Low - °C | Nhiệt độ thấp | ℃ | 55 | 5 | 50 | 60 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Dry-adidas-DEG |  | Y
Dry-adidas | [7 empty] | Temp (*) - Medium - °C | 70 | ±5 | Temp (*) - Medium - °C | Nhiệt độ trung bình | ℃ | 70 | 5 | 65 | 75 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Dry-adidas-DEG |  |  | 70 - Temp (*) - Medium - °C
Dry-adidas |  |  |  | 9: Time - sec; 10: 2700; 12: N/A; 11: Time - sec; 13: Thời gian; 14: sec; 15: 2700; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Dry-adidas-TIME |  |  | 2700 - Time - sec
Dry-adidas |  |  |  | 9: Cool down - sec; 10: 600; 12: N/A; 11: Cool down - sec; 13: Thời gian làm mát; 14: sec; 15: 600; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Dry-adidas-TIME |  | Y | 600 - Cool down - sec
Dry-adidas | [27 empty] | 0
Dry-adidas |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Dry-adidas- |  | Y | *
Dry-adidas |  |  |  | 9: 1…; 11: *; 21: #N/A |  | #N/A | #N/A | Dry-adidas- |  |  | *
Dry-adidas |  |  |  | 9: 2…; 11: *; 21: #N/A |  | #N/A | #N/A | Dry-adidas- |  |  | *
Dry-adidas |  |  |  | 9: 3…; 11: *; 21: #N/A |  | #N/A | #N/A | Dry-adidas-
Dry-adidas |  |  |  | 9: 4…; 11: *; 21: #N/A |  | #N/A | #N/A | Dry-adidas-
Dry-adidas |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Dry-adidas-
Dry-aatcc | AATCC LP1 |  | LAB-T-01 LAB-TIME-02 |  | Dryer | Máy sấy | Exhaust Temperature Time Cool down Period | Time - sec | 2700 | N/A | Time - sec | Thời gian | sec | 2700 |  |  |  | TIME |  | 0.02 | 1500000 | Service & Verification | 0.001 | 0.02 | Dry-aatcc-TIME |  | Y | 2700 - Time - sec
Dry-aatcc |  |  |  | 10: 4200; 12: N/A; 14: sec; 15: 4200; 17: 4200; 18: 4200; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Dry-aatcc-TIME | 1 |  | 4200
Dry-aatcc |  |  |  | 9: Cool down - sec; 10: 600; 12: N/A; 11: Cool down - sec; 13: Thời gian làm mát; 14: sec; 15: 600; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Dry-aatcc-TIME | 1
Dry-aatcc |  |  |  | 9: Labtex/Vortex
Dry-aatcc |  |  |  | 9: Delicate; 11: Delicate; 21: #N/A |  | #N/A | #N/A | Dry-aatcc- | 1
Dry-aatcc |  |  |  | 9: Temp (*) - min - °C; 12: N/A; 11: Temp (*) - min - °C; 13: Nhiệt độ tối thiểu; 14: ℃; 15: N/A; 16: N/A; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Dry-aatcc-DEG | 1 | Std
Dry-aatcc | [7 empty] | Temp (*) - max - °C | 60 | ±6 | Temp (*) - max - °C | Nhiệt độ tối đa | ℃ | 60 | 6 | 54 | 66 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Dry-aatcc-DEG | 1 | Std
Dry-aatcc
Dry-aatcc |  |  |  | 9: Normal & Permenant; 11: Normal & Permenant; 21: #N/A |  | #N/A | #N/A | Dry-aatcc- |  | Y
Dry-aatcc |  |  |  | 9: Temp (*) - min - °C; 12: N/A; 11: Temp (*) - min - °C; 13: Nhiệt độ tối thiểu; 14: ℃; 15: N/A; 16: N/A; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Dry-aatcc-DEG |  | Std
Dry-aatcc | [7 empty] | Temp (*) - max - °C | 68 | ±6 | Temp (*) - max - °C | Nhiệt độ tối đa | ℃ | 68 | 6 | 62 | 74 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Dry-aatcc-DEG |  | Std
Dry-aatcc |  |  |  | 9: Whirlpool | [6 empty] | 0
Dry-aatcc |  |  |  | 9: Low
Dry-aatcc |  |  |  | 9: Temp (*) - min - °C; 12: N/A; 11: Temp (*) - min - °C; 13: Nhiệt độ tối thiểu; 14: ℃; 15: N/A; 16: N/A; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Dry-aatcc-DEG
Dry-aatcc | [7 empty] | Temp (*) - max - °C | 60 | ±6 | Temp (*) - max - °C | Nhiệt độ tối đa | ℃ | 60 | 6 | 54 | 66 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Dry-aatcc-DEG
Dry-aatcc
Dry-aatcc |  |  |  | 9: Medium; 11: Medium | [6 empty] | Medium
Dry-aatcc |  |  |  | 9: Temp (*) - min - °C; 12: N/A; 11: Temp (*) - min - °C; 13: Nhiệt độ tối thiểu; 14: ℃; 15: N/A; 16: N/A; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Dry-aatcc-DEG |  |  | Temp (*) - min - °C
Dry-aatcc | [7 empty] | Temp (*) - max - °C | 68 | ±6 | Temp (*) - max - °C | Nhiệt độ tối đa | ℃ | 68 | 6 | 62 | 74 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Dry-aatcc-DEG |  |  | 68 - Temp (*) - max - °C
Dry-aatcc | [27 empty] | 0
Dry-aatcc |  |  |  | 9: High; 11: High | [6 empty] | High
Dry-aatcc |  |  |  | 9: Temp (*) - min - °C; 12: N/A; 11: Temp (*) - min - °C; 13: Nhiệt độ tối thiểu; 14: ℃; 15: N/A; 16: N/A; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Dry-aatcc-DEG |  |  | Temp (*) - min - °C
Dry-aatcc | [7 empty] | Temp (*) - max - °C | 68 | ±6 | Temp (*) - max - °C | Nhiệt độ tối đa | ℃ | 68 | 6 | 62 | 74 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Dry-aatcc-DEG
Dry-aatcc
Dry-aatcc |  |  |  | 9: Others; 11: *
Dry-aatcc | bar: 9; 11
Dry-aatcc | bar: 9; 11
Dry-aatcc | bar: 9; 11
Dry-aatcc | bar: 9; 11
Dry-aatcc | bar: 9; 11
Wash-I | NAL.TM.0020, Adidas 4.03, AATCC LP1 |  | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | Labtex, Vortex | Top Loading Washing Machine | Máy giặt cửa trên | Water level Agitation speed Spin speed Washing time Spin time | Water level - L | 57 | ±4 | Water level - L | Mực nước | L | 57 | 4 | 53 | 61 | VOL |  | 0.2 | 1500000 | Service & Verification | 0,01 | 0.07 | Wash-I-VOL |  | Y
Wash-I |  |  |  | 10: 68; 12: ±4; 14: L; 15: 68; 16: 4; 17: 64; 18: 72; 19: VOL; 21: 0.07 |  | 0,01 | 0.07 | Wash-I-VOL |  | Y
Wash-I |  |  |  | 10: 72; 12: ±4; 14: L; 15: 72; 16: 4; 17: 68; 18: 76; 19: VOL
Wash-I |  |  |  | 9: Normal; 11: Normal; 21: #N/A |  | #N/A | #N/A | Wash-I- |  | Y | Normal
Wash-I | [7 empty] | Agitation - spm | 86 | ±2 | Agitation - spm | Tốc độ lắc | spm | 86 | 2 | 5 | 10 | SPM |  | 0.02 |  |  | 0.001 | 0.02 | Wash-I-SPM |  | Y
Wash-I | [7 empty] | Wash time - sec | 960 | ±60 | Wash time - sec | Thời gian giặt | sec | 960 | 60 | 900 | 1020 | TIME |  | 0.02 | 1500000 | Service & Verification | 0.001 | 0.02 | Wash-I-TIME |  | Y
Wash-I | [7 empty] | Spin Speed (*) - rpm | 660 | ±15 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 660 | 15 | 645 | 675 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-I-RPM |  | Y
Wash-I | [7 empty] | Spin time - sec | 300 | ±60 | Spin time - sec | Thời gian vắt | sec | 300 | 60 | 240 | 360 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-I-TIME
Wash-I | [27 empty] | 0
Wash-I |  |  |  | 9: Delicate; 11: Delicate; 21: #N/A |  | #N/A | #N/A | Wash-I- |  |  | Delicate
Wash-I | [7 empty] | Agitation - spm | 27 | ±2 | Agitation - spm | Tốc độ lắc | spm | 27 | 2 | 25 | 29 | SPM |  | 0.02 | 1500000 | Service & Verification | 0.001 | 0.02 | Wash-I-SPM |  | Y | 27 - Agitation - spm
Wash-I | [7 empty] | Wash time - sec | 510 | ±60 | Wash time - sec | Thời gian giặt | sec | 510 | 60 | 450 | 570 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-I-TIME |  | Y | 510 - Wash time - sec
Wash-I | [7 empty] | Spin Speed (*) - rpm | 500 | ±15 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 500 | 15 | 485 | 515 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-I-RPM |  |  | 500 - Spin Speed (*) - rpm
Wash-I | [7 empty] | Spin time - sec | 300 | ±60 | Spin time - sec | Thời gian vắt | sec | 300 | 60 | 240 | 360 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-I-TIME |  |  | 300 - Spin time - sec
Wash-I | [27 empty] | 0
Wash-I |  |  |  | 9: Permanent; 11: Permanent; 21: #N/A |  | #N/A | #N/A | Wash-I-
Wash-I | [7 empty] | Agitation - spm | 86 | ±2 | Agitation - spm | Tốc độ lắc | spm | 86 | 2 | 84 | 88 | SPM |  | 0.02 |  |  | 0.001 | 0.02 | Wash-I-SPM
Wash-I | [7 empty] | Wash time - sec | 720 | ±60 | Wash time - sec | Thời gian giặt | sec | 720 | 60 | 660 | 780 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-I-TIME
Wash-I | [7 empty] | Spin Speed (*) - rpm | 500 | ±15 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 500 | 15 | 485 | 515 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-I-RPM
Wash-I | [7 empty] | Spin time - sec | 300 | ±60 | Spin time - sec | Thời gian vắt | sec | 300 | 60 | 240 | 360 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-I-TIME
Wash-I |  |  |  | 21: #N/A |  | #N/A | #N/A | Wash-I-
Wash-I | [7 empty] | Wash Temp (*) - °C | 27 | ±3 | Wash Temp (*) - °C | Nhiệt độ | ℃ | 27 | 3 | 24 | 30 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-I-DEG
Wash-I |  |  |  | 10: 41; 12: ±3; 14: ℃; 15: 41; 16: 3; 17: 38; 18: 44; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-I-DEG
Wash-I |  |  |  | 10: 49; 12: ±3; 14: ℃; 15: 49; 16: 3; 17: 46; 18: 52; 19: DEG
Wash-I |  |  |  | 10: 60; 12: ±3; 14: ℃; 15: 60; 16: 3; 17: 57; 18: 63; 19: DEG
Wash-I |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-I-
Wash-I |  |  |  | 9: 1…; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-I- |  | Y
Wash-I | bar: 9; 11
Wash-I | [27 empty] | * | bar: 9; 11
Wash-I | bar: 9; 11
Wash-I | [27 empty] | * | bar: 9; 11
Wash-IID | AATCC LPI |  | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | Labtex, Vortex, Whirlpool 4800, 4840 | Washer | Máy giặt | Water level Agitation speed Spin speed Washing time Spin time | Water level - 1.8kg - L | 72 | ±8 | Water level - 1.8kg - L | Mực nước | L | 72 | 8 | 64 | 80 | VOL |  | 0.2 | 1500000 | Service & Verification | 0,01 | 0.07 | Wash-IID-VOL |  | Y
Wash-IID |  |  | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 |  |  |  |  | Water level - 3.6kg - L | 83 | ±4 | Water level - 3.6kg - L | Mực nước | L | 83 | 4 | 79 | 87 | VOL |  | 0.07 |  |  | 0,01 | 0.07 | Wash-IID-VOL |  | Y
Wash-IID |  |  |  | 21: #N/A |  | #N/A | #N/A | Wash-IID- |  | Y
Wash-IID |  |  |  | 9: Normal; 11: Normal; 21: #N/A |  | #N/A | #N/A | Wash-IID- |  | Y
Wash-IID | [7 empty] | Agitation - spm | 86 | ±5 | Agitation - spm | Tốc độ lắc | spm | 86 | 5 | 81 | 91 | SPM |  | 0.02 |  |  | 0.001 | 0.02 | Wash-IID-SPM |  | Y
Wash-IID | [7 empty] | Wash time - sec | 960 | ±120 | Wash time - sec | Thời gian giặt | sec | 960 | 120 | 840 | 1080 | TIME |  | 0.02 | 1500000 | Service & Verification | 0.001 | 0.02 | Wash-IID-TIME |  | Y
Wash-IID | [7 empty] | Spin Speed (*) - rpm | 660 | ±15 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 660 | 15 | 645 | 675 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-IID-RPM |  | Y
Wash-IID | [7 empty] | Spin time - sec | 300-600 | 300-600 | Spin time - sec | Thời gian vắt | sec | 450 | 150 | 300 | 600 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-IID-TIME
Wash-IID |  |  |  | 21: #N/A |  | #N/A | #N/A | Wash-IID-
Wash-IID |  |  |  | 9: Delicate; 11: Delicate; 21: #N/A |  | #N/A | #N/A | Wash-IID-
Wash-IID | [7 empty] | Agitation - spm | 27 | ±5 | Agitation - spm | Tốc độ lắc | spm | 27 | 5 | 22 | 32 | SPM |  | 0.02 | 1500000 | Service & Verification | 0.001 | 0.02 | Wash-IID-SPM |  | Y
Wash-IID | [7 empty] | Wash time - sec | 510 | ±60 | Wash time - sec | Thời gian giặt | sec | 510 | 60 | 450 | 570 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-IID-TIME |  | Y
Wash-IID | [7 empty] | Spin Speed (*) - rpm | 500 | ±15 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 500 | 15 | 485 | 515 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-IID-RPM
Wash-IID | [7 empty] | Spin time - sec | 300-600 | 300-600 | Spin time - sec | Thời gian vắt | sec | 450 | 150 | 300 | 600 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-IID-TIME
Wash-IID |  |  |  | 21: #N/A |  | #N/A | #N/A | Wash-IID-
Wash-IID |  |  |  | 9: Wash Temp (*) - °C; 11: Wash Temp (*) - °C; 13: Nhiệt độ
Wash-IID | [7 empty] | Cold | 16 | ±4.2 | Cold |  | ℃ | 16 | 4.2 | 11.8 | 20.2 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-IID-DEG
Wash-IID | [7 empty] | Cool | 30 | ±4.2 | Cool |  | ℃ | 30 | 4.2 | 25.8 | 34.2 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-IID-DEG
Wash-IID | [7 empty] | Warm | 44 | ±4.2 | Warm |  | ℃ | 44 | 4.2 | 39.8 | 48.2 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-IID-DEG
Wash-IID | [7 empty] | Hot | 54 | ±4.2 | Hot |  | ℃ | 54 | 4.2 | 49.8 | 58.2 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-IID-DEG
Wash-IID |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-IID-
Wash-IID |  |  |  | 9: 1…; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-IID- |  | Y
Wash-IID | bar: 9; 11
Wash-IID | bar: 9; 11
Wash-IID | bar: 9; 11
Wash-IID | bar: 9; 11
Wash-IIIB | AATCC LP1, Table IIIB |  | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | Labtex, Vortex, Whirlpool 4815(2018) | Washer | Máy giặt | Water level Agitation speed Spin speed Washing time Spin time | Water level - L | 68 (18gal) | ±8 | Water level - L | Mực nước |  | 68 | 8 | 60 | 76 | VOL |  | 0.2 | 1500000 | Service & Verification | 0,01 | 0.07 | Wash-IIIB-VOL |  | Y
Wash-IIIB |  |  |  | 9: Normal (Jean); 11: Normal (Jean); 21: #N/A |  | #N/A | #N/A | Wash-IIIB- |  | Y
Wash-IIIB | [7 empty] | Agitation - spm | 60 | ±5 | Agitation - spm | Tốc độ lắc | spm | 60 | 5 | 5 | 18 | SPM |  | 0.02 |  |  | 0.001 | 0.02 | Wash-IIIB-SPM |  | Y
Wash-IIIB | [7 empty] | Wash time - sec | 840 | ±120 | Wash time - sec | Thời gian giặt | sec | 840 | 120 | 720 | 960 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-IIIB-TIME |  | Y
Wash-IIIB | [6 empty] | Water level Agitation speed Spin speed Washing time Spin time | Spin Speed (*) - rpm | 660 | ±20 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 660 | 20 | 640 | 680 | RPM |  | 0.1 | 1500000 | Service & Verification | 0.1 | 0.1 | Wash-IIIB-RPM |  | Y
Wash-IIIB | [7 empty] | Spin time - sec | 300-600 | 300-600 | Spin time - sec | Thời gian vắt | sec | 450 | 150 | 300 | 600 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-IIIB-TIME |  | Y
Wash-IIIB |  |  |  | 21: #N/A |  | #N/A | #N/A | Wash-IIIB-
Wash-IIIB |  |  |  | 9: Wash Temp (*) - °C; 11: Wash Temp (*) - °C; 13: Nhiệt độ; 21: #N/A |  | #N/A | #N/A | Wash-IIIB-
Wash-IIIB |  |  |  | 9: Cold; 10: 16; 12: ±4.2; 14: ℃; 15: 16; 16: 4.2; 17: 11.8; 18: 20.2; 19: DEG
Wash-IIIB |  |  |  | 9: Warm; 10: 35; 12: ±4.2; 14: ℃; 15: 35; 16: 4.2; 17: 30.8; 18: 39.2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-IIIB-DEG
Wash-IIIB |  |  |  | 9: Hot; 10: 54; 12: ±4.2; 14: ℃; 15: 54; 16: 4.2; 17: 49.8; 18: 58.2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-IIIB-DEG
Wash-IIIB |  |  |  | 9: Delicate; 11: Delicate; 21: #N/A |  | #N/A | #N/A | Wash-IIIB-
Wash-IIIB | [7 empty] | Agitation - spm | 70 | ±5 | Agitation - spm | Tốc độ lắc | spm | 70 | 5 | 65 | 75 | SPM |  | 0.02 |  |  | 0.001 | 0.02 | Wash-IIIB-SPM
Wash-IIIB | [7 empty] | Wash time - sec | 480 | ±120 | Wash time - sec | Thời gian giặt | sec | 480 | 120 | 360 | 600 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-IIIB-TIME
Wash-IIIB | [7 empty] | Spin Speed (*) - rpm | 660 | ±20 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 660 | 20 | 640 | 680 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-IIIB-RPM
Wash-IIIB | [7 empty] | Spin time - sec | 60-360 | 60-360 | Spin time - sec | Thời gian vắt | sec | 210 | 150 | 60 | 360 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-IIIB-TIME
Wash-IIIB |  |  |  | 21: #N/A |  | #N/A | #N/A | Wash-IIIB-
Wash-IIIB |  |  |  | 9: Wash Temp (*) - °C; 11: Wash Temp (*) - °C; 13: Nhiệt độ; 21: #N/A |  | #N/A | #N/A | Wash-IIIB-
[8 empty] | Cold | 16 | ±4.2 |  |  | ℃ | 16 | 4.2 | 11.8 | 20.2 | DEG
Wash-IIIB |  |  |  | 9: Warm; 10: 30; 12: ±4.2; 14: ℃; 15: 30; 16: 4.2; 17: 25.8; 18: 34.2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-IIIB-DEG
Wash-IIIB |  |  |  | 9: Hot; 10: 54; 12: ±4.2; 14: ℃; 15: 54; 16: 4.2; 17: 49.8; 18: 58.2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-IIIB-DEG
Wash-IIIB |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-IIIB-
Wash-IIIB |  |  |  | 9: 1…; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-IIIB- |  | Y
Wash-IIIB | bar: 9; 11
Wash-IIIB | bar: 9; 11
Wash-IIIB | bar: 9; 11
Wash-IIIB | bar: 9; 11
Wash-ISO-Top | ISO 6330 - Table D.1 |  | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | 5: Labtex M6T, Vortex; 6: Washer; 7: Máy giặt; 8: Water level Agitation speed Spin speed Washing time Spin time; 9: Table D.1; 11: Table D.1; 21: 0.2; 22: 1500000 | Service & Verification | #N/A | #N/A | Wash-ISO-Top- |  | Y
Wash-ISO-Top |  |  |  | 9: Wash; 11: Wash; 21: #N/A |  | #N/A | #N/A | Wash-ISO-Top- |  | Y
Wash-ISO-Top | [7 empty] | Temp (*) - °C - 1B | 60 | ±3 | Temp (*) - °C | Nhiệt độ | ℃ | 60 | 3 | 57 | 63 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-ISO-Top-DEG |  | Y
Wash-ISO-Top |  |  |  | 9: Temp (*) - °C - 2B, 3B; 10: 49; 12: ±3; 14: ℃; 15: 49; 16: 3; 17: 5; 18: 18; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-Top-DEG |  | Y
Wash-ISO-Top |  |  |  | 9: Temp (*) - °C - 4B, 5B; 10: 41; 12: ±3; 14: ℃; 15: 41; 16: 3; 17: 5; 18: 10; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-Top-DEG |  | Y
Wash-ISO-Top | [7 empty] | Temp (*) - °C - 6B, 7B, 8B | 27 | ±3 |  |  | ℃ | 27 | 3 | 24 | 30 | DEG |  | 0.2 | 1500000 | Service & Verification | 0.1 | 0.2 | Wash-ISO-Top-DEG |  | Y
Wash-ISO-Top |  |  |  | 9: Temp (*) - °C - 9B, 10B, 11B; 10: 16; 12: ±3; 14: ℃; 15: 16; 16: 3; 17: 13; 18: 19; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-Top-DEG |  | Y
Wash-ISO-Top | [7 empty] | Water level -mm - 1B→10B | 297 | ±25 | Water level -mm | Mực nước | mm | 297 | 25 | 272 | 322 | VOL |  | 0.07 |  |  | 0,01 | 0.07 | Wash-ISO-Top-VOL
Wash-ISO-Top |  |  |  | 9: Water level -mm - 11B; 10: 398.5; 12: ±17.8; 14: mm; 15: 398.5; 16: 17.8; 17: 380.7; 18: 416.3; 19: VOL; 21: 0.07 |  | 0,01 | 0.07 | Wash-ISO-Top-VOL
Wash-ISO-Top | [7 empty] | Wash time - sec - 1B, 2B, 4B, 6B, 9N | 720 | N/A | Wash time - sec | Thời gian giặt | sec | 720 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-Top-TIME
Wash-ISO-Top |  |  |  | 9: Wash time - sec - 3B, 5B, 7B, 10B; 10: 600; 12: N/A; 14: sec; 15: 600; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-Top-TIME
Wash-ISO-Top |  |  |  | 9: Wash time - sec - 8B, 11B; 10: 480; 12: N/A; 14: sec; 15: 480; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-Top-TIME
Wash-ISO-Top |  |  |  | 9: Rinse; 11: Rinse; 21: #N/A |  | #N/A | #N/A | Wash-ISO-Top-
Wash-ISO-Top | [7 empty] | Water level -mm - 1B→11B | 297 | ±25 | Water level -mm | Mực nước | mm | 297 | 25 | 272 | 322 | VOL |  | 0.07 |  |  | 0,01 | 0.07 | Wash-ISO-Top-VOL
Wash-ISO-Top | [7 empty] | Rinse time - sec | 180 | N/A | Rinse time - sec - sec | Thời gian xả | sec | 180 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-Top-TIME
Wash-ISO-Top
Wash-ISO-Top |  |  |  | 9: Spin; 11: Spin; 21: #N/A |  | #N/A | #N/A | Wash-ISO-Top-
Wash-ISO-Top | [7 empty] | Spin Speed (*) - rpm - 1B, 2B, 4B, 6B, 9B | 613 to 640 | 613 to 640 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 613 | 613 to 640 | 613 | 640 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-ISO-Top-RPM
Wash-ISO-Top | [7 empty] | Spin Speed (*) - rpm - 3B, 5B, 7B, 8B, 10B, 11B | 399 to 420 | 399 to 420 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 399 | 399 to 420 | 399 | 420 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-ISO-Top-RPM |  |  | Cotton - 30℃ - 30℃
Wash-ISO-Top | [7 empty] | Spin time - sec- 1B, 2B, 4B, 6B, 9B | 360 | N/A | Spin time - sec | Thời gian vắt | sec | 360 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-Top-TIME
Wash-ISO-Top |  |  |  | 9: Spin time - sec - 3B, 5B, 7B, 8B, 10B, 11B; 10: 240; 12: N/A; 14: sec; 15: 240; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-Top-TIME |  | Y
Wash-ISO-Top |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-ISO-Top- |  |  | Cotton - 40⁰C - 40⁰C
Wash-ISO-Top |  |  |  | 9: 1…; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-ISO-Top- |  | Y
Wash-ISO-Top | bar: 9; 11
Wash-ISO-Top | [27 empty] | Cotton - 60⁰C - 60⁰C | bar: 9; 11
Wash-ISO-Top | bar: 9; 11
Wash-ISO-Top | bar: 9; 11
Wash-nike | NAL.TM.0020 |  | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | Labtex, Vortex | Top Loading Washing Machine | Máy giặt cửa trên | Water level Agitation speed Spin speed Washing time Spin time | Water level - L | 45 (12gal) | ±4 | Water level - L | Mực nước | L | 45 | 4 | 41 | 49 | VOL |  | 0.2 | 1500000 | Service & Verification | 0,01 | 0.07 | Wash-nike-VOL |  | Y
Wash-nike |  |  |  | 10: 60 (16gal); 12: ±4; 14: L; 15: 60; 16: 4; 17: 56; 18: 64; 19: VOL; 21: 0.07 |  | 0,01 | 0.07 | Wash-nike-VOL |  | Y
Wash-nike |  |  |  | 9: Normal; 11: Normal; 21: #N/A |  | #N/A | #N/A | Wash-nike- |  | Y | Normal
Wash-nike | [7 empty] | Agitation - spm | 86 | ±2 | Agitation - spm | Tốc độ lắc | spm | 86 | 2 | 5 | 10 | SPM |  | 0.02 |  |  | 0.001 | 0.02 | Wash-nike-SPM |  | Y
Wash-nike | [7 empty] | Wash time - sec | 960 | ±60 | Wash time - sec | Thời gian giặt | sec | 960 | 60 | 900 | 1020 | TIME |  | 0.02 | 1500000 | Service & Verification | 0.001 | 0.02 | Wash-nike-TIME |  | Y
Wash-nike | [7 empty] | Spin Speed (*) - rpm | 660 | ±15 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 660 | 15 | 645 | 675 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-nike-RPM |  | Y
Wash-nike | [7 empty] | Spin time - sec | 300 | ±60 | Spin time - sec | Thời gian vắt | sec | 300 | 60 | 240 | 360 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-nike-TIME
Wash-nike | [27 empty] | 0
Wash-nike |  |  |  | 9: Delicate; 11: Delicate; 21: #N/A |  | #N/A | #N/A | Wash-nike- |  |  | Delicate
Wash-nike | [7 empty] | Agitation - spm | 27 | ±2 | Agitation - spm | Tốc độ lắc | spm | 27 | 2 | 25 | 29 | SPM |  | 0.02 | 1500000 | Service & Verification | 0.001 | 0.02 | Wash-nike-SPM |  | Y | 27 - Agitation - spm
Wash-nike | [7 empty] | Wash time - sec | 510 | ±60 | Wash time - sec | Thời gian giặt | sec | 510 | 60 | 450 | 570 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-nike-TIME |  | Y | 510 - Wash time - sec
Wash-nike | [7 empty] | Spin Speed (*) - rpm | 500 | ±15 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 500 | 15 | 485 | 515 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-nike-RPM |  |  | 500 - Spin Speed (*) - rpm
Wash-nike | [7 empty] | Spin time - sec | 300 | ±60 | Spin time - sec | Thời gian vắt | sec | 300 | 60 | 240 | 360 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-nike-TIME |  |  | 300 - Spin time - sec
Wash-nike | [27 empty] | 0
Wash-nike |  |  |  | 9: Permanent; 11: Permanent; 21: #N/A |  | #N/A | #N/A | Wash-nike-
Wash-nike | [7 empty] | Agitation - spm | 86 | ±2 | Agitation - spm | Tốc độ lắc | spm | 86 | 2 | 84 | 88 | SPM |  | 0.02 |  |  | 0.001 | 0.02 | Wash-nike-SPM
Wash-nike | [7 empty] | Wash time - sec | 720 | ±60 | Wash time - sec | Thời gian giặt | sec | 720 | 60 | 660 | 780 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-nike-TIME
Wash-nike | [7 empty] | Spin Speed (*) - rpm | 500 | ±15 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 500 | 15 | 485 | 515 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-nike-RPM
Wash-nike | [7 empty] | Spin time - sec | 300 | ±60 | Spin time - sec | Thời gian vắt | sec | 300 | 60 | 240 | 360 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-nike-TIME
Wash-nike |  |  |  | 21: #N/A |  | #N/A | #N/A | Wash-nike-
Wash-nike | [7 empty] | Wash Temp (*) - °C | 30 | ±2 | Wash Temp (*) - °C | Nhiệt độ | ℃ | 30 | 2 | 28 | 32 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-nike-DEG
Wash-nike |  |  |  | 10: 40; 12: ±2; 14: ℃; 15: 40; 16: 2; 17: 38; 18: 42; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-nike-DEG
Wash-nike |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-nike-
Wash-nike |  |  |  | 9: 1…; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-nike- |  | Y
Wash-nike | bar: 9; 11
Wash-nike | [27 empty] | * | bar: 9; 11
Wash-nike | bar: 9; 11
Wash-nike | [27 empty] | * | bar: 9; 11
Miele-nike-6065 | Nike Apparel Wash Protocol - Front load |  | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | Miele PW 6055, 6065 | Washer | Máy giặt | Water level Agitation speed (spm) Spin speed (rpm) Washing time Spin time | Water level - L | 12.5 | ±1.5 | Water level - L | Mực nước | L | 12.5 | 1.5 | 11 | 14 | VOL |  | 0.07 | 2000000 | Service & Calibration | 0,01 | 0.07 | Miele-nike-6065-VOL
Miele-nike-6065 |  |  |  | 9: Cottons Uni 30; 11: Cotton Uni; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Miele-nike-6065-
Miele-nike-6065 | [7 empty] | Wash Temp (*) - °C | 30 | ±3 | Wash Temp (*) - °C | Nhiệt độ | ℃ | 30 | 3 | 27 | 33 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-nike-6065-DEG
Miele-nike-6065 | [7 empty] | Spin Speed (*) - rpm | 1000 | ±50 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 1000 | 50 | 950 | 1050 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-nike-6065-RPM
Miele-nike-6065 | [7 empty] | Total time - sec | 3600 | ±300 | Total time - sec | Tổng thời gian | sec | 3600 | 300 | 3300 | 3900 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-nike-6065-TIME |  |  | 1000rpm - 1000rpm
Miele-nike-6065 |  |  |  | 9: Cottons Uni 40
Miele-nike-6065 | [7 empty] | Wash Temp (*) - °C | 40 | ±3 | Wash Temp (*) - °C | Nhiệt độ | ℃ | 40 | 3 | 37 | 43 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-nike-6065-DEG
Miele-nike-6065 | [7 empty] | Spin Speed (*) - rpm | 1000 | ±50 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 1000 | 50 | 950 | 1050 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-nike-6065-RPM
Miele-nike-6065 | [7 empty] | Total time - sec | 3600 | ±300 | Total time - sec | Tổng thời gian | sec | 3600 | 300 | 3300 | 3900 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-nike-6065-TIME |  |  | 1000rpm - 1000rpm
Miele-nike-6065 |  |  |  | 9: Cottons Uni 60
Miele-nike-6065 | [7 empty] | Wash Temp (*) - °C | 60 | ±3 | Wash Temp (*) - °C | Nhiệt độ | ℃ | 60 | 3 | 57 | 63 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-nike-6065-DEG
Miele-nike-6065 | [7 empty] | Spin Speed (*) - rpm | 1000 | ±50 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 1000 | 50 | 950 | 1050 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-nike-6065-RPM
Miele-nike-6065 | [7 empty] | Total time - sec | 3600 | ±300 | Total time - sec | Tổng thời gian | sec | 3600 | 300 | 3300 | 3900 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-nike-6065-TIME |  |  | 1000rpm - 1000rpm
Miele-nike-6065 |  |  |  | 9: Woolens 30; 11: Woolen; 21: #N/A |  | #N/A | #N/A | Miele-nike-6065-
Miele-nike-6065 | [7 empty] | Wash Temp (*) - °C | 30 | ±3 | Wash Temp (*) - °C | Nhiệt độ | ℃ | 30 | 3 | 27 | 33 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-nike-6065-DEG
Miele-nike-6065 | [7 empty] | Spin Speed (*) - rpm | 800 | ±50 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 800 | 50 | 750 | 850 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-nike-6065-RPM
Miele-nike-6065 | [7 empty] | Total time - sec | 3600 | ±300 | Total time - sec | Tổng thời gian | sec | 3600 | 300 | 3300 | 3900 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-nike-6065-TIME |  |  | 1000rpm - 1000rpm
Miele-nike-6065 |  |  |  | 9: Woolens 40
Miele-nike-6065 | [7 empty] | Wash Temp (*) - °C | 40 | ±3 | Wash Temp (*) - °C | Nhiệt độ | ℃ | 40 | 3 | 37 | 43 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-nike-6065-DEG
Miele-nike-6065 | [7 empty] | Spin Speed (*) - rpm | 800 | ±50 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 800 | 50 | 750 | 850 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-nike-6065-RPM
Miele-nike-6065 | [7 empty] | Total time - sec | 3600 | ±300 | Total time - sec | Tổng thời gian | sec | 3600 | 300 | 3300 | 3900 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-nike-6065-TIME |  |  | 1000rpm - 1000rpm
Miele-nike-6065 |  |  |  | 9: Delicate
Miele-nike-6065 | [7 empty] | Wash Temp (*) - °C | 30 | ±3 | Wash Temp (*) - °C | Nhiệt độ | ℃ | 30 | 3 | 27 | 33 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-nike-6065-DEG
Miele-nike-6065 | [7 empty] | Spin Speed (*) - rpm | 600 | ±50 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 600 | 50 | 550 | 650 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-nike-6065-RPM
Miele-nike-6065 | [7 empty] | Total time - sec | N/A | N/A | Total time - sec | Tổng thời gian | sec | N/A | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-nike-6065-TIME |  |  | 1000rpm - 1000rpm
Miele-nike-6065 |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Miele-nike-6065-
Miele-nike-6065 |  |  |  | 9: Nếu đo từng chương trình thì lấy thông số cho từng chương trình. Đo gộp thì chỉ đo 1 thông sô đại diện cho các chương trình còn lại
Miele-nike-6065 |  |  |  | 9: 1…; 11: *; 21: #N/A |  | #N/A | #N/A | Miele-nike-6065-
Miele-nike-6065 |  |  |  | 9: 2…; 11: *; 21: #N/A |  | #N/A | #N/A | Miele-nike-6065-
Miele-nike-6065 |  |  |  | 9: 3…; 11: *; 21: #N/A |  | #N/A | #N/A | Miele-nike-6065-
Miele-nike-6065 |  |  |  | 9: 4…; 11: *; 21: #N/A |  | #N/A | #N/A | Miele-nike-6065-
Miele-nike-6065 | bar: 9; 11
Miele-nike-tav | Nike Apparel Wash Protocol - Front load |  | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | Miele PWM 906 | Washer | Máy giặt | Water level Agitation speed (spm) Spin speed (rpm) Washing time Spin time | Water level - L | 12.5 | ±1.5 | Water level - L | Mực nước | L | 12.5 | 1.5 | 11 | 14 | VOL |  | 0.07 | 2000000 | Service & Calibration | 0,01 | 0.07 | Miele-nike-tav-VOL
Miele-nike-tav |  |  |  | 9: Cottons Plus 30; 11: TAV-N30.1; 21: #N/A |  | #N/A | #N/A | Miele-nike-tav-
Miele-nike-tav | [7 empty] | Wash Temp (*) - °C | 30 | ±3 | Wash Temp (*) - °C | Nhiệt độ | ℃ | 30 | 3 | 27 | 33 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-nike-tav-DEG
Miele-nike-tav | [7 empty] | Spin Speed (*) - rpm | 1000 | ±50 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 1000 | 50 | 950 | 1050 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-nike-tav-RPM |  |  | 40spm @ 33min - 40spm @ 33min | Thông số này chưa có trong COC hiện tại
Miele-nike-tav | [7 empty] | Total time - sec | 3600 | ±300 | Total time - sec | Tổng thời gian | sec | 3600 | 300 | 3300 | 3900 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-nike-tav-TIME |  |  | 33min - 33min
Miele-nike-tav |  |  |  | 9: Cottons Plus 40
Miele-nike-tav |  |  |  | 9: Wash Temp (*) - °C; 10: 40; 12: ±3; 14: ℃; 15: 40; 16: 3; 17: 37; 18: 43; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Miele-nike-tav-DEG
Miele-nike-tav | [7 empty] | Spin Speed (*) - rpm | 1000 | ±50 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 1000 | 50 | 950 | 1050 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-nike-tav-RPM |  |  | 40spm @ 33min - 40spm @ 33min | Thông số này chưa có trong COC hiện tại
Miele-nike-tav | [7 empty] | Total time - sec | 3600 | ±300 | Total time - sec | Tổng thời gian | sec | 3600 | 300 | 3300 | 3900 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-nike-tav-TIME |  |  | 33min - 33min
Miele-nike-tav |  |  |  | 9: Cottons Plus 60
Miele-nike-tav |  |  |  | 9: Wash Temp (*) - °C; 10: 60; 12: ±3; 14: ℃; 15: 60; 16: 3; 17: 57; 18: 63; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Miele-nike-tav-DEG
Miele-nike-tav | [7 empty] | Spin Speed (*) - rpm | 1000 | ±50 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 1000 | 50 | 950 | 1050 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-nike-tav-RPM |  |  | 40spm @ 33min - 40spm @ 33min | Thông số này chưa có trong COC hiện tại
Miele-nike-tav | [7 empty] | Total time - sec | 3600 | ±300 | Total time - sec | Tổng thời gian | sec | 3600 | 300 | 3300 | 3900 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-nike-tav-TIME |  |  | 33min - 33min
Miele-nike-tav |  |  |  | 9: Woolens 30; 11: TAV-N30.9; 21: #N/A |  | #N/A | #N/A | Miele-nike-tav-
Miele-nike-tav | [7 empty] | Wash Temp (*) - °C | 30 | ±3 | Wash Temp (*) - °C | Nhiệt độ | ℃ | 30 | 3 | 27 | 33 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-nike-tav-DEG
Miele-nike-tav | [7 empty] | Spin Speed (*) - rpm | 800 | ±50 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 800 | 50 | 750 | 850 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-nike-6065-RPM
Miele-nike-tav | [7 empty] | Total time - sec | 3600 | ±300 | Total time - sec | Tổng thời gian | sec | 3600 | 300 | 3300 | 3900 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-nike-tav-TIME |  |  | 33min - 33min
Miele-nike-tav |  |  |  | 9: Woolens 40
Miele-nike-tav | [7 empty] | Wash Temp (*) - °C | 40 | ±3 | Wash Temp (*) - °C | Nhiệt độ | ℃ | 40 | 3 | 37 | 43 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-nike-tav-DEG
Miele-nike-tav | [7 empty] | Spin Speed (*) - rpm | 800 | ±50 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 800 | 50 | 750 | 850 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-nike-tav-RPM |  |  | 40spm @ 33min - 40spm @ 33min | Thông số này chưa có trong COC hiện tại
Miele-nike-tav | [7 empty] | Total time - sec | 3600 | ±300 | Total time - sec | Tổng thời gian | sec | 3600 | 300 | 3300 | 3900 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-nike-tav-TIME |  |  | 33min - 33min
Miele-nike-tav |  |  |  | 9: Delicate
Miele-nike-tav | [7 empty] | Wash Temp (*) - °C | 30 | ±3 | Wash Temp (*) - °C | Nhiệt độ | ℃ | 30 | 3 | 27 | 33 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-nike-tav-DEG
Miele-nike-tav | [7 empty] | Spin Speed (*) - rpm | 600 | ±50 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 600 | 50 | 550 | 650 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-nike-tav-RPM |  |  | 40spm @ 33min - 40spm @ 33min | Thông số này chưa có trong COC hiện tại
Miele-nike-tav | [7 empty] | Total time - sec | N/A | N/A | Total time - sec | Tổng thời gian | sec | N/A | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-nike-tav-TIME |  |  | 33min - 33min
Miele-nike-tav |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Miele-nike-tav-
Miele-nike-tav |  |  |  | 9: 1…; 11: *; 21: #N/A |  | #N/A | #N/A | Miele-nike-tav-
Miele-nike-tav | bar: 9; 11
Miele-nike-tav | bar: 9; 11
Miele-nike-tav | bar: 9; 11
Miele-nike-tav | bar: 9; 11
Miele-adidas | Adidas PHX-AP0701 |  | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | Miele PW 6055, 6065, PW 906 | Washer | Máy giặt | Water level Agitation speed (spm) Spin speed (rpm) Washing time Spin time | Water level - L | 6 - 10 | 6 - 10 | Water level - L | Mực nước | L | 6 | 6 - 10 | 6 | 10 | VOL |  | 0.07 | 2000000 | Service & Calibration | 0,01 | 0.07 | Miele-adidas-VOL
Miele-adidas | [7 empty] | Spin Speed (*) - rpm | 600 | ±50 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 400 | 50 | 350 | 450 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-adidas-RPM |  | Y
Miele-adidas |  |  |  | 9: TAV-A2 (Tùy cài đặt của thiết bị); 11: TAV-A2; 21: #N/A |  | #N/A | #N/A | Miele-adidas-
Miele-adidas | [7 empty] | Wash time - sec | 1200 | ±180 | Wash time - sec | Thời gian giặt | sec | 1200 | 180 | 1020 | 1380 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  | Y
Miele-adidas | [7 empty] | Temp (*) - °C | 30 | ±3 | Temp (*) - °C | Nhiệt độ | ℃ | 30 | 3 | 27 | 33 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-adidas-DEG |  |  | 12.5L - 12.5L | Thông số này chưa có trong COC hiện tại
Miele-adidas | [7 empty] | Final Spin 1 - sec | 120 | ±60 | Final Spin 1 - sec | Thời gian vắt 1 | sec | 120 | 60 | 60 | 180 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Rinse 1 - sec | 240 | ±180 | Rinse 1 - sec | Thời gian xả 1 | sec | 240 | 180 | 60 | 420 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Total Spin 1 - sec | N/A | N/A | Total Spin 1 - sec | Thời gian vắt 1 | sec | N/A | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Final Spin 2 - sec | 120 | ±60 | Final Spin 2 - sec | Thời gian vắt 2 | sec | 120 | 60 | 60 | 180 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Total Spin 2 - sec | N/A | N/A | Total Spin 2 - sec | Thời gian vắt 2 | sec | N/A | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Rinse 2 - sec | 240 | ±180 | Rinse 2 - sec | Thời gian xả 2 | sec | 240 | 180 | 60 | 420 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  |  | 500rpm - 500rpm
Miele-adidas | [7 empty] | Final Spin 3 - sec | 240 | ±120 | Final Spin 3 - sec | Thời gian vắt 3 | sec | N/A | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Total Spin 3 - sec | N/A | N/A | Total Spin 3 - sec | Thời gian vắt 3 | sec | N/A | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas |  |  |  | 9: Rinse 3 - sec; 11: Rinse 3 - sec; 13: Thời gian xả 3; 14: sec; 15: 0; 16: 120; 17: -120; 18: 120; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Miele-adidas-TIME |  |  | 60mins - 60mins
Miele-adidas |  |  |  | 9: TAV-A3/A8/A5; 11: TAV-A3; 21: #N/A |  | #N/A | #N/A | Miele-adidas-
Miele-adidas | [7 empty] | Wash time - sec | 1500 | ±300 | Wash time - sec | Thời gian giặt | sec | 1500 | 300 | 1200 | 1800 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  | Y
Miele-adidas | [7 empty] | Temp (*) - °C | 30 | ±3 | Temp (*) - °C | Nhiệt độ | ℃ | 30 | 3 | 27 | 33 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-adidas-DEG |  |  | 12.5L - 12.5L | Thông số này chưa có trong COC hiện tại
Miele-adidas | [7 empty] | Spin 1 - sec | 240 | ±120 | Spin 1 - sec | Thời gian vắt 1 | sec | 240 | 120 | 120 | 360 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Rinse 1 - sec | 420 | ±120 | Rinse 1 - sec | Thời gian xả 1 | sec | 420 | 120 | 300 | 540 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Spin 2 - sec | 180 | ±120 | Spin 2 - sec | Thời gian vắt 2 | sec | 180 | 120 | 60 | 300 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Rinse 2 - sec | 420 | ±120 | Rinse 2 - sec | Thời gian xả 2 | sec | 420 | 120 | 300 | 540 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  |  | 500rpm - 500rpm
Miele-adidas | [7 empty] | Spin 3 - sec | 180 | ±120 | Spin 3 - sec | Thời gian vắt 3 | sec | 180 | 120 | 60 | 300 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Rinse 3 - sec | 360 | ±120 | Rinse 3 - sec | Thời gian xả 3 | sec | 360 | 120 | 240 | 480 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  |  | 60mins - 60mins
Miele-adidas | [7 empty] | Spin 4 - sec | 360 | ±120 | Spin 4 - sec | Thời gian vắt 4 | sec | 360 | 120 | 240 | 480 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas |  |  |  | 9: TAV-A4; 11: TAV-A4/9; 21: #N/A |  | #N/A | #N/A | Miele-adidas-
Miele-adidas | [7 empty] | Wash time - sec | 1800 | ±300 | Wash time - sec | Thời gian giặt | sec | 1800 | 300 | 1500 | 2100 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  | Y
Miele-adidas | [7 empty] | Temp (*) - °C | 30 | ±3 | Temp (*) - °C | Nhiệt độ | ℃ | 30 | 3 | 27 | 33 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-adidas-DEG |  | Y
Miele-adidas | [7 empty] | Spin 1 - sec | 180 | ±120 | Spin 1 - sec | Thời gian vắt 1 | sec | 180 | 120 | 60 | 300 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Rinse 1 - sec | 420 | ±120 | Rinse 1 - sec | Thời gian xả 1 | sec | 420 | 120 | 300 | 540 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Spin 2 - sec | 180 | ±120 | Spin 2 - sec | Thời gian vắt 2 | sec | 180 | 120 | 60 | 300 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Rinse 2 - sec | 420 | ±120 | Rinse 2 - sec | Thời gian xả 2 | sec | 420 | 120 | 300 | 540 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  |  | Cotton - 30℃ - 30℃
Miele-adidas | [7 empty] | Spin 3 - sec | 360 | ±120 | Spin 3 - sec | Thời gian vắt 3 | sec | 360 | 120 | 240 | 480 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas |  |  |  | 9: TAV-A5; 11: TAV-A5; 21: #N/A |  | #N/A | #N/A | Miele-adidas- |  |  | 12.5L - 12.5L | Thông số này chưa có trong COC hiện tại
Miele-adidas | [7 empty] | Wash time - sec | 1500 | ±300 | Wash time - sec | Thời gian giặt | sec | 1500 | 300 | 1200 | 1800 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  | Y | 500rpm - 500rpm
Miele-adidas | [7 empty] | Temp (*) - °C | 40 | ±3 | Temp (*) - °C | Nhiệt độ | ℃ | 40 | 3 | 37 | 43 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-adidas-DEG |  | Y | 60mins - 60mins
Miele-adidas | [7 empty] | Spin 1 - sec | 240 | ±120 | Spin 1 - sec | Thời gian vắt 1 | sec | 240 | 120 | 120 | 360 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  |  | Cotton - 30℃ - 30℃
Miele-adidas | [7 empty] | Rinse 1 - sec | 420 | ±120 | Rinse 1 - sec | Thời gian xả 1 | sec | 420 | 120 | 300 | 540 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Spin 2 - sec | 180 | ±120 | Spin 2 - sec | Thời gian vắt 2 | sec | 180 | 120 | 60 | 300 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Rinse 2 - sec | 420 | ±120 | Rinse 2 - sec | Thời gian xả 2 | sec | 420 | 120 | 300 | 540 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Spin 3 - sec | 180 | ±120 | Spin 3 - sec | Thời gian vắt 3 | sec | 180 | 120 | 60 | 300 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Rinse 3 - sec | 360 | ±120 | Rinse 3 - sec | Thời gian xả 3 | sec | 360 | 120 | 240 | 480 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Spin 4 - sec | 360 | ±120 | Spin 4 - sec | Thời gian vắt 4 | sec | 360 | 120 | 240 | 480 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas |  |  |  | 9: TAV-A6/7; 11: TAV-A6/7; 21: #N/A |  | #N/A | #N/A | Miele-adidas- |  |  | 12.5L - 12.5L | Thông số này chưa có trong COC hiện tại
Miele-adidas | [7 empty] | Wash time - sec | 1800 | ±300 | Wash time - sec | Thời gian giặt | sec | 1800 | 300 | 1500 | 2100 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  |  | 500rpm - 500rpm
Miele-adidas | [7 empty] | Temp (*) - °C | 40 | ±3 | Temp (*) - °C | Nhiệt độ | ℃ | 40 | 3 | 37 | 43 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-adidas-DEG |  |  | 60mins - 60mins
Miele-adidas | [7 empty] | Spin 1 - sec | 180 | ±120 | Spin 1 - sec | Thời gian vắt 1 | sec | 180 | 120 | 60 | 300 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Rinse 1 - sec | 420 | ±120 | Rinse 1 - sec | Thời gian xả 1 | sec | 420 | 120 | 300 | 540 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  |  | Cotton - 30℃ - 30℃
Miele-adidas | [7 empty] | Spin 2 - sec | 180 | ±120 | Spin 2 - sec | Thời gian vắt 2 | sec | 180 | 120 | 60 | 300 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Rinse 2 - sec | 420 | ±120 | Rinse 2 - sec | Thời gian xả 2 | sec | 420 | 120 | 300 | 540 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  |  | 12.5L - 12.5L | Thông số này chưa có trong COC hiện tại
Miele-adidas | [7 empty] | Spin 3 - sec | 360 | ±120 | Spin 3 - sec | Thời gian vắt 3 | sec | 360 | 120 | 240 | 480 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  |  | 500rpm - 500rpm
Miele-adidas | [27 empty] | 60mins - 60mins
Miele-adidas | [27 empty] | 0
Miele-adidas
Miele-adidas
Miele-adidas
Miele-adidas
Miele-adidas |  |  |  | 9: TAV-A8; 11: TAV-A8; 21: #N/A |  | #N/A | #N/A | Miele-adidas- |  | Y
Miele-adidas | [7 empty] | Wash time - sec | 1500 | ±300 | Wash time - sec | Thời gian giặt | sec | 1500 | 300 | 1200 | 1800 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  | Y
Miele-adidas | [7 empty] | Temp (*) - °C | 60 | ±3 | Temp (*) - °C | Nhiệt độ | ℃ | 60 | 3 | 57 | 63 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-adidas-DEG |  | Y
Miele-adidas | [7 empty] | Spin 1 - sec | 240 | ±120 | Spin 1 - sec | Thời gian vắt 1 | sec | 240 | 120 | 120 | 360 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Rinse 1 - sec | 420 | ±120 | Rinse 1 - sec | Thời gian xả 1 | sec | 420 | 120 | 300 | 540 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME
Miele-adidas | [7 empty] | Spin 2 - sec | 180 | ±120 | Spin 2 - sec | Thời gian vắt 2 | sec | 180 | 120 | 60 | 300 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  |  | 180 - Spin 2 - sec
Miele-adidas | [7 empty] | Rinse 2 - sec | 420 | ±120 | Rinse 2 - sec | Thời gian xả 2 | sec | 420 | 120 | 300 | 540 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  |  | 420 - Rinse 2 - sec
Miele-adidas | [7 empty] | Spin 3 - sec | 180 | ±120 | Spin 3 - sec | Thời gian vắt 3 | sec | 180 | 120 | 60 | 300 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  |  | 180 - Spin 3 - sec
Miele-adidas | [7 empty] | Rinse 3 - sec | 360 | ±120 | Rinse 3 - sec | Thời gian xả 3 | sec | 360 | 120 | 240 | 480 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  |  | 360 - Rinse 3 - sec
Miele-adidas | [7 empty] | Spin 4 - sec | 360 | ±120 | Spin 4 - sec | Thời gian vắt 4 | sec | 360 | 120 | 240 | 480 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-adidas-TIME |  |  | 360 - Spin 4 - sec
Miele-adidas |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Miele-adidas- |  |  | *
Miele-adidas |  |  |  | 9: Chương trình A3/ A5/ A8 thông số như nhau khác nhau thông số nhiệt độ (A3-30, A5-40, A8-60) |  |  |  | Miele-adidas-
Miele-adidas |  |  |  | 9: 1…; 11: *; 21: #N/A |  | #N/A | #N/A | Miele-adidas-
Miele-adidas | bar: 9; 11
Miele-adidas | [27 empty] | * | bar: 9; 11
Miele-adidas | bar: 9; 11
Miele-adidas | bar: 9; 11
Miele-decathlon | DS0034 Version 4 | Dimensional & Appearance Change - International, Finished Garment - European | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | 5: Miele Washer - Model PWM 906 DP; 6: Washer; 7: Máy giặt; 8: Wash Water Level Spin Speed Wash Time Agitation/Wash Time Spin time Total Run Time; 9: Colors (Normal); 11: Colors (Normal); 21: #N/A |  | #N/A | #N/A | Miele-decathlon- |  |  | Colors (Normal) | Thông số này chưa có trong COC hiện tại
Miele-decathlon | [7 empty] | Temp (*) - °C | 40 | ±5 | Temp (*) - °C | Nhiệt độ | ℃ | 40 | 5 | 35 | 45 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-decathlon-DEG |  |  | 40 - Temp (*) - °C
Miele-decathlon | [7 empty] | Time spent at the required T ͦC - sec | 3600 | ±600 | Time spent at the required T ͦC - sec | Thời gian giặt giữ nhiệt | sec | 3600 | 600 | 3000 | 4200 | TIME |  | 0.02 | 2000000 | Service & Calibration | 0.001 | 0.02 | Miele-decathlon-TIME |  |  | 3600 - Time spent at the required T ͦC - sec
Miele-decathlon | [7 empty] | Spin Speed (*) - rpm | 1000 | ±100 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 1000 | 100 | 900 | 1100 | RPM |  | 0.02 |  |  | 0.1 | 0.1 | Miele-decathlon-RPM |  |  | 1000 - Spin Speed (*) - rpm
Miele-decathlon | [7 empty] | Time of total cycle - sec | 7200 | ±600 | Time of total cycle - sec | Tổng thời gian | sec | 7200 | 600 | 6600 | 7800 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-decathlon-TIME
Miele-decathlon | [7 empty] | Time of rising & spinning - sec | 900 - 1800 | 900 - 1800 | Time of rising & spinning - sec | Thời gian xả & vắt | sec | 1350 | 450 | 900 | 1800 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-decathlon-TIME
Miele-decathlon |  |  |  | 9: Minimum iron (Delicate); 11: Minimum iron (Delicate); 21: #N/A |  | #N/A | #N/A | Miele-decathlon- |  |  | Minimum iron (Delicate)
Miele-decathlon | [7 empty] | Temp (*) - °C | 30 | ±5 | Temp (*) - °C | Nhiệt độ | ℃ | 30 | 5 | 25 | 35 | DEG |  | 0.2 | 2000000 | Service & Calibration | 0.1 | 0.2 | Miele-decathlon-DEG |  |  | 30 - Temp (*) - °C
Miele-decathlon | [7 empty] | Time spent at the required T ͦC - sec | 1650 | ±300 | Time spent at the required T ͦC - sec | Thời gian giặt giữ nhiệt | sec | 1650 | 300 | 1350 | 1950 | TIME |  | 0.02 | 2000000 | Service & Calibration | 0.001 | 0.02 | Miele-decathlon-TIME |  |  | 1650 - Time spent at the required T ͦC - sec
Miele-decathlon | [7 empty] | Spin Speed (*) - rpm | 700 | ±100 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 700 | 100 | 600 | 800 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-decathlon-RPM |  |  | 700 - Spin Speed (*) - rpm
Miele-decathlon | [7 empty] | Time of total cycle - sec | 3300 | ±600 | Time of total cycle - sec | Tổng thời gian | sec | 3300 | 600 | 2700 | 3900 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-decathlon-TIME |  |  | 3300 - Time of total cycle - sec | Thông số này chưa có trong COC hiện tại
Miele-decathlon | [7 empty] | Time of rising & spinning - sec | 900 - 1800 | 900 - 1800 | Time of rising & spinning - sec | Thời gian xả & vắt | sec | 1350 | 450 | 900 | 1800 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-decathlon-TIME
Miele-decathlon |  |  |  | 9: Clean ECE A; 11: Clean ECE A; 21: #N/A |  | #N/A | #N/A | Miele-decathlon- |  |  | Clean ECE A
Miele-decathlon | [7 empty] | Temp (*) - °C | 90 | ±5 | Temp (*) - °C | Nhiệt độ | ℃ | 90 | 5 | 85 | 95 | DEG |  | 0.2 | 2000000 | Service & Calibration | 0.1 | 0.2 | Miele-decathlon-DEG |  |  | 90 - Temp (*) - °C
Miele-decathlon |  |  |  | 9: Spin Speed (*) - rpm; 10: 900; 12: N/A; 11: Spin Speed (*) - rpm; 13: Tốc độ vắt; 14: rpm; 15: 900; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Miele-decathlon-RPM |  |  | 900 - Spin Speed (*) - rpm
Miele-decathlon |  |  |  | 9: Time of total cycle - sec; 10: 5400; 12: N/A; 11: Time of total cycle - sec; 13: Tổng thời gian; 14: sec; 15: 5400; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Miele-decathlon-TIME |  |  | 5400 - Time of total cycle - sec | Thông số này chưa có trong COC hiện tại
Miele-decathlon | [7 empty] | Time of rising & spinning - sec | 900 - 1800 | 900 - 1800 | Time of rising & spinning - sec | Thời gian xả & vắt | sec | 1350 | 450 | 900 | 1800 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-decathlon-TIME
Miele-decathlon |  |  |  | 9: Check Fixed Capacity
Miele-decathlon |  |  |  | 9: Normal Program (Cottons plus) with load/ no load
Miele-decathlon | [7 empty] | Temp (*) - °C | 40 | ±5 | Temp (*) - °C | Nhiệt độ | ℃ | 40 | 5 | 35 | 45 | DEG |  | 0.2 | 2000000 | Service & Calibration | 0.1 | 0.2 | Miele-decathlon-DEG |  |  | 40 - Temp (*) - °C
Miele-decathlon | [7 empty] | Time spent at the required T ͦC - sec | 60 | ±10 | Time spent at the required T ͦC - sec | Thời gian giặt giữ nhiệt | sec | 60 | 10 | 50 | 70 | TIME |  | 0.02 | 2000000 | Service & Calibration | 0.001 | 0.02 | Miele-decathlon-TIME |  |  | 60 - Time spent at the required T ͦC - sec
Miele-decathlon | [7 empty] | Time of total cycle - sec | 120 | ±10 | Time of total cycle - sec | Tổng thời gian | sec | 120 | 10 | 110 | 130 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-decathlon-TIME |  |  | 120 - Time of total cycle - sec | Thông số này chưa có trong COC hiện tại
Miele-decathlon |  |  |  | 9: Time of rising & spinning - sec; 12: 15-30; 11: Time of rising & spinning - sec; 13: Thời gian xả & vắt; 14: sec; 17: 0; 18: 0; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Miele-decathlon-TIME
Miele-decathlon |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Miele-decathlon- |  |  | *
Miele-decathlon |  |  |  | 9: 1…; 11: *; 21: #N/A |  | #N/A | #N/A | Miele-decathlon- |  |  | *
Miele-decathlon |  |  |  | 9: 2…; 11: *; 21: #N/A |  | #N/A | #N/A | Miele-decathlon- |  |  | *
Miele-decathlon |  |  |  | 9: 3…; 11: *; 21: #N/A |  | #N/A | #N/A | Miele-decathlon-
Miele-decathlon |  |  |  | 9: 4…; 11: *; 21: #N/A |  | #N/A | #N/A | Miele-decathlon-
Miele-decathlon |  |  |  | 9: 5…; 11: *; 21: #N/A; 22: 2000000 | Service & Calibration | #N/A | #N/A | Miele-decathlon-
Miele-decathlon1 | DS179 | Dimensional & Appearance Change - International, Finished Garment - European | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | 5: Miele Washer - Model PWM 906 DP; 6: Washer; 7: Máy giặt; 8: Wash Water Level Spin Speed Wash Time Agitation/Wash Time Spin time Total Run Time; 9: Aggressive; 11: Aggressive; 21: #N/A |  | #N/A | #N/A | Miele-decathlon1- |  |  | Aggressive | Thông số này chưa có trong COC hiện tại
Miele-decathlon1 | [7 empty] | Temp (*) - °C | 90 | ±5 | Temp (*) - °C | Nhiệt độ | ℃ | 90 | 5 | 85 | 95 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-decathlon1-DEG |  |  | 90 - Temp (*) - °C
Miele-decathlon1 | [7 empty] | Time spent at the required T ͦC - sec | 3600 | ±600 | Time spent at the required T ͦC - sec | Thời gian giặt giữ nhiệt | sec | 3600 | 600 | 3000 | 4200 | TIME |  | 0.02 | 2000000 | Service & Calibration | 0.001 | 0.02 | Miele-decathlon1-TIME |  |  | 3600 - Time spent at the required T ͦC - sec
Miele-decathlon1 | [7 empty] | Spin Speed (*) - rpm | 1000 | ±100 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 1000 | 100 | 900 | 1100 | RPM |  | 0.02 |  |  | 0.1 | 0.1 | Miele-decathlon1-RPM |  |  | 1000 - Spin Speed (*) - rpm
Miele-decathlon1 | [7 empty] | Time of total cycle - sec | 7200 | ±600 | Time of total cycle - sec | Tổng thời gian | sec | 7200 | 600 | 6600 | 7800 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-decathlon1-TIME
Miele-decathlon1 | [7 empty] | Time of rising & spinning - sec | 900 - 1800 | 900 - 1800 | Time of rising & spinning - sec | Thời gian xả & vắt | sec | 1350 | 450 | 900 | 1800 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-decathlon1-TIME
Miele-decathlon1 |  |  |  | 9: Synthetic; 11: Synthetic; 21: #N/A |  | #N/A | #N/A | Miele-decathlon1- |  |  | Synthetic
Miele-decathlon1 | [7 empty] | Temp (*) - °C | 60 | ±5 | Temp (*) - °C | Nhiệt độ | ℃ | 60 | 5 | 55 | 65 | DEG |  | 0.2 | 2000000 | Service & Calibration | 0.1 | 0.2 | Miele-decathlon1-DEG |  |  | 60 - Temp (*) - °C
Miele-decathlon1 | [7 empty] | Time spent at the required T ͦC - sec | 1650 | ±300 | Time spent at the required T ͦC - sec | Thời gian giặt giữ nhiệt | sec | 1650 | 300 | 1350 | 1950 | TIME |  | 0.02 | 2000000 | Service & Calibration | 0.001 | 0.02 | Miele-decathlon1-TIME |  |  | 1650 - Time spent at the required T ͦC - sec
Miele-decathlon1 | [7 empty] | Spin Speed (*) - rpm | 700 | ±100 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 700 | 100 | 600 | 800 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-decathlon1-RPM |  |  | 700 - Spin Speed (*) - rpm
Miele-decathlon1 | [7 empty] | Time of total cycle - sec | 6000 | ±600 | Time of total cycle - sec | Tổng thời gian | sec | 6000 | 600 | 5400 | 6600 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-decathlon1-TIME |  |  | 6000 - Time of total cycle - sec | Thông số này chưa có trong COC hiện tại
Miele-decathlon1 | [7 empty] | Time of rising & spinning - sec | 900 - 1800 | 900 - 1800 | Time of rising & spinning - sec | Thời gian xả & vắt | sec | 1350 | 450 | 900 | 1800 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-decathlon1-TIME
Miele-decathlon1 |  |  |  | 9: Cottons 90; 11: Cotton 90; 21: #N/A |  | #N/A | #N/A | Miele-decathlon1- |  |  | Cotton 90
Miele-decathlon1 | [7 empty] | Temp (*) - °C | 90 | ±5 | Temp (*) - °C | Nhiệt độ | ℃ | 90 | 5 | 85 | 95 | DEG |  | 0.2 | 2000000 | Service & Calibration | 0.1 | 0.2 | Miele-decathlon1-DEG |  |  | 90 - Temp (*) - °C
Miele-decathlon1 | [7 empty] | Spin Speed (*) - rpm | 1000 | N/A | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 1000 | N/A | #VALUE! | #VALUE! | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-decathlon1-RPM |  |  | 1000 - Spin Speed (*) - rpm
Miele-decathlon1 | [7 empty] | Time of total cycle - sec | 3600 | N/A | Time of total cycle - sec | Tổng thời gian | sec | 3600 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-decathlon1-TIME |  |  | 3600 - Time of total cycle - sec | Thông số này chưa có trong COC hiện tại
Miele-decathlon1 | [7 empty] | Time of rising & spinning - sec | 900 - 1800 | 900 - 1800 | Time of rising & spinning - sec | Thời gian xả & vắt | sec | 1350 | 450 | 900 | 1800 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-decathlon1-TIME |  |  | 900 - 1800 - Time of rising & spinning - sec
Miele-decathlon1 |  |  |  | 9: Others; 21: #N/A |  | #N/A | #N/A | Miele-decathlon1- |  |  | 0
Miele-decathlon1 | [22 empty] | #N/A | #N/A | Miele-decathlon1- |  |  | 0 | bar: 9
Miele-decathlon1 |  |  |  | 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Miele-decathlon1-DEG |  |  | 0
Miele-decathlon1 |  |  |  | 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Miele-decathlon1-RPM |  |  | 0
Miele-decathlon1 |  |  |  | 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Miele-decathlon1-TIME |  |  | 0
Miele-decathlon1 |  |  |  | 19: TIME; 21: 0.02; 22: 2000000 | Service & Calibration | 0.001 | 0.02 | Miele-decathlon1-TIME |  |  | 0
Miele-puma | Puma - F-01 |  | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | 6: Washer; 7: Máy giặt; 9: TAV-P1-30 (Puma - 30)/ Cottons Plus 30; 11: TAV-P1-30 (Puma - 30) |  | #N/A | #N/A | Miele-puma-
Miele-puma | [7 empty] | Temp (*) - °C | 30 | ±3 | Temp (*) - °C | Nhiệt độ | ℃ | 30 | 3 | 27 | 33 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-puma-DEG |  |  | 30 - Temp (*) - °C
Miele-puma | [7 empty] | Total time - sec | 3600 | ±10% | Total time - sec | Tổng thời gian | sec | 3600 | 360 | 3240 | 3960 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-puma-TIME |  | Y
Miele-puma |  |  |  | 9: Washing time - sec; 11: Washing time - sec; 13: Thời gian giặt; 14: sec; 17: 0; 18: 0; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Miele-puma-TIME
Miele-puma | [7 empty] | Spin Speed (*) - rpm | 600 | ±100 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 600 | 100 | 500 | 700 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-puma-RPM |  |  | 600 - Spin Speed (*) - rpm
Miele-puma |  |  |  | 9: TAV-P2-40 (Puma - 40)/ Cottons Plus 40; 11: TAV-P2-40 (Puma - 40) | [5 empty] | Y | TAV-P2-40 (Puma - 40)
Miele-puma | [7 empty] | Temp (*) - °C | 40 | ±3 | Temp (*) - °C | Nhiệt độ | ℃ | 40 | 3 | 37 | 43 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Miele-puma-DEG |  |  | 40 - Temp (*) - °C
Miele-puma | [7 empty] | Total time - sec | 3600 | ±10% | Total time - sec | Tổng thời gian | sec | 3600 | 360 | 3240 | 3960 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Miele-puma-TIME |  | Y
Miele-puma |  |  |  | 9: Washing time - sec; 11: Washing time - sec; 13: Thời gian giặt; 14: sec; 17: 0; 18: 0; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Miele-puma-TIME
Miele-puma | [7 empty] | Spin Speed (*) - rpm | 600 | ±100 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 600 | 100 | 500 | 700 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Miele-puma-RPM |  |  | 600 - Spin Speed (*) - rpm
Miele-puma |  |  |  | 9: Others; 11: * | [6 empty] | *
Miele-puma |  |  |  | 9: Cottons Plus 30, 40 tương tự chương trình Puma 30, 40 tùy thuộc vào thiết bị khách hàng cài đặt sẽ hiển thị.; 11: * | [6 empty] | *
Miele-puma | [27 empty] | * | bar: 9; 11
Miele-puma | [27 empty] | * | bar: 9; 11
Miele-puma | [27 empty] | * | bar: 9; 11
Miele-puma | [27 empty] | * | bar: 9; 11
Miele-open |  |  | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | Miele PWM 906 | Washer | Máy giặt |  | Spin Speed (*) - rpm | 400 | ±50 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 400 | 50 | 350 | 450 | RPM |  | 0.1 |  | Service & Calibration | 0.1 | 0.1 | Miele-open-RPM |  |  | 400 - Spin Speed (*) - rpm
Miele-open |  |  |  | 10: 800; 12: ±50; 14: rpm; 15: 800; 16: 50; 17: 750; 18: 850; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Miele-open-RPM |  |  | 800 | Thông số này chưa có trong COC hiện tại
Miele-open |  |  |  | 10: 1000; 12: ±50; 14: rpm; 15: 1000; 16: 50; 17: 950; 18: 1050; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Miele-open-RPM |  |  | 1000
Miele-open |  |  |  | 10: 1200; 12: ±50; 14: rpm; 15: 1200; 16: 50; 17: 1150; 18: 1250; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Miele-open-RPM |  |  | 1200
Miele-open |  |  |  | 10: 1400; 12: ±50; 14: rpm; 15: 1400; 16: 50; 17: 1350; 18: 1450; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Miele-open-RPM |  |  | 1400
Miele-open |  |  |  | 10: 1600; 12: ±50; 14: rpm; 15: 1600; 16: 50; 17: 1550; 18: 1650; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Miele-open-RPM |  |  | 1600
Miele-open |  |  |  | 14: rpm; 16: 50; 17: -50; 18: 50; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Miele-open-RPM |  |  | 0
Miele-open |  |  |  | 9: Cotton Plus; 11: Cotton Plus; 17: 0; 18: 0; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Miele-open-RPM |  |  | Cotton Plus
Miele-open |  |  |  | 9: Water level - L; 11: Water level - L; 13: Mực nước; 14: L; 17: 0; 18: 0; 19: VOL; 21: 0.07 |  | 0,01 | 0.07 | Miele-open-VOL |  |  | Water level - L
Miele-open |  |  |  | 9: Total time - sec; 10: 3600; 11: Total time - sec; 13: Tổng thời gian; 14: sec; 15: 3600; 17: 3600; 18: 3600; 19: TIME; 21: 0.02 |  | #N/A | 0.02 | Miele-open-TIME |  |  | 3600 - Total time - sec
Miele-open | [7 empty] | Wash Temp (*) - °C | 30 | ±3 | Wash Temp (*) - °C | Nhiệt độ | ℃ | 30 | 3 | 27 | 33 | DEG |  | 0.2 |  |  | #N/A | 0.2 | Miele-open-DEG |  |  | 30 - Wash Temp (*) - °C
Miele-open |  |  |  | 10: 40; 15: 40; 16: 3; 17: 37; 18: 43; 19: DEG; 21: 0.2 |  | #N/A | 0.2 | Miele-open-DEG |  |  | 40
Miele-open |  |  |  | 10: 60; 12: ±3; 15: 60; 16: 3; 17: 57; 18: 63; 19: DEG; 21: 0.2 |  | #N/A | 0.2 | Miele-open-DEG |  |  | 60
Miele-open |  |  |  | 9: Minimum Iron 30; 11: Minimum Iron 30; 17: 0; 18: 0; 19: DEG; 21: 0.2 |  | #N/A | 0.2 | Miele-open-DEG |  |  | Minimum Iron 30
Miele-open |  |  |  | 9: Water level - L; 12: ±3; 11: Water level - L; 13: Mực nước; 14: L; 17: 0; 18: 0; 19: VOL; 21: 0.07 |  | 0,01 | 0.07 | Miele-open-VOL |  |  | Water level - L
Miele-open |  |  |  | 9: Total time - sec; 10: 2700; 11: Total time - sec; 13: Tổng thời gian; 14: sec; 15: 2700; 17: 2700; 18: 2700; 19: TIME; 21: 0.02 |  | #N/A | 0.02 | Miele-open-TIME |  |  | 2700 - Total time - sec
Miele-open | [7 empty] | Wash Temp (*) - °C | 30 | ±3 | Wash Temp (*) - °C | Nhiệt độ | ℃ | 30 | 3 | 27 | 33 | DEG |  | 0.2 |  |  | #N/A | 0.2 | Miele-open-DEG |  |  | 30 - Wash Temp (*) - °C
Miele-open |  |  |  | 9: Minimum Iron 40; 11: Minimum Iron 40; 17: 0; 18: 0; 19: DEG; 21: 0.2 |  | #N/A | 0.2 | Miele-open-DEG |  |  | Minimum Iron 40
Miele-open |  |  |  | 9: Water level - L; 11: Water level - L; 13: Mực nước; 14: L; 17: 0; 18: 0; 19: VOL; 21: 0.07 |  | 0,01 | 0.07 | Miele-open-VOL |  |  | Water level - L
Miele-open |  |  |  | 9: Total time - sec; 10: 2700; 11: Total time - sec; 13: Tổng thời gian; 14: sec; 15: 2700; 17: 2700; 18: 2700; 19: TIME; 21: 0.02 |  | #N/A | 0.02 | Miele-open-TIME |  |  | 2700 - Total time - sec
Miele-open | [7 empty] | Wash Temp (*) - °C | 40 |  | Wash Temp (*) - °C | Nhiệt độ | ℃ | 40 | 3 | 37 | 43 | DEG |  | 0.2 |  |  | #N/A | 0.2 | Miele-open-DEG |  |  | 40 - Wash Temp (*) - °C
Miele-open |  |  |  | 17: 0; 18: 0; 19: DEG; 21: 0.2 |  | #N/A | 0.2 | Miele-open-DEG |  |  | 0
Miele-open |  |  |  | 17: 0; 18: 0; 19: DEG; 21: 0.2 |  | #N/A | 0.2 | Miele-open-DEG |  |  | 0
Wash-other | On request |  | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | 5: Electrolux, LG, Panasonic; 6: Washer; 7: Máy giặt; 8: Water level Agitation speed (spm) Spin speed (rpm) Washing time Spin time; 9: Program 1 (theo chương trình máy); 11: Program 1; 21: #N/A; 22: 2000000 | Service & Calibration | #N/A | #N/A | Wash-other- |  | Y | Program 1
Wash-other |  |  |  | 9: Water level - L; 12: ±; 11: Water level - L; 13: Mực nước; 14: L; 17: 0; 18: 0; 19: VOL; 21: 0.07 |  | 0,01 | 0.07 | Wash-other-VOL
Wash-other |  |  |  | 9: Agitation - spm; 12: ±; 11: Agitation - spm; 13: Tốc độ lắc; 14: spm; 17: 0; 18: 0; 19: SPM; 21: 0.02 |  | 0.001 | 0.02 | Wash-other-SPM |  | Y
Wash-other |  |  |  | 9: Wash time - sec; 12: ±; 11: Wash time - sec; 13: Thời gian giặt; 14: sec; 17: 0; 18: 0; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-other-TIME |  | Y | Wash time - sec
Wash-other |  |  |  | 9: Spin Speed (*) - rpm; 12: ±; 11: Spin Speed (*) - rpm; 13: Tốc độ vắt; 14: rpm; 17: 0; 18: 0; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Wash-other-RPM |  | Y | Spin Speed (*) - rpm
Wash-other |  |  |  | 9: Spin time - sec; 12: ±; 11: Spin time - sec; 13: Thời gian vắt; 14: sec; 17: 0; 18: 0; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-other-TIME |  | Y | Spin time - sec | Thông số này chưa có trong COC hiện tại
Wash-other |  |  |  | 9: Wash Temp (*) - °C; 12: ±; 11: Wash Temp (*) - °C; 13: Nhiệt độ; 14: G; 17: 0; 18: 0; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-other-DEG |  |  | Wash Temp (*) - °C | Thông số này chưa có trong COC hiện tại
Wash-other | [27 empty] | 0
Wash-other |  |  |  | 9: Program 2; 11: Program 2; 21: #N/A |  | #N/A | #N/A | Wash-other- |  |  | Program 2
Wash-other |  |  |  | 9: Water level - L; 12: ±; 11: Water level - L; 13: Mực nước; 14: L; 17: 0; 18: 0; 19: VOL; 21: 0.07 |  | 0,01 | 0.07 | Wash-other-VOL
Wash-other |  |  |  | 9: Agitation - spm; 12: ±; 11: Agitation - spm; 13: Tốc độ lắc; 14: spm; 17: 0; 18: 0; 19: SPM; 21: 0.02 |  | 0.001 | 0.02 | Wash-other-SPM |  |  | Agitation - spm
Wash-other |  |  |  | 9: Wash time - sec; 12: ±; 11: Wash time - sec; 13: Thời gian giặt; 14: sec; 17: 0; 18: 0; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-other-TIME |  |  | Wash time - sec
Wash-other |  |  |  | 9: Spin Speed (*) - rpm; 12: ±; 11: Spin Speed (*) - rpm; 13: Tốc độ vắt; 14: rpm; 17: 0; 18: 0; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Wash-other-RPM |  | Y
Wash-other |  |  |  | 9: Spin time - sec; 12: ±; 11: Spin time - sec; 13: Thời gian vắt; 14: sec; 17: 0; 18: 0; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-other-TIME |  | Y
Wash-other |  |  |  | 9: Wash Temp (*) - °C; 12: ±; 11: Wash Temp (*) - °C; 13: Nhiệt độ; 14: ℃; 17: 0; 18: 0; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-other-DEG |  | Y
Wash-other
Wash-other |  |  |  | 9: Program 3; 11: Program 3; 21: #N/A |  | #N/A | #N/A | Wash-other- |  | Y
Wash-other |  |  |  | 9: Water level - L; 12: ±; 11: Water level - L; 13: Mực nước; 14: L; 17: 0; 18: 0; 19: VOL; 21: 0.07 |  | 0,01 | 0.07 | Wash-other-VOL
Wash-other |  |  |  | 9: Agitation - spm; 12: ±; 11: Agitation - spm; 13: Tốc độ lắc; 14: spm; 17: 0; 18: 0; 19: SPM; 21: 0.02 |  | 0.001 | 0.02 | Wash-other-SPM |  | Y
Wash-other |  |  |  | 9: Wash time - sec; 12: ±; 11: Wash time - sec; 13: Thời gian giặt; 14: sec; 17: 0; 18: 0; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-other-TIME |  | Y
Wash-other |  |  |  | 9: Spin Speed (*) - rpm; 12: ±; 11: Spin Speed (*) - rpm; 13: Tốc độ vắt; 14: rpm; 17: 0; 18: 0; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Wash-other-RPM |  | Y
Wash-other |  |  |  | 9: Spin time - sec; 12: ±; 11: Spin time - sec; 13: Thời gian vắt; 14: sec; 17: 0; 18: 0; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-other-TIME |  | Y
Wash-other |  |  |  | 9: Wash Temp (*) - °C; 12: ±; 11: Wash Temp (*) - °C; 13: Nhiệt độ; 14: ℃; 17: 0; 18: 0; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-other-DEG
Wash-other
Wash-other |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-other-
Wash-other | bar: 9; 11
Wash-other | bar: 9; 11
Wash-other | bar: 9; 11
Wash-other | bar: 9; 11
Wash-other | bar: 9; 11
Wash-other | bar: 9; 11
Wascator-tav.fw | Wascator Calibration Guidelines, Adidas FGT-14 | Standardised European Washing | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | Wascator | Washer | Máy giặt |  | Water level - mm | 100 | ±5 | Water level - mm | Mực nước | mm | 100 | 5 | 95 | 105 | DIM |  | 0.017 | 3000000 | Service & Calibration | 0.5 | 0.017 | Wascator-tav.fw-DIM |  | Y
Wascator-tav.fw |  |  |  | 10: 130; 12: ±5; 14: mm; 15: 130; 16: 5; 17: 125; 18: 135; 19: DIM; 21: 0.017 |  | 0.5 | 0.017 | Wascator-tav.fw-DIM |  | Y
Wascator-tav.fw |  |  |  | 10: 160; 12: ±5; 14: mm; 15: 160; 16: 5; 17: 155; 18: 165; 19: DIM; 21: 0.017 |  | 0.5 | 0.017 | Wascator-tav.fw-DIM |  | Y
Wascator-tav.fw |  |  |  | 10: 200; 12: ±5; 14: mm; 15: 200; 16: 5; 17: 195; 18: 205; 19: DIM; 21: 0.017 |  | 0.5 | 0.017 | Wascator-tav.fw-DIM
Wascator-tav.fw | [7 empty] | Agitation (*) - rpm | 52 | ±2 | Agitation (*) - rpm | Tốc độ giặt | rpm | 52 | 2 | 50 | 54 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wascator-tav.fw-RPM
Wascator-tav.fw | [7 empty] | Time - sec | 100 | N/A | Time - sec | Thời gian | sec | 100 |  | 100 | 100 | TIME |  | 0.02 |  |  | #N/A | 0.02 | Wascator-tav.fw-TIME |  | Y
Wascator-tav.fw | [7 empty] | Spin Speed (*) - rpm | 400 | ±50 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 400 | 50 | 350 | 450 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wascator-tav.fw-RPM
Wascator-tav.fw |  |  |  | 10: 500; 12: ±50; 14: rpm; 15: 500; 16: 50; 17: 450; 18: 550; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Wascator-tav.fw-RPM
Wascator-tav.fw |  |  |  | 10: 600; 12: ±50; 14: rpm; 15: 600; 16: 50; 17: 550; 18: 650; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Wascator-tav.fw-RPM |  | Y
Wascator-tav.fw |  |  |  | 10: 700; 12: ±50; 14: rpm; 15: 700; 16: 50; 17: 650; 18: 750; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Wascator-tav.fw-RPM
Wascator-tav.fw |  |  |  | 10: 800; 12: ±50; 14: rpm; 15: 800; 16: 50; 17: 750; 18: 850; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Wascator-tav.fw-RPM
Wascator-tav.fw |  |  |  | 10: 900; 12: ±50; 14: rpm; 15: 900; 16: 50; 17: 850; 18: 950; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Wascator-tav.fw-RPM |  | Y
Wascator-tav.fw |  |  |  | 10: 1000; 12: ±50; 14: rpm; 15: 1000; 16: 50; 17: 950; 18: 1050; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Wascator-tav.fw-RPM |  | Y
Wascator-tav.fw |  |  |  | 10: 1100; 12: ±50; 14: rpm; 15: 1100; 16: 50; 17: 1050; 18: 1150; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Wascator-tav.fw-RPM |  | Y
Wascator-tav.fw | [7 empty] | Wash Temp (*) - °C | 25 | ±3 | Wash Temp (*) - °C | Nhiệt độ | ℃ | 25 | 3 | 22 | 28 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wascator-tav.fw-DEG |  | Y
Wascator-tav.fw |  |  |  | 10: 30; 12: ±3; 14: ℃; 15: 30; 16: 3; 17: 27; 18: 33; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wascator-tav.fw-DEG |  | Y
Wascator-tav.fw |  |  |  | 10: 40; 12: ±3; 14: ℃; 15: 40; 16: 3; 17: 37; 18: 43; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wascator-tav.fw-DEG |  | Y
Wascator-tav.fw |  |  |  | 10: 50; 12: ±3; 14: ℃; 15: 50; 16: 3; 17: 47; 18: 53; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wascator-tav.fw-DEG |  | Y
Wascator-tav.fw |  |  |  | 10: 60; 12: ±3; 14: ℃; 15: 60; 16: 3; 17: 57; 18: 63; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wascator-tav.fw-DEG |  | Y
Wascator-tav.fw |  |  |  | 10: 70; 12: ±3; 14: ℃; 15: 70; 16: 3; 17: 67; 18: 73; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wascator-tav.fw-DEG |  | Y
Wascator-tav.fw |  |  |  | 10: 80; 12: ±3; 14: ℃; 15: 80; 16: 3; 17: 77; 18: 83; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wascator-tav.fw-DEG |  | Y
Wascator-tav.fw |  |  |  | 10: 90; 12: ±3; 14: ℃; 15: 90; 16: 3; 17: 87; 18: 93; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wascator-tav.fw-DEG |  | Y
Wascator-tav.fw |  |  |  | 9: Others; 11: *
Wascator-tav.fw | bar: 9; 11
Wascator-tav.fw | bar: 9; 11
Wascator-tav.fw | bar: 9; 11
Wascator-tav.fw | [27 empty] | * | bar: 9; 11
Wascator-tav.fw | [27 empty] | * | bar: 9; 11
Wash-ISO-front | ISO 6330:2012 - Table B1 |  | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | 5: Wascator FOM71CLS; 6: Washer; 7: Máy giặt; 9: Program 1: ….; 11: *; 21: #N/A; 22: 3000000 | Service & Calibration | #N/A | #N/A | Wash-ISO-front- |  | Y | *
Wash-ISO-front |  |  |  | 9: Wash; 11: Wash; 21: #N/A |  | #N/A | #N/A | Wash-ISO-front- |  | Y | Wash
Wash-ISO-front | [7 empty] | Temp (*) - °C - 3M, 3N | 30 | ±3 | Temp (*) - °C | Nhiệt độ | ℃ | 30 | 3 | 27 | 33 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-ISO-front-DEG |  |  | 30 - Temp (*) - °C
Wash-ISO-front |  |  |  | 9: Temp (*) - °C - 4M, 4N; 10: 40; 12: ±3; 14: ℃; 15: 40; 16: 3; 17: 37; 18: 43; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-front-DEG |  | Y | 40
Wash-ISO-front |  |  |  | 9: Temp (*) - °C - 5M, 5N; 10: 50; 12: ±3; 14: ℃; 15: 50; 16: 3; 17: 47; 18: 53; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-front-DEG
Wash-ISO-front |  |  |  | 9: Temp (*) - °C - 6M, 6N; 10: 60; 12: ±3; 14: ℃; 15: 60; 16: 3; 17: 57; 18: 63; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-front-DEG |  | Y | 60
Wash-ISO-front |  |  |  | 9: Temp (*) - °C - 7N; 10: 70; 12: ±3; 14: ℃; 15: 70; 16: 3; 17: 67; 18: 73; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-front-DEG |  | Y
Wash-ISO-front |  |  |  | 9: Temp (*) - °C - 9N; 10: 92; 12: ±3; 14: ℃; 15: 92; 16: 3; 17: 89; 18: 95; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-front-DEG |  | Y | 92
Wash-ISO-front | [7 empty] | Water level -mm - 3M, 3N, 4H, 4M, 4N,5M, 5N, 6M, 6N, 7N, 9N | 100 | ±5 | Water level -mm | Mực nước | mm | 100 | 5 | 95 | 105 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Wash-ISO-front-DIM |  | Y | 100 - Water level -mm
Wash-ISO-front |  |  |  | 9: Water level -mm - 4G; 10: 130; 12: ±5; 14: mm; 15: 130; 16: 5; 17: 125; 18: 135; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wash-ISO-front-DIM |  | Y | 130
Wash-ISO-front | [7 empty] | Wash time - sec - 3G, 3M, 3N, 4M, 4N, 5M, 5N, 6M, 6N, 7N, 9N | 900 | N/A | Wash time - sec | Thời gian giặt | sec | 900 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-front-TIME |  |  | 900 - Wash time - sec
Wash-ISO-front |  |  |  | 9: Wash time - sec - 4G; 10: 180; 12: N/A; 14: sec; 15: 180; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-front-TIME
Wash-ISO-front |  |  |  | 9: Rinse 1; 11: Rinse 1; 21: #N/A |  | #N/A | #N/A | Wash-ISO-front- |  | Y | Rinse 1
Wash-ISO-front | [7 empty] | Water level -mm - 3M, 3N, 4G, 4M, 4N, 5M, 5N, 6M, 6N, 7N, 9N | 130 | ±5 | Water level -mm | Mực nước | mm | 130 | 5 | 125 | 135 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Wash-ISO-front-DIM |  |  | 130 - Water level -mm
Wash-ISO-front | [7 empty] | Rinse time - sec - 3M, 3N, 4G, 4M, 4N, 5M, 5N, 6M, 6N, 7N, 9N | 180 | N/A | Rinse time - sec | Thời gian xả | sec | 180 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-front-TIME |  |  | 180 - Rinse time - sec
Wash-ISO-front |  |  |  | 9: Rinse 2; 11: Rinse 2; 21: #N/A |  | #N/A | #N/A | Wash-ISO-front- |  | Y
Wash-ISO-front | [7 empty] | Water level -mm - 3M, 3N, 4G, 4M, 4N,5M, 5N, 6M, 6N, 7N, 9N | 130 | ±5 | Water level -mm | Mực nước | mm | 130 | 5 | 125 | 135 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Wash-ISO-front-DIM |  | Y | 130 - Water level -mm
Wash-ISO-front | [7 empty] | Rinse time - sec - 3N, 4G, 4N, 5N, 6N, 7N, 9N | 180 | N/A | Rinse time - sec | Thời gian xả | sec | 180 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-front-TIME |  | Y
Wash-ISO-front |  |  |  | 9: Rinse time - sec - 3M, 4M, 5M, 6M; 10: 120; 12: N/A; 14: sec; 15: 120; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-front-TIME |  | Y | 120
Wash-ISO-front | [7 empty] | Spin time - sec - 4G | 60 | N/A | Spin time - sec | Thời gian vắt | sec | 60 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-front-TIME |  | Y | 60 - Spin time - sec
Wash-ISO-front
Wash-ISO-front |  |  |  | 9: Rinse 3; 11: Rinse 3; 21: #N/A |  | #N/A | #N/A | Wash-ISO-front- |  | Y
Wash-ISO-front | [7 empty] | Water level -mm - 3M, 3N, 4G, 4M, 4N, 5M, 5N, 6M, 6N, 7N, 9N | 130 | ±5 | Water level -mm | Mực nước | mm | 130 | 5 | 125 | 135 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Wash-ISO-front-DIM |  | Y
Wash-ISO-front |  |  |  | 9: Rinse time - sec - 3M, 3N, 4G, 4M, 4N, 5M, 5N, 6M, 6N, 7N, 9N; 10: 120; 12: N/A; 11: Rinse time - sec; 13: Thời gian xả; 14: sec; 15: 120; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-front-TIME
Wash-ISO-front |  |  |  | 9: Spin time - sec - 3M, 4M, 5M, 6M; 10: 120; 12: N/A; 11: Spin time - sec; 13: Thời gian vắt; 14: sec; 15: 120; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-front-TIME
Wash-ISO-front |  |  |  | 9: Spin time - sec - 4G; 10: 360; 12: N/A; 14: sec; 15: 360; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-front-TIME
Wash-ISO-front |  |  |  | 9: Rinse 4; 11: Rinse 4
Wash-ISO-front |  |  |  | 9: Water level -mm - 3N, 4N, 5N, 6N, 7N, 9N; 10: 130; 12: ±5; 11: Water level -mm; 13: Mực nước; 14: mm; 15: 130; 16: 5; 19: DIM
Wash-ISO-front |  |  |  | 9: Rinse time - sec - 3N, 4N, 5N, 6N, 7N, 9N; 10: 120; 12: N/A; 11: Rinse time - sec; 13: Thời gian xả; 14: sec; 15: 120; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-front-TIME
Wash-ISO-front |  |  |  | 9: Spin time - sec - 3N, 4N, 5N, 6N, 7N, 9N; 10: 300; 12: N/A; 11: Spin time - sec; 13: Thời gian vắt; 14: sec; 15: 300; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-front-TIME
Wash-ISO-front
Wash-ISO-front |  |  |  | 9: Program 2: ….; 11: *
Wash-ISO-front |  |  |  | 9: Wash; 11: Wash
Wash-ISO-front |  |  |  | 9: Temp (*) - °C - 3M, 3N; 10: 30; 12: ±3; 11: Temp (*) - °C; 13: Nhiệt độ; 14: ℃; 15: 30; 16: 3; 19: DEG
Wash-ISO-front |  |  |  | 9: Temp (*) - °C - 4M, 4N; 10: 40; 12: ±3; 14: ℃; 15: 40; 16: 3; 19: DEG
Wash-ISO-front |  |  |  | 9: Temp (*) - °C - 5M, 5N; 10: 50; 12: ±3; 14: ℃; 15: 50; 16: 3; 19: DEG
Wash-ISO-front |  |  |  | 9: Temp (*) - °C - 6M, 6N; 10: 60; 12: ±3; 14: ℃; 15: 60; 16: 3; 19: DEG
Wash-ISO-front |  |  |  | 9: Temp (*) - °C - 7N; 10: 70; 12: ±3; 14: ℃; 15: 70; 16: 3; 19: DEG
Wash-ISO-front |  |  |  | 9: Temp (*) - °C - 9N; 10: 92; 12: ±3; 14: ℃; 15: 92; 16: 3; 19: DEG
Wash-ISO-front |  |  |  | 9: Water level -mm - 3M, 3N, 4H, 4M, 4N,5M, 5N, 6M, 6N, 7N, 9N; 10: 100; 12: N/A; 11: Water level -mm; 13: Mực nước; 14: mm; 15: 100; 16: N/A; 19: DIM
Wash-ISO-front |  |  |  | 9: Water level -mm - 4G; 10: 130; 12: N/A; 14: mm; 15: 130; 16: N/A; 19: DIM
Wash-ISO-front |  |  |  | 9: Wash time - sec - 3M, 3N, 4M, 4N, 5M, 5N, 6M, 6N, 7N, 9N; 10: 900; 12: N/A; 11: Wash time - sec; 13: Thời gian giặt; 14: sec; 15: 900; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-front-TIME
Wash-ISO-front |  |  |  | 9: Wash time - sec - 4G; 10: 180; 12: N/A; 14: sec; 15: 180; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-front-TIME
Wash-ISO-front |  |  |  | 9: Rinse 1; 11: Rinse 1
Wash-ISO-front |  |  |  | 9: Water level -mm - 3M, 3N, 4G, 4M, 4N, 5M, 5N, 6M, 6N, 7N, 9N; 10: 130; 12: ±5; 11: Water level -mm; 13: Mực nước; 14: mm; 15: 130; 16: 5; 19: DIM
Wash-ISO-front |  |  |  | 9: Rinse time - sec - 3M, 3N, 4G, 4M, 4N, 5M, 5N, 6M, 6N, 7N, 9N; 10: 180; 12: N/A; 11: Rinse time - sec; 13: Thời gian xả; 14: sec; 15: 180; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-front-TIME
Wash-ISO-front |  |  |  | 9: Rinse 2; 11: Rinse 2
Wash-ISO-front |  |  |  | 9: Water level -mm - 3M, 3N, 4G, 4M, 4N,5M, 5N, 6M, 6N, 7N, 9N; 10: 130; 12: ±5; 11: Water level -mm; 13: Mực nước; 14: mm; 15: 130; 16: 5; 19: DIM
Wash-ISO-front | [7 empty] | Rinse time - sec - 3N, 4G, 4N, 5N, 6N, 7N, 9N | 180 | N/A | Rinse time - sec | Thời gian xả | sec | 180 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-front-TIME |  | Y
Wash-ISO-front |  |  |  | 9: Rinse time - sec - 3M, 4M, 5M, 6M; 10: 120; 12: N/A; 14: sec; 15: 120; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-front-TIME |  | Y | 120
Wash-ISO-front | [7 empty] | Spin time - sec - 4G | 60 | N/A | Spin time - sec | Thời gian vắt | sec | 60 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-front-TIME |  | Y | 60 - Spin time - sec
Wash-ISO-front |  |  |  | 9: Rinse 3; 11: Rinse 3
Wash-ISO-front |  |  |  | 9: Water level -mm - 3M, 3N, 4G, 4M, 4N, 5M, 5N, 6M, 6N, 7N, 9N; 10: 130; 12: ±5; 11: Water level -mm; 13: Mực nước; 14: mm; 15: 130; 16: 5; 19: DIM
Wash-ISO-front |  |  |  | 9: Rinse time - sec - 3M, 3N, 4G, 4M, 4N, 5M, 5N, 6M, 6N, 7N, 9N; 10: 120; 12: N/A; 11: Rinse time - sec; 13: Thời gian xả; 14: sec; 15: 120; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-front-TIME
Wash-ISO-front |  |  |  | 9: Spin time - sec - 3M, 4M, 5M, 6M; 10: 120; 12: N/A; 11: Spin time - sec; 13: Thời gian vắt; 14: sec; 15: 120; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-front-TIME
Wash-ISO-front |  |  |  | 9: Spin time - sec - 4G; 10: 360; 12: N/A; 14: sec; 15: 360; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-front-TIME
Wash-ISO-front |  |  |  | 9: Rinse 4; 11: Rinse 4
Wash-ISO-front |  |  |  | 9: Water level -mm; 10: 130; 12: ±5; 11: Water level -mm; 13: Mực nước; 14: mm; 15: 130; 16: 5; 19: DIM
Wash-ISO-front |  |  |  | 9: Rinse time - sec - 3N, 4N, 5N, 6N, 7N, 9N; 10: 120; 12: N/A; 11: Rinse time - sec; 13: Thời gian xả; 14: sec; 15: 120; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-front-TIME
Wash-ISO-front |  |  |  | 9: Spin time - sec - 3N, 4N, 5N, 6N, 7N, 9N; 10: 300; 12: N/A; 11: Spin time - sec; 13: Thời gian vắt; 14: sec; 15: 300; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-front-TIME
Wash-ISO-front
Wash-ISO-front |  |  |  | 9: Others; 11: *
Wash-ISO-front | bar: 9; 11
Wash-ISO-front | bar: 9; 11
Wash-ISO-front | bar: 9; 11
Wash-ISO-front | bar: 9; 11
Wash-ISO-front | bar: 9; 11
Wash-ISO-typeB | ISO 6330 Type B |  | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | 5: Labtex washer - LBT-M6T; 6: Washer; 7: Máy giặt; 9: Program 1: ….; 11: *; 21: #N/A; 22: 3000000 | Service & Calibration | #N/A | #N/A | Wash-ISO-typeB- |  | Y | *
Wash-ISO-typeB | [7 empty] | Agitation - spm - 1B, 2B, 3B, 4B, 5B, 6B, 7B, 9B, 10B: Normal | 179 | ±2 | Agitation - spm - 1B, 2B, 3B, 4B, 5B, 6B, 7B, 9B, 10B: Normal | Tốc độ lắc | spm | 179 | 2 | 177 | 181 | SPM |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-typeB-SPM |  | Y | 179 - Agitation - spm - 1B, 2B, 3B, 4B, 5B, 6B, 7B, 9B, 10B: Normal
Wash-ISO-typeB | [7 empty] | Agitation – spm – 8B, 11B: Gentle | 119 | ±2 | Agitation – spm – 8B, 11B: Gentle | Tốc độ lắc | spm | 119 | 2 | 117 | 121 | SPM |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-typeB-SPM |  | Y | 119 - Agitation – spm – 8B, 11B: Gentle
Wash-ISO-typeB |  |  |  | 9: Wash; 11: Wash; 21: #N/A |  | #N/A | #N/A | Wash-ISO-typeB- |  | Y | Wash
Wash-ISO-typeB |  |  |  | 9: Temp (*) - °C - 1B; 10: 60; 12: ±3; 14: ℃; 15: 60; 16: 3; 17: 57; 18: 63; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-typeB-DEG |  | Y | 60
Wash-ISO-typeB |  |  |  | 9: Temp (*) - °C - 2B, 3B; 10: 49; 12: ±3; 14: ℃; 15: 49; 16: 3; 17: 46; 18: 52; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-typeB-DEG
Wash-ISO-typeB |  |  |  | 9: Temp (*) - °C - 4B, 5B; 10: 41; 12: ±3; 14: ℃; 15: 41; 16: 3; 17: 38; 18: 44; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-typeB-DEG |  | Y | 41
Wash-ISO-typeB |  |  |  | 9: Temp (*) - °C - 6B,7B, 8B; 10: 27; 12: ±3; 14: ℃; 15: 27; 16: 3; 17: 24; 18: 30; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-typeB-DEG |  | Y
Wash-ISO-typeB |  |  |  | 9: Temp (*) - °C - 9B, 10B, 11B; 10: 16; 12: ±3; 14: ℃; 15: 16; 16: 3; 17: 13; 18: 19; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-typeB-DEG |  | Y | 16
Wash-ISO-typeB | [7 empty] | Water level -mm - 1B, 2B, 3B, 4B, 5B, 6B, 7B, 8B, 9B, 10B | 297 | ±25 | Water level -mm | Mực nước | mm | 297 | 25 | 272 | 322 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Wash-ISO-typeB-DIM |  | Y | 297 - Water level -mm
Wash-ISO-typeB |  |  |  | 9: Water level -mm - 11B; 10: 398.5; 12: ±17.8; 14: mm; 15: 398.5; 16: 17.8; 17: 380.7; 18: 416.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wash-ISO-typeB-DIM |  | Y | 398.5
Wash-ISO-typeB | [7 empty] | Wash time - sec -1B, 2B, 4B, 6B, 9B | 720 | N/A | Wash time - sec | Thời gian giặt | sec | 720 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-typeB-TIME |  |  | 720 - Wash time - sec
Wash-ISO-typeB |  |  |  | 9: Wash time - sec - 3B, 5B, 7B, 10B; 10: 600; 12: N/A; 14: sec; 15: 600; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-typeB-TIME
Wash-ISO-typeB |  |  |  | 9: Wash time - sec - 8B, 11B; 10: 480; 12: N/A; 14: sec; 15: 480; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-typeB-TIME
Wash-ISO-typeB |  |  |  | 9: Rinse; 11: Rinse 1; 21: #N/A |  | #N/A | #N/A | Wash-ISO-typeB- |  | Y | Rinse 1
Wash-ISO-typeB | [7 empty] | Water level -mm - 1B, 2B, 3B, 4B, 5B, 6B, 7B, 8B, 9B, 10B, 11B | 297 | ±25 | Water level -mm | Mực nước | mm | 297 | 25 | 272 | 322 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Wash-ISO-typeB-DIM |  |  | 297 - Water level -mm
Wash-ISO-typeB | [7 empty] | Rinse time - sec - 1B, 2B, 3B, 4B, 5B, 6B, 7B, 8B, 9B, 10B, 11B | 180 | N/A | Rinse time - sec | Thời gian xả | sec | 180 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-typeB-TIME |  |  | 180 - Rinse time - sec
Wash-ISO-typeB |  |  |  | 9: Spin; 11: Spin; 21: #N/A |  | #N/A | #N/A | Wash-ISO-typeB- |  | Y
Wash-ISO-typeB | [7 empty] | Spin Speed (*) - rpm - 1B, 2B, 4B, 6B, 9B | 613 to 640 | 613 to 640 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 613 | 613 to 640 | 613 | 640 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-ISO-typeB-RPM
Wash-ISO-typeB | [7 empty] | Spin Speed (*) - rpm - 3B, 5B, 7B, 8B, 10B, 11B | 399 to 420 | 399 to 420 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 399 | 399 to 420 | 399 | 420 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-ISO-typeB-RPM |  |  | Cotton - 30℃ - 30℃
Wash-ISO-typeB | [7 empty] | Spin time - sec - 1B, 2B, 4B, 6B, 9B | 360 | N/A | Spin time - sec | Thời gian vắt | sec | 360 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-typeB-TIME |  | Y | 360 - Spin time - sec
Wash-ISO-typeB | [7 empty] | Spin time - sec - 3B, 5B, 7B, 8B, 10B, 11B | 240 | N/A | Spin time - sec | Thời gian vắt | sec | 240 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-typeB-TIME |  | Y | 240 - Spin time - sec
Wash-ISO-typeB
Wash-ISO-typeB |  |  |  | 9: Program 2: ….; 11: *
Wash-ISO-typeB | [7 empty] | Agitation - spm - 1B, 2B, 3B, 4B, 5B, 6B, 7B, 9B, 10B: Normal | 179 | ±2 | Agitation - spm - 1B, 2B, 3B, 4B, 5B, 6B, 7B, 9B, 10B: Normal | Tốc độ lắc | spm | 179 | 2 | 177 | 181 | SPM |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-typeB-SPM |  | Y | 179 - Agitation - spm - 1B, 2B, 3B, 4B, 5B, 6B, 7B, 9B, 10B: Normal
Wash-ISO-typeB | [7 empty] | Agitation – spm – 8B, 11B: Gentle | 119 | ±2 | Agitation – spm – 8B, 11B: Gentle | Tốc độ lắc | spm | 119 | 2 | 117 | 121 | SPM |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-typeB-SPM |  | Y | 119 - Agitation – spm – 8B, 11B: Gentle
Wash-ISO-typeB |  |  |  | 9: Wash; 11: Wash; 21: #N/A |  | #N/A | #N/A | Wash-ISO-typeB- |  | Y | Wash
Wash-ISO-typeB |  |  |  | 9: Temp (*) - °C - 1B; 10: 60; 12: ±3; 14: ℃; 15: 60; 16: 3; 17: 57; 18: 63; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-typeB-DEG |  | Y | 60
Wash-ISO-typeB |  |  |  | 9: Temp (*) - °C - 2B, 3B; 10: 49; 12: ±3; 14: ℃; 15: 49; 16: 3; 17: 46; 18: 52; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-typeB-DEG
Wash-ISO-typeB |  |  |  | 9: Temp (*) - °C - 4B, 5B; 10: 41; 12: ±3; 14: ℃; 15: 41; 16: 3; 17: 38; 18: 44; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-typeB-DEG |  | Y | 41
Wash-ISO-typeB |  |  |  | 9: Temp (*) - °C - 6B,7B, 8B; 10: 27; 12: ±3; 14: ℃; 15: 27; 16: 3; 17: 24; 18: 30; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-typeB-DEG |  | Y
Wash-ISO-typeB |  |  |  | 9: Temp (*) - °C - 9B, 10B, 11B; 10: 16; 12: ±3; 14: ℃; 15: 16; 16: 3; 17: 13; 18: 19; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-ISO-typeB-DEG |  | Y | 16
Wash-ISO-typeB | [7 empty] | Water level -mm - 1B, 2B, 3B, 4B, 5B, 6B, 7B, 8B, 9B, 10B | 297 | ±25 | Water level -mm | Mực nước | mm | 297 | 25 | 272 | 322 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Wash-ISO-typeB-DIM |  | Y | 297 - Water level -mm
Wash-ISO-typeB |  |  |  | 9: Water level -mm - 11B; 10: 398.5; 12: ±17.8; 14: mm; 15: 398.5; 16: 17.8; 17: 380.7; 18: 416.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wash-ISO-typeB-DIM |  | Y | 398.5
Wash-ISO-typeB | [7 empty] | Wash time - sec -1B, 2B, 4B, 6B, 9B | 720 | N/A | Wash time - sec | Thời gian giặt | sec | 720 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-typeB-TIME |  |  | 720 - Wash time - sec
Wash-ISO-typeB |  |  |  | 9: Wash time - sec - 3B, 5B, 7B, 10B; 10: 600; 12: N/A; 14: sec; 15: 600; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-typeB-TIME
Wash-ISO-typeB |  |  |  | 9: Wash time - sec - 8B, 11B; 10: 480; 12: N/A; 14: sec; 15: 480; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-ISO-typeB-TIME
Wash-ISO-typeB |  |  |  | 9: Rinse; 11: Rinse 1; 21: #N/A |  | #N/A | #N/A | Wash-ISO-typeB- |  | Y | Rinse 1
Wash-ISO-typeB | [7 empty] | Water level -mm - 1B, 2B, 3B, 4B, 5B, 6B, 7B, 8B, 9B, 10B, 11B | 297 | ±25 | Water level -mm | Mực nước | mm | 297 | 25 | 272 | 322 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Wash-ISO-typeB-DIM |  |  | 297 - Water level -mm
Wash-ISO-typeB | [7 empty] | Rinse time - sec - 1B, 2B, 3B, 4B, 5B, 6B, 7B, 8B, 9B, 10B, 11B | 180 | N/A | Rinse time - sec | Thời gian xả | sec | 180 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-typeB-TIME |  |  | 180 - Rinse time - sec
Wash-ISO-typeB |  |  |  | 9: Spin; 11: Spin; 21: #N/A |  | #N/A | #N/A | Wash-ISO-typeB- |  | Y
Wash-ISO-typeB | [7 empty] | Spin Speed (*) - rpm - 1B, 2B, 4B, 6B, 9B | 613 to 640 | 613 to 640 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 613 | 613 to 640 | 613 | 640 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-ISO-typeB-RPM
Wash-ISO-typeB | [7 empty] | Spin Speed (*) - rpm - 3B, 5B, 7B, 8B, 10B, 11B | 399 to 420 | 399 to 420 | Spin Speed (*) - rpm | Tốc độ vắt | rpm | 399 | 399 to 420 | 399 | 420 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-ISO-typeB-RPM |  |  | Cotton - 30℃ - 30℃
Wash-ISO-typeB | [7 empty] | Spin time - sec - 1B, 2B, 4B, 6B, 9B | 360 | N/A | Spin time - sec | Thời gian vắt | sec | 360 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-typeB-TIME |  | Y | 360 - Spin time - sec
Wash-ISO-typeB | [7 empty] | Spin time - sec - 3B, 5B, 7B, 8B, 10B, 11B | 240 | N/A | Spin time - sec | Thời gian vắt | sec | 240 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-ISO-typeB-TIME |  | Y | 240 - Spin time - sec
Wash-ISO-typeB
Wash-ISO-typeB |  |  |  | 9: Others; 11: *
Wash-ISO-typeB | bar: 9; 11
Wash-ISO-typeB | bar: 9; 11
Wash-ISO-typeB | bar: 9; 11
Wash-ISO-typeB | bar: 9; 11
Wash-ISO-typeB | bar: 9; 11
Wash-HM | H&M | Washing - H&M | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | 5: Electrolux W555H; 6: Wash; 8: Water level Temperature Washing time Agitation speed (spm) Spin speed (rpm) Spin time; 9: Normal 60 °C; 11: Normal 60 °C; 21: #N/A; 22: 1800000 | Service & Calibration | #N/A | #N/A | Wash-HM- |  | Y
Wash-HM | [7 empty] | Total time - sec | 2820 | N/A | Total time - sec | Tổng thời gian | sec | 2820 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM-TIME |  | Y
Wash-HM | [7 empty] | Temp (*) - °C | 60 | ±3 | Temp (*) - °C | Nhiệt độ | ℃ | 60 | 3 | 57 | 63 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-HM-DEG |  | Y
Wash-HM |  |  |  | 9: Water Consumpsion - L; 10: 65; 11: Water Consumpsion - L; 13: Mức nước tiêu thụ; 14: L; 15: 65; 17: 65; 18: 65; 19: VOL; 21: 0.07 |  | 0,01 | 0.07 | Wash-HM-VOL |  | Y
Wash-HM | [7 empty] | Spin Speed (*) - max - rpm | 1300 | N/A | Spin Speed (*) - max - rpm | Tốc độ vắt | rpm | 1300 | N/A | #VALUE! | #VALUE! | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-HM-RPM |  | Y
Wash-HM | H&M | Washing - H&M | [5 empty] | Spin time - sec | 360 | N/A | Spin time - sec | Thời gian vắt | sec | 360 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM-TIME |  | Y
Wash-HM
Wash-HM |  |  |  | 9: Normal 40 °C; 11: Normal 40 °C; 21: #N/A |  | #N/A | #N/A | Wash-HM- |  | Y
Wash-HM | [7 empty] | Total time - sec | 2520 | N/A | Total time - sec | Tổng thời gian | sec | 2520 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM-TIME
Wash-HM | [7 empty] | Temp (*) - °C | 40 | ±3 | Temp (*) - °C | Nhiệt độ | ℃ | 40 | 3 | 37 | 43 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-HM-DEG |  | Y
Wash-HM |  |  |  | 9: Water Consumpsion - L; 10: 65; 11: Water Consumpsion - L; 13: Mức nước tiêu thụ; 14: L; 15: 65; 17: 65; 18: 65; 19: VOL; 21: 0.07 |  | 0,01 | 0.07 | Wash-HM-VOL |  | Y
Wash-HM | H&M | Washing - H&M | [5 empty] | Spin Speed (*) - max - rpm | 1300 | N/A | Spin Speed (*) - max - rpm | Tốc độ vắt | rpm | 1300 | N/A | #VALUE! | #VALUE! | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-HM-RPM |  | Y
Wash-HM | [7 empty] | Spin time - sec | 360 | N/A | Spin time - sec | Thời gian vắt | sec | 360 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM-TIME |  | Y
Wash-HM
Wash-HM |  |  |  | 9: Synthetic 40 °C; 11: Synthetic 40 °C; 21: #N/A |  | #N/A | #N/A | Wash-HM- |  | Y
Wash-HM | [7 empty] | Total time - sec | 1740 | N/A | Total time - sec | Tổng thời gian | sec | 1740 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM-TIME |  | Y
Wash-HM | [7 empty] | Temp (*) - °C | 40 | ±3 | Temp (*) - °C | Nhiệt độ | ℃ | 40 | 3 | 37 | 43 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-HM-DEG |  | Y
Wash-HM | H&M | Washing - H&M |  | 9: Water Consumpsion - L; 10: 26; 11: Water Consumpsion - L; 13: Mức nước tiêu thụ; 14: L; 15: 26; 17: 26; 18: 26; 19: VOL; 21: 0.07 |  | 0,01 | 0.07 | Wash-HM-VOL |  | Y
Wash-HM | [7 empty] | Spin Speed (*) - max - rpm | 1100 | N/A | Spin Speed (*) - max - rpm | Tốc độ vắt | rpm | 1100 | N/A | #VALUE! | #VALUE! | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-HM-RPM
Wash-HM | [7 empty] | Spin time - sec | 90 | N/A | Spin time - sec | Thời gian vắt | sec | 90 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM-TIME
Wash-HM
Wash-HM |  |  |  | 9: Synthetic 30 °C; 11: Synthetic 30 °C; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-HM-
Wash-HM | [7 empty] | Total time - sec | 1680 | N/A | Total time - sec | Tổng thời gian | sec | 1680 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM-TIME
Wash-HM | [7 empty] | Temp (*) - °C | 30 | ±3 | Temp (*) - °C | Nhiệt độ | ℃ | 30 | 3 | 27 | 33 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-HM-DEG
Wash-HM |  |  |  | 9: Water Consumpsion - L; 10: 42; 11: Water Consumpsion - L; 13: Mức nước tiêu thụ; 14: L; 15: 42; 17: 42; 18: 42; 19: VOL; 21: 0.07 |  | 0,01 | 0.07 | Wash-HM-VOL
Wash-HM | [7 empty] | Spin Speed (*) - max - rpm | 850 | N/A | Spin Speed (*) - max - rpm | Tốc độ vắt | rpm | 850 | N/A | #VALUE! | #VALUE! | RPM |  |  |  |  | 0.1 | 0.1 | Wash-HM-RPM
Wash-HM | [7 empty] | Spin time - sec | 90 | N/A | Spin time - sec | Thời gian vắt | sec | 90 | N/A | #VALUE! | #VALUE! | TIME |  |  |  |  | 0.001 | 0.02 | Wash-HM-TIME
Wash-HM
Wash-HM |  |  |  | 9: Wool 40 °C; 11: Wool 40 °C |  | #N/A | #N/A | Wash-HM-
Wash-HM | [7 empty] | Total time - sec | 1560 | N/A | Total time - sec | Tổng thời gian | sec | 1560 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM-TIME
Wash-HM | [7 empty] | Temp (*) - °C | 40 | ±3 | Temp (*) - °C | Nhiệt độ | ℃ | 40 | 3 | 37 | 43 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-HM-DEG
Wash-HM |  |  |  | 9: Water Consumpsion - L; 10: 33; 11: Water Consumpsion - L; 13: Mức nước tiêu thụ; 14: L; 15: 33; 17: 33; 18: 33; 19: VOL; 21: 0.07 |  | 0,01 | 0.07 | Wash-HM-VOL
Wash-HM | [7 empty] | Spin Speed (*) - max - rpm | 550 | N/A | Spin Speed (*) - max - rpm | Tốc độ vắt | rpm | 550 | N/A | #VALUE! | #VALUE! | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-HM-RPM
Wash-HM | [7 empty] | Spin time - sec | 90 | N/A | Spin time - sec | Thời gian vắt | sec | 90 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM-TIME
Wash-HM
Wash-HM |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-HM-
Wash-HM |  |  |  | 9: 1…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-HM-
Wash-HM |  |  |  | 9: 2…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-HM-
Wash-HM |  |  |  | 9: 3…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-HM-
Wash-HM |  |  |  | 9: 4…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-HM- |  |  | *
Wash-HM |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-HM- |  |  | *
Wash-HM71 | H&M | Washing - H&M | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | 5: Wascator FOM71 CLS; 6: Wash; 8: Water level Temperature Washing time Agitation speed (spm) Spin speed (rpm) Spin time; 9: Normal 40 (Cotton 40); 11: Normal 40 (Cotton 40); 21: #N/A; 22: 1800000 | Service & Calibration | #N/A | #N/A | Wash-HM71- |  | Y
Wash-HM71 | [7 empty] | Wash Temp (*) - °C | 40 | ±3 | Wash Temp (*) - °C | Nhiệt độ | ℃ | 40 | 3 | 37 | 43 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-HM71-DEG |  | Y
Wash-HM71 | [7 empty] | Water level - mm | 100 | ±5 | Water level - mm | Mức nước | mm | 100 | 5 | 95 | 105 | VOL |  | 0.07 |  |  | 0,01 | 0.07 | Wash-HM71-VOL |  | Y
Wash-HM71 | [7 empty] | Wash time - sec | 900 | N/A | Wash time - sec | Thời gian | sec | 900 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM71-TIME
Wash-HM71 |  | Washing - H&M | [5 empty] | Water level - mm | 130 | ±5 | Water level - mm | Mức nước | mm | 130 | 5 | 125 | 135 | VOL |  | 0.07 |  |  | 0,01 | 0.07 | Wash-HM71-VOL |  | Y
Wash-HM71 | [7 empty] | Rinse time 1 | 180 | N/A | Rinse time 1 | Thời gian xả 1 | sec | 180 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM71-TIME |  | Y
Wash-HM71 | [7 empty] | Water level - mm | 130 | ±5 | Water level - mm | Mức nước | mm | 130 | 5 | 125 | 135 | VOL
Wash-HM71 |  |  |  | 9: Rinse time 2; 10: 180; 12: N/A; 11: Rinse time 2; 13: Thời gian xả 2; 14: sec; 15: 180; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-HM71-TIME |  | Y
Wash-HM71 | [7 empty] | Water level - mm | 130 | ±5 | Water level - mm | Mức nước | mm | 130 | 5 | 125 | 135 | VOL |  | 0.07 |  |  | 0,01 | 0.07 | Wash-HM71-VOL |  | Y
Wash-HM71 | [7 empty] | Rinse time 3 | 120 | N/A | Rinse time 3 | Thời gian xả 3 | sec | 120 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM71-TIME |  | Y
Wash-HM71 |  | Washing - H&M | [5 empty] | Water level - mm | 130 | ±5 | Water level - mm | Mức nước | mm | 130 | 5 | 125 | 135 | VOL |  | 0.07 |  |  | 0,01 | 0.07 | Wash-HM71-VOL |  | Y
Wash-HM71 | [7 empty] | Rinse time 4 | 120 | N/A | Rinse time 4 | Thời gian xả 4 | sec | 120 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM71-TIME
Wash-HM71 | [7 empty] | Spin time - sec | 360 | N/A | Spin time - sec | Thời gian vắt | sec | 360 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM71-TIME
Wash-HM71
Wash-HM71 |  |  |  | 9: Wool 40; 11: Wool 40 |  | #N/A | #N/A | Wash-HM71-
Wash-HM71 | [7 empty] | Wash Temp (*) - °C | 40 | ±3 | Wash Temp (*) - °C | Nhiệt độ | sec | 40 | 3 | 37 | 43 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-HM71-DEG
Wash-HM71 | H&M | Washing - H&M | [5 empty] | Water level - mm | 130 | ±5 | Water level - mm | Mức nước | mm | 130 | 5 | 125 | 135 | VOL |  | 0.07 |  |  | 0,01 | 0.07 | Wash-HM71-VOL |  | Y
Wash-HM71 | [7 empty] | Rinse time 1 | 180 | N/A | Rinse time 1 | Thời gian xả 1 | sec | 180 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM71-TIME |  | Y
Wash-HM71 | [7 empty] | Water level - mm | 130 | ±5 | Water level - mm | Mức nước | mm | 130 | 5 | 125 | 135 | VOL
Wash-HM71 |  |  |  | 9: Rinse time 2; 10: 180; 12: N/A; 11: Rinse time 2; 13: Thời gian xả 2; 14: sec; 15: 180; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-HM71-TIME |  | Y
Wash-HM71 | [7 empty] | Water level - mm | 130 | ±5 | Water level - mm | Mức nước | mm | 130 | 5 | 125 | 135 | VOL |  | 0.07 |  |  | 0,01 | 0.07 | Wash-HM71-VOL |  | Y
Wash-HM71 | [7 empty] | Rinse time 3 | 120 | N/A | Rinse time 3 | Thời gian xả 3 | sec | 120 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM71-TIME |  | Y
Wash-HM71 | [7 empty] | Spin time 1 - sec | 90 | N/A | Spin time 1 - sec | Thời gian vắt | sec | 90 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM71-TIME
Wash-HM71 | [7 empty] | Spin time 2 - sec | 90 | N/A | Spin time 2 - sec | Thời gian vắt | sec | 90 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-HM71-TIME
Wash-HM71
Wash-HM71 |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-HM71-
Wash-HM71 |  |  |  | 9: 1…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-HM71-
Wash-HM71 |  |  |  | 9: 2…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-HM71-
Wash-HM71 |  |  |  | 9: 3…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-HM71-
Wash-HM71 |  |  |  | 9: 4…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-HM71- |  |  | *
Wash-HM71 |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-HM71- |  |  | *
Wash-colorfastness | AATCC 132, Adidas 5.01, ISO 105-D01 | Colourfastness to Washing | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | Gyrowash, Launder Ometer, Rotawash | Wash - Colourfastness | Thiết bị thử bền màu giặt | Wash Temp Time Rotating speed Temperature display Washing temperature Time Speed (rpm) Vessel Volume | Wash Temp (*) | 30 | ±2 | Wash Temp (*) - ℃ | Nhiệt độ giặt | ℃ | 30 | 2 | 28 | 32 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-colorfastness-DEG |  |  | 30 - Wash Temp (*) - ℃
Wash-colorfastness |  |  |  | 10: 40; 12: ±2; 14: ℃; 15: 40; 16: 2; 17: 38; 18: 42; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-colorfastness-DEG |  |  | 40
Wash-colorfastness |  |  |  | 10: 49; 12: ±2; 14: ℃; 15: 49; 16: 2; 17: 47; 18: 51; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-colorfastness-DEG |  |  | 49
Wash-colorfastness |  |  |  | 10: 60; 12: ±2; 14: ℃; 15: 60; 16: 2; 17: 58; 18: 62; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-colorfastness-DEG |  |  | 60
Wash-colorfastness |  |  |  | 10: 70; 12: ±2; 14: ℃; 15: 70; 16: 2; 17: 68; 18: 72; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-colorfastness-DEG |  |  | 70
Wash-colorfastness |  |  |  | 10: 80; 12: ±2; 14: ℃; 15: 80; 16: 2; 17: 78; 18: 82; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-colorfastness-DEG |  |  | 80
Wash-colorfastness |  |  |  | 10: 90; 12: ±2; 14: ℃; 15: 90; 16: 2; 17: 88; 18: 92; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-colorfastness-DEG |  |  | 90
Wash-colorfastness |  |  |  | 9: Temp deviation - ⁰C/min; 10: 1.5; 12: N/A; 11: Temp deviation - ⁰C/min; 13: Độ lệch nhiệt độ; 14: M1; 15: 1.5; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-colorfastness-DEG |  |  | 1.5 - Temp deviation - ⁰C/min
Wash-colorfastness |  |  |  | 9: Time - sec (customer to choose 1800 or 2700 secs); 10: 1800; 12: N/A; 11: Time - sec; 13: Thời gian; 14: sec; 15: 1800; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-colorfastness-TIME |  |  | 1800 - Time - sec
Wash-colorfastness |  |  |  | 9: Time; 10: 2700; 12: N/A; 11: Time - sec; 13: Thời gian; 14: sec; 15: 2700; 16: N/A; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wash-colorfastness-TIME |  |  | 2700 - Time - sec
Wash-colorfastness | [7 empty] | Speed (*) | 40 | ±2 | Speed (*) - rpm | Tốc độ | rpm | 40 | 2 | 38 | 42 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-colorfastness-RPM |  |  | 40 - Speed (*) - rpm
Wash-colorfastness |  |  |  | 10: 60; 12: ±2; 14: rpm; 15: 60; 16: 2; 17: 58; 18: 62; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Wash-colorfastness-RPM |  |  | 60
Wash-colorfastness | [7 empty] | Vessel Volume - ml | Vessel 1 - 550 | ±50 | Vessel Volume - ml | Thể tích cốc | g | 550 | 50 | 500 | 600 | VOL |  | 0.1 |  |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 1 - 550 - Vessel Volume - ml
Wash-colorfastness |  |  |  | 10: Vessel 2 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 2 - 550
Wash-colorfastness |  |  |  | 10: Vessel 3 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 3 - 550
Wash-colorfastness |  |  |  | 10: Vessel 4 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 4 - 550
Wash-colorfastness |  |  |  | 10: Vessel 5 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 5 - 550
Wash-colorfastness |  |  |  | 10: Vessel 6 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 6 - 550
Wash-colorfastness |  |  |  | 10: Vessel 7 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 7 - 550
Wash-colorfastness |  |  |  | 10: Vessel 8 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 8 - 550
Wash-colorfastness |  |  |  | 10: Vessel 9 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 9 - 550
Wash-colorfastness |  |  |  | 10: Vessel 10 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL
Wash-colorfastness |  |  |  | 10: Vessel 11 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 11 - 550
Wash-colorfastness |  |  |  | 10: Vessel 12 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 12 - 550
Wash-colorfastness |  |  |  | 10: Vessel 13 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 13 - 550
Wash-colorfastness |  |  |  | 10: Vessel 14 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 14 - 550
Wash-colorfastness |  |  |  | 10: Vessel 15 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 15 - 550
Wash-colorfastness |  |  |  | 10: Vessel 16 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 16 - 550
Wash-colorfastness |  |  |  | 10: Vessel 17 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 17 - 550
Wash-colorfastness |  |  |  | 10: Vessel 18 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 18 - 550
Wash-colorfastness |  |  |  | 10: Vessel 19 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 19 - 550
Wash-colorfastness |  |  |  | 10: Vessel 20 - 550; 12: ±50; 14: g; 15: 550; 16: 50; 17: 500; 18: 600; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 20 - 550
Wash-colorfastness |  |  |  | 10: Vessel 1 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 1 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 2 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 2 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 3 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 3 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 4 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 4 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 5 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 5 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 6 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 6 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 7 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 7 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 8 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 8 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 9 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 9 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 10 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 10 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 11 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 11 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 12 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 12 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 13 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 13 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 14 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 14 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 15 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 15 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 16 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 16 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 17 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 17 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 18 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 18 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 19 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 19 - 1200
Wash-colorfastness |  |  |  | 10: Vessel 20 - 1200; 12: ±100; 14: g; 15: 1200; 16: 100; 17: 1100; 18: 1300; 19: VOL; 21: 0.1 |  | 0.01 | 0.1 | Wash-colorfastness-VOL |  |  | Vessel 20 - 1200
Wash-colorfastness |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-colorfastness- |  |  | *
Wash-colorfastness |  |  |  | 9: 1…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-colorfastness- |  |  | *
Wash-colorfastness |  |  |  | 9: 2…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-colorfastness- |  |  | *
Wash-colorfastness |  |  |  | 9: 3…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-colorfastness- |  |  | *
Wash-colorfastness |  |  |  | 9: 4…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-colorfastness- |  |  | *
Wash-colorfastness |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-colorfastness- |  |  | *
Wash-durability | M&S C15, M&S P5/P6/P7/P69, Next TM8, BHS TM12A/TM12B/TM12C, ARCADIA CA10a | Garment and printed fabric durability | LAB-T-02 LAB-V-10 LAB-TIME-02 LAB-F01:2023 | Dynawash, Durawash | Durability Tester | Thiết bị thử độ bền mẫu in khi giặt | Wash Temp Wash Speed Time | Wash Temp (*) | 40 | ±2 | Wash Temp (*) - ℃ | Nhiệt độ giặt | ℃ | 40 | 2 | 38 | 42 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Wash-durability-DEG |  |  | 40 - Wash Temp (*) - ℃
Wash-durability |  |  |  | 10: 50; 12: ±2; 14: ℃; 15: 50; 16: 2; 17: 48; 18: 52; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-durability-DEG |  |  | 50
Wash-durability |  |  |  | 10: 60; 12: ±2; 14: ℃; 15: 60; 16: 2; 17: 58; 18: 62; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Wash-durability-DEG |  |  | 60
Wash-durability | [7 empty] | Wash Speed (*) | 560 | ±2% | Wash Speed (*) - rpm | Tốc độ giặt | rpm | 560 | 11.2 | 548.8 | 571.2 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Wash-durability-RPM |  |  | 560 - Wash Speed (*) - rpm
Wash-durability | [7 empty] | Time | 300 | N/A | Time - sec | Thời gian | sec | 300 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wash-durability-TIME |  |  | 300 - Time - sec
Wash-durability |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-durability- |  |  | *
Wash-durability |  |  |  | 9: 1…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-durability- |  |  | *
Wash-durability |  |  |  | 9: 2…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-durability- |  |  | *
Wash-durability |  |  |  | 9: 3…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-durability- |  |  | *
Wash-durability |  |  |  | 9: 4…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Wash-durability- |  |  | *
Wash-durability |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Wash-durability- |  |  | *
Martindale | ISO 12947-1, ISO 12945-2, ASTM D4970 | Martindale | LAB-LN-03 LAB-M-01 LAB-F01:2023 LAB-A-01 | Martindale | Abrasion & Pilling Tester - Martindale | Thiết bị thử độ mài mòn và xù lông của vải | Lissajous Figures Speed Pressing weights Thickness PACKAGE | Speed (*) | Right - 47.5 | ±2.5 | Speed (*) - rpm | Tốc độ | rpm | 47.5 | 2.5 | 45 | 50 | RPM |  | 0.1 | 2500000 | Service & Calibration - Abrasion >6 heads | 0.1 | 0.1 | Martindale-RPM |  |  | Right - 47.5 - Speed (*) - rpm
Martindale |  |  |  | 10: Inner - 44.5; 12: ±2.4; 14: rpm; 15: 44.5; 16: 2.4; 17: 42.1; 18: 46.9; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Martindale-RPM |  |  | Inner - 44.5
Martindale |  |  |  | 10: Left - 47.5; 12: ±2.5; 14: rpm; 15: 47.5; 16: 2.5; 17: 45; 18: 50; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Martindale-RPM |  |  | Left - 47.5
Martindale | [7 empty] | Tensioning weight (*) | 2500 | ±500 | Tensioning weight (*) - g | Khối lượng tải ép mẫu | g | 2500 | 500 | 2000 | 3000 | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Martindale-MASS |  |  | 2500 - Tensioning weight (*) - g
Martindale | [7 empty] | Pressing Diameter | 120 | ±10 | Pressing Diameter - mm | Đường kính tải ép mẫu | mm | 120 | 10 | 110 | 130 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Martindale-DIM |  |  | 120 - Pressing Diameter - mm
Martindale | [7 empty] | Clamping Ring Inner Diameter | Station 1 - 127 | ±0.5 | Clamping Ring Inner Diameter - mm | Đường kính trong vòng kẹp | mm | 127 | 0.5 | 126.5 | 127.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Martindale-DIM |  |  | Station 1 - 127 - Clamping Ring Inner Diameter - mm
Martindale |  |  |  | 10: Station 2 - 127; 12: ±0.5; 14: mm; 15: 127; 16: 0.5; 17: 126.5; 18: 127.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM |  |  | Station 2 - 127
Martindale |  |  |  | 10: Station 3 - 127; 12: ±0.5; 14: mm; 15: 127; 16: 0.5; 17: 126.5; 18: 127.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM |  |  | Station 3 - 127
Martindale |  |  |  | 10: Station 4 - 127; 12: ±0.5; 14: mm; 15: 127; 16: 0.5; 17: 126.5; 18: 127.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM |  |  | Station 4 - 127
Martindale |  |  |  | 10: Station 5 - 127; 12: ±0.5; 14: mm; 15: 127; 16: 0.5; 17: 126.5; 18: 127.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM |  |  | Station 5 - 127
Martindale |  |  |  | 10: Station 6 - 127; 12: ±0.5; 14: mm; 15: 127; 16: 0.5; 17: 126.5; 18: 127.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM |  |  | Station 6 - 127
Martindale |  |  |  | 10: Station 7 - 127; 12: ±0.5; 14: mm; 15: 127; 16: 0.5; 17: 126.5; 18: 127.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM |  |  | Station 7 - 127
Martindale |  |  |  | 10: Station 8 - 127; 12: ±0.5; 14: mm; 15: 127; 16: 0.5; 17: 126.5; 18: 127.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM |  |  | Station 8 - 127
Martindale |  |  |  | 10: Station 9 - 127; 12: ±0.5; 14: mm; 15: 127; 16: 0.5; 17: 126.5; 18: 127.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM |  |  | Station 9 - 127
Martindale | [7 empty] | Sample Holder Insert Protrusion | Station 1 - 1.05 | ±0.1 | Sample Holder Insert Protrusion - mm | Độ lồi chén giữ mẫu | mm | 1.05 | 0.5 | 0.55 | 1.55 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Martindale-DIM |  |  | Station 1 - 1.05 - Sample Holder Insert Protrusion - mm
Martindale |  |  |  | 10: Station 2 - 1.05; 12: ±0.1; 14: mm; 15: 1.05; 16: 0.5; 17: 0.55; 18: 1.55; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 3 - 1.05; 12: ±0.1; 14: mm; 15: 1.05; 16: 0.5; 17: 0.55; 18: 1.55; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 4 - 1.05; 12: ±0.1; 14: mm; 15: 1.05; 16: 0.5; 17: 0.55; 18: 1.55; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 5 - 1.05; 12: ±0.1; 14: mm; 15: 1.05; 16: 0.5; 17: 0.55; 18: 1.55; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 6 - 1.05; 12: ±0.1; 14: mm; 15: 1.05; 16: 0.5; 17: 0.55; 18: 1.55; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 7 - 1.05; 12: ±0.1; 14: mm; 15: 1.05; 16: 0.5; 17: 0.55; 18: 1.55; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 8 - 1.05; 12: ±0.1; 14: mm; 15: 1.05; 16: 0.5; 17: 0.55; 18: 1.55; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 9 - 1.05; 12: ±0.1; 14: mm; 15: 1.05; 16: 0.5; 17: 0.55; 18: 1.55; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale
Martindale | [7 empty] | Lissajous Graphic | Station 1 - 60.5x60.5 | ±0.5 | Lissajous Graphic - mm | Kích thước hình Lissajous | mm | 60.5 | 0.5 | 60 | 61 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 2 - 60.5x60.5; 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 3 - 60.5x60.5; 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 4 - 60.5x60.5; 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 5 - 60.5x60.5; 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 6 - 60.5x60.5; 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 7 - 60.5x60.5; 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 8 - 60.5x60.5; 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 9 - 60.5x60.5; 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale | [7 empty] | Abrasion Sample Holder Mass (*) | Station 1 - 198 | ±2 | Abrasion Sample Holder Mass (*) - g | Khối lượng chén giữ mẫu | g | 198 | 2 | 196 | 200 | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 2 - 198; 12: ±2; 14: g; 15: 198; 16: 2; 17: 196; 18: 200; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 3 - 198; 12: ±2; 14: g; 15: 198; 16: 2; 17: 196; 18: 200; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 4 - 198; 12: ±2; 14: g; 15: 198; 16: 2; 17: 196; 18: 200; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 5 - 198; 12: ±2; 14: g; 15: 198; 16: 2; 17: 196; 18: 200; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 6 - 198; 12: ±2; 14: g; 15: 198; 16: 2; 17: 196; 18: 200; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 7 - 198; 12: ±2; 14: g; 15: 198; 16: 2; 17: 196; 18: 200; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 8 - 198; 12: ±2; 14: g; 15: 198; 16: 2; 17: 196; 18: 200; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 9 - 198; 12: ±2; 14: g; 15: 198; 16: 2; 17: 196; 18: 200; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale | [7 empty] | Abrasion mass 9kpa (*) | Station 1 - 595 | ±7 | Abrasion mass 9kpa (*) - g | Khối lượng tải mài mòn (9Kpa) | g | 595 | 7 | 588 | 602 | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 2 - 595; 12: ±7; 14: g; 15: 595; 16: 7; 17: 588; 18: 602; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 3 - 595; 12: ±7; 14: g; 15: 595; 16: 7; 17: 588; 18: 602; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 4 - 595; 12: ±7; 14: g; 15: 595; 16: 7; 17: 588; 18: 602; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 5 - 595; 12: ±7; 14: g; 15: 595; 16: 7; 17: 588; 18: 602; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 6 - 595; 12: ±7; 14: g; 15: 595; 16: 7; 17: 588; 18: 602; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 7 - 595; 12: ±7; 14: g; 15: 595; 16: 7; 17: 588; 18: 602; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 8 - 595; 12: ±7; 14: g; 15: 595; 16: 7; 17: 588; 18: 602; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 9 - 595; 12: ±7; 14: g; 15: 595; 16: 7; 17: 588; 18: 602; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale | [7 empty] | Abrasion mass 12kpa (*) | Station 1 - 795 | ±7 | Abrasion mass 12kpa (*) - g | Khối lượng tải mài mòn (12Kpa) | g | 795 | 7 | 788 | 802 | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 2 - 795; 12: ±7; 14: g; 15: 795; 16: 7; 17: 788; 18: 802; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 3 - 795; 12: ±7; 14: g; 15: 795; 16: 7; 17: 788; 18: 802; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 4 - 795; 12: ±7; 14: g; 15: 795; 16: 7; 17: 788; 18: 802; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 5 - 795; 12: ±7; 14: g; 15: 795; 16: 7; 17: 788; 18: 802; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 6 - 795; 12: ±7; 14: g; 15: 795; 16: 7; 17: 788; 18: 802; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 7 - 795; 12: ±7; 14: g; 15: 795; 16: 7; 17: 788; 18: 802; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 8 - 795; 12: ±7; 14: g; 15: 795; 16: 7; 17: 788; 18: 802; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS |  |  | Station 8 - 795 | Station 8 - 795 -
Martindale |  |  |  | 10: Station 9 - 795; 12: ±7; 14: g; 15: 795; 16: 7; 17: 788; 18: 802; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS |  |  | Station 9 - 795 | Station 9 - 795 -
Martindale | [7 empty] | Lissajous Graphic | Station 1 - 24x24 | ±0.5 | Lissajous Graphic - mm | Kích thước hình Lissajous | mm | 24 | 0.5 | 23.5 | 24.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Martindale-DIM |  |  | Station 1 - 24x24 - Lissajous Graphic - mm | Station 1 - 24x24 - Lissajous Graphic - mm
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM |  |  | 0 | -
Martindale |  |  |  | 10: Station 2 - 24x24; 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM |  |  | Station 2 - 24x24 | Station 2 - 24x24 -
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM |  |  | 0 | -
Martindale |  |  |  | 10: Station 3 - 24x24; 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM |  |  | Station 3 - 24x24 | Station 3 - 24x24 -
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM |  |  | 0 | -
Martindale |  |  |  | 10: Station 4 - 24x24; 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM |  |  | Station 4 - 24x24 | Station 4 - 24x24 -
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM |  |  | 0 | -
Martindale |  |  |  | 10: Station 5 - 24x24; 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 6 - 24x24; 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 7 - 24x24; 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 8 - 24x24; 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 10: Station 9 - 24x24; 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale |  |  |  | 12: ±0.5; 14: mm; 15: 24; 16: 0.5; 17: 23.5; 18: 24.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Martindale-DIM
Martindale | [7 empty] | Sample holder mass (*) | Station 1 - 155 | ±1 | Sample holder mass (*) - g | Khối lượng thanh giữ mẫu thử | g | 155 | 1 | 154 | 156 | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 2 - 155; 12: ±1; 14: g; 15: 155; 16: 1; 17: 154; 18: 156; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 3 - 155; 12: ±1; 14: g; 15: 155; 16: 1; 17: 154; 18: 156; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 4 - 155; 12: ±1; 14: g; 15: 155; 16: 1; 17: 154; 18: 156; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 5 - 155; 12: ±1; 14: g; 15: 155; 16: 1; 17: 154; 18: 156; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 6 - 155; 12: ±1; 14: g; 15: 155; 16: 1; 17: 154; 18: 156; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 7 - 155; 12: ±1; 14: g; 15: 155; 16: 1; 17: 154; 18: 156; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 8 - 155; 12: ±1; 14: g; 15: 155; 16: 1; 17: 154; 18: 156; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS
Martindale |  |  |  | 10: Station 9 - 155; 12: ±1; 14: g; 15: 155; 16: 1; 17: 154; 18: 156; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS |  |  | Station 9 - 155 | Station 9 - 155 -
Martindale | [7 empty] | Pilling mass (*) | Station 1 - 260 | ±1 | Pilling mass (*) - g | Khối lượng miếng tải trọng thử vón gút | g | 260 | 1 | 259 | 261 | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Martindale-MASS |  |  | Station 1 - 260 - Pilling mass (*) - g | Station 1 - 260 - Pilling mass (*) - g
Martindale |  |  |  | 10: Station 2 - 260; 12: ±1; 14: g; 15: 260; 16: 1; 17: 259; 18: 261; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS |  |  | Station 2 - 260 | Station 2 - 260 -
Martindale |  |  |  | 10: Station 3 - 260; 12: ±1; 14: g; 15: 260; 16: 1; 17: 259; 18: 261; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS |  |  | Station 3 - 260 | Station 3 - 260 -
Martindale |  |  |  | 10: Station 4 - 260; 12: ±1; 14: g; 15: 260; 16: 1; 17: 259; 18: 261; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS |  |  | Station 4 - 260 | Station 4 - 260 -
Martindale |  |  |  | 10: Station 5 - 260; 12: ±1; 14: g; 15: 260; 16: 1; 17: 259; 18: 261; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS |  |  | Station 5 - 260 | Station 5 - 260 -
Martindale |  |  |  | 10: Station 6 - 260; 12: ±1; 14: g; 15: 260; 16: 1; 17: 259; 18: 261; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS |  |  | Station 6 - 260 | Station 6 - 260 -
Martindale |  |  |  | 10: Station 7 - 260; 12: ±1; 14: g; 15: 260; 16: 1; 17: 259; 18: 261; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS |  |  | Station 7 - 260 | Station 7 - 260 -
Martindale |  |  |  | 10: Station 8 - 260; 12: ±1; 14: g; 15: 260; 16: 1; 17: 259; 18: 261; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS |  |  | Station 8 - 260 | Station 8 - 260 -
Martindale |  |  |  | 10: Station 9 - 260; 12: ±1; 14: g; 15: 260; 16: 1; 17: 259; 18: 261; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Martindale-MASS |  |  | Station 9 - 260 | Station 9 - 260 -
Martindale |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Martindale-
Martindale |  |  |  | 9: 1…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Martindale-
Martindale |  |  |  | 9: 2…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Martindale-
Martindale |  |  |  | 9: 3…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Martindale-
Martindale |  |  |  | 9: 4…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Martindale-
Martindale |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Martindale-
Martindale-fw | Adidas PHM-FW0263 | Abrasion | LAB-LN-03 LAB-M-01 LAB-F01:2023 LAB-A-01 | Martindale | Abrasion & Pilling Tester - Martindale | Thiết bị thử độ mài mòn và xù lông của vải | Lissajous Figures Speed Pressing weights Thickness PACKAGE | Speed (*) | Right - 47.5 | ±2.5 | Speed (*) - rpm | Tốc độ | rpm | 47.5 | 2.5 | 45 | 50 | RPM |  | 0.2 | 2500000 | Service & Calibration - Abrasion >6 heads | 0.1 | 0.2 | Martindale-fw-RPM
Martindale-fw |  |  |  | 10: Inner - 44.5; 12: ±2.4; 14: rpm; 15: 44.5; 16: 2.4; 17: 42.1; 18: 46.9; 19: RPM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-RPM
Martindale-fw |  |  |  | 10: Left - 47.5; 12: ±2.5; 14: rpm; 15: 47.5; 16: 2.5; 17: 45; 18: 50; 19: RPM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-RPM
Martindale-fw | [7 empty] | Tensioning weight (*) | 2500 | ±500 | Tensioning weight (*) - g | Khối lượng tải ép mẫu | g | 2500 | 500 | 2000 | 3000 | MASS |  | 0.2 |  |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw | [7 empty] | Pressing Diameter | 120 | ±10 | Pressing Diameter - mm | Đường kính tải ép mẫu | mm | 120 | 10 | 110 | 130 | DIM |  | 0.2 |  |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw | [7 empty] | Clamping Ring Inner Diameter | Station 1 - 127 | ±0.5 | Clamping Ring Inner Diameter - mm | Đường kính trong vòng kẹp | mm | 127 | 0.5 | 126.5 | 127.5 | DIM |  | 0.2 |  |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 2 - 127; 12: ±0.5; 14: mm; 15: 127; 16: 0.5; 17: 126.5; 18: 127.5; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 3 - 127; 12: ±0.5; 14: mm; 15: 127; 16: 0.5; 17: 126.5; 18: 127.5; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 4 - 127; 12: ±0.5; 14: mm; 15: 127; 16: 0.5; 17: 126.5; 18: 127.5; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 5 - 127; 12: ±0.5; 14: mm; 15: 127; 16: 0.5; 17: 126.5; 18: 127.5; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 6 - 127; 12: ±0.5; 14: mm; 15: 127; 16: 0.5; 17: 126.5; 18: 127.5; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 7 - 127; 12: ±0.5; 14: mm; 15: 127; 16: 0.5; 17: 126.5; 18: 127.5; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 8 - 127; 12: ±0.5; 14: mm; 15: 127; 16: 0.5; 17: 126.5; 18: 127.5; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 9 - 127; 12: ±0.5; 14: mm; 15: 127; 16: 0.5; 17: 126.5; 18: 127.5; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | Station 9 - 127 | Station 9 - 127 -
Martindale-fw | [7 empty] | Sample Holder Insert Protrusion | Station 1 - 1.05 | ±0.1 | Sample Holder Insert Protrusion - mm | Độ lồi chén giữ mẫu | mm | 1.05 | 0.5 | 0.55 | 1.55 | DIM |  | 0.2 |  |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | Point 1 - 70°C | Point 1 - 70°C -
Martindale-fw |  |  |  | 10: Station 2 - 1.05; 12: ±0.1; 14: mm; 15: 1.05; 16: 0.5; 17: 0.55; 18: 1.55; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 3 - 1.05; 12: ±0.1; 14: mm; 15: 1.05; 16: 0.5; 17: 0.55; 18: 1.55; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 4 - 1.05; 12: ±0.1; 14: mm; 15: 1.05; 16: 0.5; 17: 0.55; 18: 1.55; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 5 - 1.05; 12: ±0.1; 14: mm; 15: 1.05; 16: 0.5; 17: 0.55; 18: 1.55; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 6 - 1.05; 12: ±0.1; 14: mm; 15: 1.05; 16: 0.5; 17: 0.55; 18: 1.55; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 7 - 1.05; 12: ±0.1; 14: mm; 15: 1.05; 16: 0.5; 17: 0.55; 18: 1.55; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 8 - 1.05; 12: ±0.1; 14: mm; 15: 1.05; 16: 0.5; 17: 0.55; 18: 1.55; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 9 - 1.05; 12: ±0.1; 14: mm; 15: 1.05; 16: 0.5; 17: 0.55; 18: 1.55; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw | [7 empty] | Lissajous Graphic | Station 1 - 60.5x60.5 | ±0.5 | Lissajous Graphic - mm | Kích thước hình Lissajous | mm | 60.5 | 0.5 | 60 | 61 | DIM |  | 0.2 |  |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | Point 1 - 80°C | Point 1 - 80°C -
Martindale-fw |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | 0 | -
Martindale-fw |  |  |  | 10: Station 2 - 60.5x60.5; 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | Station 2 - 60.5x60.5 | Station 2 - 60.5x60.5 -
Martindale-fw |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | 0 | -
Martindale-fw |  |  |  | 10: Station 3 - 60.5x60.5; 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | Station 3 - 60.5x60.5 | Station 3 - 60.5x60.5 -
Martindale-fw |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | 0 | -
Martindale-fw |  |  |  | 10: Station 4 - 60.5x60.5; 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | Station 4 - 60.5x60.5 | Station 4 - 60.5x60.5 -
Martindale-fw |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | 0 | -
Martindale-fw |  |  |  | 10: Station 5 - 60.5x60.5; 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | Station 5 - 60.5x60.5 | Station 5 - 60.5x60.5 -
Martindale-fw |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | 0 | -
Martindale-fw |  |  |  | 10: Station 6 - 60.5x60.5; 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | Station 6 - 60.5x60.5 | Station 6 - 60.5x60.5 -
Martindale-fw |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | 0 | -
Martindale-fw |  |  |  | 10: Station 7 - 60.5x60.5; 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | Station 7 - 60.5x60.5 | Station 7 - 60.5x60.5 -
Martindale-fw |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | 0 | -
Martindale-fw |  |  |  | 10: Station 8 - 60.5x60.5; 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | Station 8 - 60.5x60.5 | Station 8 - 60.5x60.5 -
Martindale-fw |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | 0 | -
Martindale-fw |  |  |  | 10: Station 9 - 60.5x60.5; 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | Station 9 - 60.5x60.5 | Station 9 - 60.5x60.5 -
Martindale-fw |  |  |  | 12: ±0.5; 14: mm; 15: 60.5; 16: 0.5; 17: 60; 18: 61; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM |  |  | 0 | -
Martindale-fw | [7 empty] | Abrasion Sample Holder Mass (*) | Station 1 - 198 | ±2 | Abrasion Sample Holder Mass (*) - g | Khối lượng chén giữ mẫu | g | 198 | 2 | 196 | 200 | MASS |  | 0.2 |  |  | 0.1 | 0.2 | Martindale-fw-MASS |  |  | Station 1 - 198 - Abrasion Sample Holder Mass (*) - g | Station 1 - 198 - Abrasion Sample Holder Mass (*) - g
Martindale-fw |  |  |  | 10: Station 2 - 198; 12: ±2; 14: g; 15: 198; 16: 2; 17: 196; 18: 200; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS |  |  | Station 2 - 198 | Station 2 - 198 -
Martindale-fw |  |  |  | 10: Station 3 - 198; 12: ±2; 14: g; 15: 198; 16: 2; 17: 196; 18: 200; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS |  |  | Station 3 - 198
Martindale-fw |  |  |  | 10: Station 4 - 198; 12: ±2; 14: g; 15: 198; 16: 2; 17: 196; 18: 200; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw |  |  |  | 10: Station 5 - 198; 12: ±2; 14: g; 15: 198; 16: 2; 17: 196; 18: 200; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS |  |  | Station 5 - 198
Martindale-fw |  |  |  | 10: Station 6 - 198; 12: ±2; 14: g; 15: 198; 16: 2; 17: 196; 18: 200; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS |  |  | Station 6 - 198
Martindale-fw |  |  |  | 10: Station 7 - 198; 12: ±2; 14: g; 15: 198; 16: 2; 17: 196; 18: 200; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw |  |  |  | 10: Station 8 - 198; 12: ±2; 14: g; 15: 198; 16: 2; 17: 196; 18: 200; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS |  |  | Station 8 - 198
Martindale-fw |  |  |  | 10: Station 9 - 198; 12: ±2; 14: g; 15: 198; 16: 2; 17: 196; 18: 200; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS |  |  | Station 9 - 198
Martindale-fw | [7 empty] | Abrasion mass 9kpa (*) | Station 1 - 595 | ±7 | Abrasion mass 9kpa (*) - g | Khối lượng tải mài mòn (9Kpa) | g | 595 | 7 | 588 | 602 | MASS |  | 0.2 |  |  | 0.1 | 0.2 | Martindale-fw-MASS |  |  | Station 1 - 595 - Abrasion mass 9kpa (*) - g
Martindale-fw |  |  |  | 10: Station 2 - 595; 12: ±7; 14: g; 15: 595; 16: 7; 17: 588; 18: 602; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS |  |  | Station 2 - 595
Martindale-fw |  |  |  | 10: Station 3 - 595; 12: ±7; 14: g; 15: 595; 16: 7; 17: 588; 18: 602; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS |  |  | Station 3 - 595
Martindale-fw |  |  |  | 10: Station 4 - 595; 12: ±7; 14: g; 15: 595; 16: 7; 17: 588; 18: 602; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw |  |  |  | 10: Station 5 - 595; 12: ±7; 14: g; 15: 595; 16: 7; 17: 588; 18: 602; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw |  |  |  | 10: Station 6 - 595; 12: ±7; 14: g; 15: 595; 16: 7; 17: 588; 18: 602; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw |  |  |  | 10: Station 7 - 595; 12: ±7; 14: g; 15: 595; 16: 7; 17: 588; 18: 602; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw |  |  |  | 10: Station 8 - 595; 12: ±7; 14: g; 15: 595; 16: 7; 17: 588; 18: 602; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw |  |  |  | 10: Station 9 - 595; 12: ±7; 14: g; 15: 595; 16: 7; 17: 588; 18: 602; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw | [7 empty] | Abrasion mass 12kpa (*) | Station 1 - 795 | ±7 | Abrasion mass 12kpa (*) - g | Khối lượng tải mài mòn (12Kpa) | g | 795 | 7 | 788 | 802 | MASS |  | 0.2 |  |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw |  |  |  | 10: Station 2 - 795; 12: ±7; 14: g; 15: 795; 16: 7; 17: 788; 18: 802; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw |  |  |  | 10: Station 3 - 795; 12: ±7; 14: g; 15: 795; 16: 7; 17: 788; 18: 802; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw |  |  |  | 10: Station 4 - 795; 12: ±7; 14: g; 15: 795; 16: 7; 17: 788; 18: 802; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw |  |  |  | 10: Station 5 - 795; 12: ±7; 14: g; 15: 795; 16: 7; 17: 788; 18: 802; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw |  |  |  | 10: Station 6 - 795; 12: ±7; 14: g; 15: 795; 16: 7; 17: 788; 18: 802; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw |  |  |  | 10: Station 7 - 795; 12: ±7; 14: g; 15: 795; 16: 7; 17: 788; 18: 802; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw |  |  |  | 10: Station 8 - 795; 12: ±7; 14: g; 15: 795; 16: 7; 17: 788; 18: 802; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw |  |  |  | 10: Station 9 - 795; 12: ±7; 14: g; 15: 795; 16: 7; 17: 788; 18: 802; 19: MASS; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-MASS
Martindale-fw | [7 empty] | Angle - ⁰ | Station 1 - 45 | ±0.15 | Angle - ⁰ | Góc nghiêng chén | ⁰ | 45 | 0.15 | 44.85 | 45.15 | ANG |  | #N/A |  |  | #N/A | #N/A | Martindale-fw-ANG
Martindale-fw |  |  |  | 10: Station 2 - 45; 12: ±0.15; 14: ⁰; 15: 45; 16: 0.15; 17: 44.85; 18: 45.15; 19: ANG; 21: #N/A |  | #N/A | #N/A | Martindale-fw-ANG
Martindale-fw |  |  |  | 10: Station 3 - 45; 12: ±0.15; 14: ⁰; 15: 45; 16: 0.15; 17: 44.85; 18: 45.15; 19: ANG; 21: #N/A |  | #N/A | #N/A | Martindale-fw-ANG
Martindale-fw |  |  |  | 10: Station 4 - 45; 12: ±0.15; 14: ⁰; 15: 45; 16: 0.15; 17: 44.85; 18: 45.15; 19: ANG; 21: #N/A |  | #N/A | #N/A | Martindale-fw-ANG
Martindale-fw |  |  |  | 10: Station 5 - 45; 12: ±0.15; 14: ⁰; 15: 45; 16: 0.15; 17: 44.85; 18: 45.15; 19: ANG; 21: #N/A |  | #N/A | #N/A | Martindale-fw-ANG
Martindale-fw |  |  |  | 10: Station 6 - 45; 12: ±0.15; 14: ⁰; 15: 45; 16: 0.15; 17: 44.85; 18: 45.15; 19: ANG; 21: #N/A |  | #N/A | #N/A | Martindale-fw-ANG
Martindale-fw |  |  |  | 10: Station 7 - 45; 12: ±0.15; 14: ⁰; 15: 45; 16: 0.15; 17: 44.85; 18: 45.15; 19: ANG; 21: #N/A |  | #N/A | #N/A | Martindale-fw-ANG
Martindale-fw |  |  |  | 10: Station 8 - 45; 12: ±0.15; 14: ⁰; 15: 45; 16: 0.15; 17: 44.85; 18: 45.15; 19: ANG; 21: #N/A |  | #N/A | #N/A | Martindale-fw-ANG
Martindale-fw |  |  |  | 10: Station 9 - 45; 12: ±0.15; 14: ⁰; 15: 45; 16: 0.15; 17: 44.85; 18: 45.15; 19: ANG; 21: #N/A |  | #N/A | #N/A | Martindale-fw-ANG
Martindale-fw | [7 empty] | Diameter - mm | Station 1 - 38.1 | +0.050 -0.127 | Diameter - mm | Đường kính chén | mm | 38.1 | + 0.050 -0.127 | 37.973 | 38.15 | DIM |  | 0.2 |  |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 2 - 38.1; 12: + 0.050 -0.127; 14: mm; 15: 38.1; 16: + 0.050 -0.127; 17: 37.973; 18: 38.15; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 3 - 38.1; 12: + 0.050 -0.127; 14: mm; 15: 38.1; 16: + 0.050 -0.127; 17: 37.973; 18: 38.15; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 4 - 38.1; 12: + 0.050 -0.127; 14: mm; 15: 38.1; 16: + 0.050 -0.127; 17: 37.973; 18: 38.15; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 5 - 38.1; 12: + 0.050 -0.127; 14: mm; 15: 38.1; 16: + 0.050 -0.127; 17: 37.973; 18: 38.15; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 6 - 38.1; 12: + 0.050 -0.127; 14: mm; 15: 38.1; 16: + 0.050 -0.127; 17: 37.973; 18: 38.15; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 7 - 38.1; 12: + 0.050 -0.127; 14: mm; 15: 38.1; 16: + 0.050 -0.127; 17: 37.973; 18: 38.15; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 8 - 38.1; 12: + 0.050 -0.127; 14: mm; 15: 38.1; 16: + 0.050 -0.127; 17: 37.973; 18: 38.15; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 9 - 38.1; 12: + 0.050 -0.127; 14: mm; 15: 38.1; 16: + 0.050 -0.127; 17: 37.973; 18: 38.15; 19: DIM; 21: #N/A |  | #N/A | #N/A | OVEN70-1-DEG
Martindale-fw | [7 empty] | Heigh | Station 1 | N/A | Heigh - mm | Chiều cao chén | mm | N/A | N/A | #VALUE! | #VALUE! | DIM |  | #N/A |  |  | #N/A | #N/A | OVEN70-1-DEG
Martindale-fw |  |  |  | 10: Station 2; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: #N/A |  | #N/A | #N/A | OVEN70-1-DEG
Martindale-fw |  |  |  | 10: Station 3; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: #N/A |  | #N/A | #N/A | OVEN70-1-DEG
Martindale-fw |  |  |  | 10: Station 4; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: #N/A |  | #N/A | #N/A | OVEN70-1-DEG
Martindale-fw |  |  |  | 10: Station 5; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: #N/A |  | #N/A | #N/A | OVEN70-1-DEG
Martindale-fw |  |  |  | 10: Station 6; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: #N/A |  | #N/A | #N/A | OVEN70-1-DEG
Martindale-fw |  |  |  | 10: Station 7; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: #N/A |  | #N/A | #N/A | OVEN70-1-DEG
Martindale-fw |  |  |  | 10: Station 8; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: #N/A |  | #N/A | #N/A | OVEN70-1-DEG
Martindale-fw |  |  |  | 10: Station 9; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: #N/A |  | #N/A | #N/A | OVEN80-9-DEG
Martindale-fw | [7 empty] | Diameter Sample Holder Insert Protrusion | Station 1 - 28.65 | ± 0.05 | Diameter Sample Holder Insert Protrusion - mm | Đường kính độ lồi chén giữ mẫu | mm | 28.65 | 0.05 | 28.6 | 28.7 | DIM |  | 0.2 |  |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 2 - 28.65; 12: ± 0.05; 14: mm; 15: 28.65; 16: 0.05; 17: 28.6; 18: 28.7; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 3 - 28.65; 12: ± 0.05; 14: mm; 15: 28.65; 16: 0.05; 17: 28.6; 18: 28.7; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 4 - 28.65; 12: ± 0.05; 14: mm; 15: 28.65; 16: 0.05; 17: 28.6; 18: 28.7; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 5 - 28.65; 12: ± 0.05; 14: mm; 15: 28.65; 16: 0.05; 17: 28.6; 18: 28.7; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 6 - 28.65; 12: ± 0.05; 14: mm; 15: 28.65; 16: 0.05; 17: 28.6; 18: 28.7; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 7 - 28.65; 12: ± 0.05; 14: mm; 15: 28.65; 16: 0.05; 17: 28.6; 18: 28.7; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 8 - 28.65; 12: ± 0.05; 14: mm; 15: 28.65; 16: 0.05; 17: 28.6; 18: 28.7; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 9 - 28.65; 12: ± 0.05; 14: mm; 15: 28.65; 16: 0.05; 17: 28.6; 18: 28.7; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw | [7 empty] | Flatness of Abrading table | Station 1 | <0.10 | Flatness of Abrading table - mm | Độ phẳng bàn mài mòn | mm | 0 | 0.1 | -0.1 | 0.1 | DIM |  | 0.2 |  |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 2; 12: <0.10; 14: mm; 15: 0; 16: 0.1; 17: -0.1; 18: 0.1; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 3; 12: <0.10; 14: mm; 15: 0; 16: 0.1; 17: -0.1; 18: 0.1; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 4; 12: <0.10; 14: mm; 15: 0; 16: 0.1; 17: -0.1; 18: 0.1; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 5; 12: <0.10; 14: mm; 15: 0; 16: 0.1; 17: -0.1; 18: 0.1; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 6; 12: <0.10; 14: mm; 15: 0; 16: 0.1; 17: -0.1; 18: 0.1; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 7; 12: <0.10; 14: mm; 15: 0; 16: 0.1; 17: -0.1; 18: 0.1; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 8; 12: <0.10; 14: mm; 15: 0; 16: 0.1; 17: -0.1; 18: 0.1; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 10: Station 9; 12: <0.10; 14: mm; 15: 0; 16: 0.1; 17: -0.1; 18: 0.1; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-fw |  |  |  | 9: Spacing of lines in Lissajours figures; 12: N/A; 11: Spacing of lines in Lissajours figures; 13: Khoảng cách giữa các đường; 19: DIM; 21: 0.2 |  | 0.1 | 0.2 | Martindale-fw-DIM
Martindale-mount | EN ISO 12947-2 |  | LAB-F-01 LAB-M-01 LAB-LN-03 | 5: Specimen Preperation for Stretch; 6: Specimen Preperation for Stretch; 9: Total mass with clamp - g; 10: 1 - 100; 12: N/A; 11: Total mass with clamp - g; 14: g; 15: 100; 19: MASS |  |  |  | Martindale-mount-MASS
Martindale-mount |  |  |  | 10: 2 - 200; 12: N/A; 14: g; 15: 200; 19: MASS
Martindale-mount |  |  |  | 10: 3 - 100; 12: N/A; 14: g; 15: 100; 19: MASS
Martindale-mount |  |  |  | 10: 4 - 100; 12: N/A; 14: g; 15: 100; 19: MASS
Martindale-mount |  |  |  | 9: Total mount - mm; 10: A-B 45; 12: N/A; 11: Total mount - mm; 14: mm; 15: 45; 19: DIM
Martindale-mount |  |  |  | 10: C-D 45; 12: N/A; 14: mm; 15: 45; 19: DIM
Martindale-mount |  |  |  | 9: Lowering console - mm; 10: Ø100; 12: N/A; 11: Lowering console - mm; 14: mm; 15: 20; 19: DIM
Martindale-mount |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Martindale-mount-
Martindale-mount |  |  |  | 9: 1…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Martindale-mount-
Martindale-mount |  |  |  | 9: 2…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Martindale-mount-
Martindale-mount |  |  |  | 9: 3…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Martindale-mount-
Martindale-mount |  |  |  | 9: 4…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Martindale-mount-
Martindale-mount |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Martindale-mount-
Oven-other | Customer's Test Method | Colorfastness to Dye Migration - Solid Đyed Fabrics, materials | LAB-T-01 LAB-TIME-02 | 5: Memmert, Binder, Oven/Incubator; 6: Oven/Incubator; 7: Tủ sấy / Tủ ấm; 8: Temperature (*); 9: GT hiển thị; 20: #N/A; 22: 600000 | Service & Calibration - 1 points | #N/A | #N/A | Oven-other-
Oven-other |  |  |  | 9: Temp (*) - °C - Xác nhận với KH; 10: 1 -; 11: Temp (*) - °C; 13: Nhiệt độ; 14: ℃; 17: 0; 18: 0; 19: DEG
Oven-other |  |  |  | 10: 2 -; 14: ℃; 17: 0; 18: 0; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 3 -; 14: ℃; 17: 0; 18: 0; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 4 -; 14: ℃; 17: 0; 18: 0; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 5 -; 14: ℃; 17: 0; 18: 0; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 6 -; 14: ℃; 17: 0; 18: 0; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 7 -; 14: ℃; 17: 0; 18: 0; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 8 -; 14: ℃; 17: 0; 18: 0; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 9 -; 14: ℃; 17: 0; 18: 0; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 9: GT hiển thị |  | #N/A | #N/A | Oven-other-
Oven-other |  |  |  | 9: Temp (*) - °C; 10: 1 -; 11: Temp (*) - °C; 13: Nhiệt độ; 14: ℃; 17: 0; 18: 0; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 2 -; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 3 -; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 4 -; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 5 -; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 6 -; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 7 -; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 8 -; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 9 -; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 9: GT hiển thị |  | #N/A | #N/A | Oven-other-
Oven-other |  |  |  | 9: Temp (*) - °C; 10: 1 -; 11: Temp (*) - °C; 13: Nhiệt độ; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 2 -; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 3 -; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 4 -; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 5 -; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 6 -; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 7 -; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 8 -; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-other |  |  |  | 10: 9 -; 14: ℃; 19: DEG |  | 0.1 | 0.2 | Oven-other-DEG
Oven-all | ISO 105, Colourfastness to Migration, Perspiration |  | LAB-T-01 LAB-TIME-02 | 6: Oven; 7: Tủ sấy; 8: Temperature (*); 9: GT hiển thị; 10: 37 | Service & Calibration - 9 points | #N/A | #N/A | Oven-all-
Oven-all | [7 empty] | Temp (*) - °C | 1 - 37 | ±2 | Temp (*) - °C | Nhiệt độ | ℃ | 37 | 2 | 35 | 39 | DEG |  |  |  |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 2 - 37; 12: ±2; 14: ℃; 15: 37; 16: 2; 17: 35; 18: 39; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 3 - 37; 12: ±2; 14: ℃; 15: 37; 16: 2; 17: 35; 18: 39; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 4 - 37; 12: ±2; 14: ℃; 15: 37; 16: 2; 17: 35; 18: 39; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 5 - 37; 12: ±2; 14: ℃; 15: 37; 16: 2; 17: 35; 18: 39; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 6 - 37; 12: ±2; 14: ℃; 15: 37; 16: 2; 17: 35; 18: 39; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 7 - 37; 12: ±2; 14: ℃; 15: 37; 16: 2; 17: 35; 18: 39; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 8 - 37; 12: ±2; 14: ℃; 15: 37; 16: 2; 17: 35; 18: 39; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 9 - 37; 12: ±2; 14: ℃; 15: 37; 16: 2; 17: 35; 18: 39; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 9: GT hiển thị; 10: 38
Oven-all | [7 empty] | Temp (*) - °C | 1 - 38 | ±2 | Temp (*) - °C | Nhiệt độ | ℃ | 38 | 2 | 36 | 40 | DEG |  |  |  |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 2 - 38; 12: ±2; 14: ℃; 15: 38; 16: 2; 17: 36; 18: 40; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 3 - 38; 12: ±2; 14: ℃; 15: 38; 16: 2; 17: 36; 18: 40; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 4 - 38; 12: ±2; 14: ℃; 15: 38; 16: 2; 17: 36; 18: 40; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 5 - 38; 12: ±2; 14: ℃; 15: 38; 16: 2; 17: 36; 18: 40; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 6 - 38; 12: ±2; 14: ℃; 15: 38; 16: 2; 17: 36; 18: 40; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 7 - 38; 12: ±2; 14: ℃; 15: 38; 16: 2; 17: 36; 18: 40; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 8 - 38; 12: ±2; 14: ℃; 15: 38; 16: 2; 17: 36; 18: 40; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 9 - 38; 12: ±2; 14: ℃; 15: 38; 16: 2; 17: 36; 18: 40; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 9: GT hiển thị; 10: 50
Oven-all | [7 empty] | Temp (*) - °C | 1 - 50 | ±2 | Temp (*) - °C | Nhiệt độ | ℃ | 50 | 2 | 48 | 52 | DEG |  |  | 1200000 | Service & Calibration - 1 points | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 2 - 50; 12: ±2; 14: ℃; 15: 50; 16: 2; 17: 48; 18: 52; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 3 - 50; 12: ±2; 14: ℃; 15: 50; 16: 2; 17: 48; 18: 52; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 4 - 50; 12: ±2; 14: ℃; 15: 50; 16: 2; 17: 48; 18: 52; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 5 - 50; 12: ±2; 14: ℃; 15: 50; 16: 2; 17: 48; 18: 52; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 6 - 50; 12: ±2; 14: ℃; 15: 50; 16: 2; 17: 48; 18: 52; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 7 - 50; 12: ±2; 14: ℃; 15: 50; 16: 2; 17: 48; 18: 52; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 8 - 50; 12: ±2; 14: ℃; 15: 50; 16: 2; 17: 48; 18: 52; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 9 - 50; 12: ±2; 14: ℃; 15: 50; 16: 2; 17: 48; 18: 52; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 9: GT hiển thị; 10: 70
Oven-all | [7 empty] | Temp (*) - °C | 1 - 70 | ±2 | Temp (*) - °C | Nhiệt độ | ℃ | 70 | 2 | 68 | 72 | DEG |  |  | 1200000 | Service & Calibration - 9 points | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 2 - 70; 12: ±2; 14: ℃; 15: 70; 16: 2; 17: 68; 18: 72; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 3 - 70; 12: ±2; 14: ℃; 15: 70; 16: 2; 17: 68; 18: 72; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 4 - 70; 12: ±2; 14: ℃; 15: 70; 16: 2; 17: 68; 18: 72; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 5 - 70; 12: ±2; 14: ℃; 15: 70; 16: 2; 17: 68; 18: 72; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 6 - 70; 12: ±2; 14: ℃; 15: 70; 16: 2; 17: 68; 18: 72; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 7 - 70; 12: ±2; 14: ℃; 15: 70; 16: 2; 17: 68; 18: 72; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 8 - 70; 12: ±2; 14: ℃; 15: 70; 16: 2; 17: 68; 18: 72; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 9 - 70; 12: ±2; 14: ℃; 15: 70; 16: 2; 17: 68; 18: 72; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 9: GT hiển thị; 10: 90
Oven-all | [7 empty] | Temp (*) - °C | 1 - 90 | ±2 | Temp (*) - °C | Nhiệt độ | ℃ | 90 | 2 | 88 | 92 | DEG |  |  |  |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 2 - 90; 12: ±2; 14: ℃; 15: 90; 16: 2; 17: 88; 18: 92; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 3 - 90; 12: ±2; 14: ℃; 15: 90; 16: 2; 17: 88; 18: 92; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 4 - 90; 12: ±2; 14: ℃; 15: 90; 16: 2; 17: 88; 18: 92; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 5 - 90; 12: ±2; 14: ℃; 15: 90; 16: 2; 17: 88; 18: 92; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 6 - 90; 12: ±2; 14: ℃; 15: 90; 16: 2; 17: 88; 18: 92; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 7 - 90; 12: ±2; 14: ℃; 15: 90; 16: 2; 17: 88; 18: 92; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 8 - 90; 12: ±2; 14: ℃; 15: 90; 16: 2; 17: 88; 18: 92; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 10: 9 - 90; 12: ±2; 14: ℃; 15: 90; 16: 2; 17: 88; 18: 92; 19: DEG |  | 0.1 | 0.2 | Oven-all-DEG
Oven-all |  |  |  | 9: GT hiển thị; 11: *
Oven-all |  |  |  | 9: Temp (*) - °C; 10: 1 -; 11: *; 14: ℃; 19: DEG
Oven-all |  |  |  | 10: 2 -; 11: *; 14: ℃; 19: DEG
Oven-all |  |  |  | 10: 3 -; 11: *; 14: ℃; 19: DEG
Oven-all |  |  |  | 10: 4 -; 11: *; 14: ℃; 19: DEG
Oven-all |  |  |  | 10: 5 -; 11: *; 14: ℃; 19: DEG
Oven-all |  |  |  | 10: 6 -; 11: *; 14: ℃; 19: DEG
Oven-all |  |  |  | 10: 7 -; 11: *; 14: ℃; 19: DEG
Oven-all |  |  |  | 10: 8 -; 11: *; 14: ℃; 19: DEG
Oven-all |  |  |  | 10: 9 -; 11: *; 14: ℃; 19: DEG
Heat-therma | ISO 105-PO1, ISO 105-X11, AATCC 117, AATCC 133, M & S P10, Adidas 5.09, JIS L0850, JIS L0879, GB 5718, GB 6154 | Thermal stability, Colour fastness to hot pressing | LAB-T-01 LAB-TIME-02 | Thermaplate | Colour fastness to hot pressing | Thiết bị thử độ bền màu ép nhiệt |  | Temperature 1 (*) | 110 | ±2 | Temperature 1 (*) - ℃ |  | ℃ | 110 | 2 | 108 | 112 | DEG | 0.1 | 0.2 |  |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma |  |  |  | 10: 150; 12: ±2; 14: ℃; 15: 150; 16: 2; 17: 148; 18: 152; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma |  |  |  | 10: 180; 12: ±2; 14: ℃; 15: 180; 16: 2; 17: 178; 18: 182; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma |  |  |  | 10: 200; 12: ±2; 14: ℃; 15: 200; 16: 2; 17: 198; 18: 202; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma |  |  |  | 10: 210; 12: ±2; 14: ℃; 15: 210; 16: 2; 17: 208; 18: 212; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma | [7 empty] | Temperature 2 (*) | 110 | ±2 | Temperature 2 (*) - ℃ |  | ℃ | 110 | 2 | 108 | 112 | DEG | 0.1 | 0.2 |  |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma |  |  |  | 10: 150; 12: ±2; 14: ℃; 15: 150; 16: 2; 17: 148; 18: 152; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma |  |  |  | 10: 180; 12: ±2; 14: ℃; 15: 180; 16: 2; 17: 178; 18: 182; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma |  |  |  | 10: 200; 12: ±2; 14: ℃; 15: 200; 16: 2; 17: 198; 18: 202; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma |  |  |  | 10: 210; 12: ±2; 14: ℃; 15: 210; 16: 2; 17: 208; 18: 212; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma | [7 empty] | Temperature 3 (*) | 110 | ±2 | Temperature 3 (*) - ℃ |  | ℃ | 110 | 2 | 108 | 112 | DEG | 0.1 | 0.2 |  |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma |  |  |  | 10: 150; 12: ±2; 14: ℃; 15: 150; 16: 2; 17: 148; 18: 152; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma |  |  |  | 10: 180; 12: ±2; 14: ℃; 15: 180; 16: 2; 17: 178; 18: 182; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma |  |  |  | 10: 200; 12: ±2; 14: ℃; 15: 200; 16: 2; 17: 198; 18: 202; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma |  |  |  | 10: 210; 12: ±2; 14: ℃; 15: 210; 16: 2; 17: 208; 18: 212; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma |  |  |  | 11: *; 14: ℃; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma |  |  |  | 11: *; 14: ℃; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Heat-therma-DEG
Heat-therma | [7 empty] | Time | 5 | N/A | Time - sec | Bộ đếm thời gian | sec | 5 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Heat-therma-TIME
Heat-therma |  |  |  | 10: 10; 12: N/A; 14: sec; 15: 10; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Heat-therma-TIME
Heat-therma |  |  |  | 10: 15; 12: N/A; 14: sec; 15: 15; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Heat-therma-TIME
Heat-therma |  |  |  | 10: 30; 12: N/A; 14: sec; 15: 30; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Heat-therma-TIME
Heat-therma |  |  |  | 10: 60; 12: N/A; 14: sec; 15: 60; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Heat-therma-TIME
Heat-therma |  |  |  | 12: N/A; 11: *; 14: sec; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Heat-therma-TIME
Heat-therma |  |  |  | 12: N/A; 11: *; 14: sec; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Heat-therma-TIME
Heat-therma | [7 empty] | Dimension of top & bottom plate Kích thước |  | N/A | * | Kích thước | mm | 50 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heat-therma-DIM
Heat-therma |  |  |  | 12: N/A; 11: *; 14: mm; 15: 50; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heat-therma-DIM
Heat-therma |  |  |  | 12: N/A; 11: *; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heat-therma-DIM
Heat-therma |  |  |  | 12: N/A; 11: *; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heat-therma-DIM
Heat-therma |  |  |  | 12: N/A; 11: *; 14: mm; 15: 120; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heat-therma-DIM
Heat-therma |  |  |  | 12: N/A; 11: *; 14: mm; 15: 120; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heat-therma-DIM
Heat-therma |  |  |  | 12: N/A; 11: *; 14: mm; 15: 160; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heat-therma-DIM
Heat-therma |  |  |  | 12: N/A; 11: *; 14: mm; 15: 160; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heat-therma-DIM
Heat-therma |  |  |  | 12: N/A; 11: *; 14: mm; 15: 200; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heat-therma-DIM
Heat-therma |  |  |  | 12: N/A; 11: *; 14: mm; 15: 200; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heat-therma-DIM |  |  | *
Heat-therma |  |  |  | 12: N/A; 11: *; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heat-therma-DIM |  |  | *
Heat-therma |  |  |  | 12: N/A; 11: *; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heat-therma-DIM
Heat-therma |  |  |  | 12: N/A; 11: *; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heat-therma-DIM |  |  | *
Heat-therma |  |  |  | 12: N/A; 11: * | [6 empty] | *
Heat-therma |  |  |  | 12: N/A; 11: * | [6 empty] | *
Heat-therma | [7 empty] | Top weight (*) | N/A | N/A | Top weight (*) - g | Trọng lượng | g | N/A | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Heat-therma-MASS |  |  | N/A - Top weight (*) - g
Heat-therma |  |  |  | 11: *; 14: g; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Heat-therma-MASS |  |  | *
Stirrer |  |  | LAB-T-01 LAB-TIME-02 | 5: Heating Magnetic Stirrer; 6: Magnetic Stirrer; 7: Máy khuấy từ; 9: Temperature (*); 11: Temperature (*) - ℃; 13: Nhiệt độ; 14: ℃; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Stirrer- |  |  | Temperature (*) - ℃
Stirrer |  |  |  | 9: Min …..........; 12: ±2; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Stirrer-DEG
Stirrer |  |  |  | 9: Middle …...............; 12: ±2; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Stirrer-DEG |  |  | 0
Stirrer |  |  |  | 9: Max …..........; 12: ±2; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Stirrer-DEG |  |  | 0
Stirrer |  |  |  | 11: *; 14: ℃; 17: 0; 18: 0; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Stirrer-DEG |  |  | *
Stirrer |  |  |  | 11: *; 14: ℃; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Stirrer-DEG
Stirrer | [7 empty] | Speed (*) | 1 | ±2 | Speed (*) - rpm |  | rpm | 1 | 2 | -1 | 3 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Stirrer-RPM
Stirrer |  |  |  | 10: 2; 12: ±2; 14: rpm; 15: 2; 16: 2; 17: 0; 18: 4; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Stirrer-RPM
Stirrer |  |  |  | 10: 3; 12: ±2; 14: rpm; 15: 3; 16: 2; 17: 1; 18: 5; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Stirrer-RPM
Stirrer |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Stirrer-
Stirrer |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Stirrer-
Stirrer |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Stirrer-
Heat-iron |  |  | LAB-T-01 LAB-TIME-02 | 5: Iron, Steam presss; 6: Iron; 9: Temperature (*); 11: Temperature (*); 13: Nhiệt độ; 21: #N/A |  | #N/A | #N/A | Heat-iron-
Heat-iron |  |  |  | 9: Minimum iron/ Synthetics; 10: Left; 12: N/A; 11: Minimum iron/ Synthetics; 14: ℃; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Heat-iron-DEG |  |  | Left - Minimum iron/ Synthetics
Heat-iron |  |  |  | 10: Center; 12: N/A; 14: ℃; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Heat-iron-DEG
Heat-iron |  |  |  | 10: Right; 12: N/A; 14: ℃; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Heat-iron-DEG
Heat-iron | [7 empty] | Woolens/ Silks | Left | N/A | Woolens/ Silks |  | ℃ | N/A | N/A | #VALUE! | #VALUE! | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Heat-iron-DEG
Heat-iron |  |  |  | 10: Center; 12: N/A; 14: ℃; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Heat-iron-DEG
Heat-iron |  |  |  | 10: Right; 12: N/A; 14: ℃; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Heat-iron-DEG
Heat-iron | [7 empty] | Linen/ Cotton | Left | N/A | Linen/ Cotton |  | ℃ | N/A | N/A | #VALUE! | #VALUE! | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Heat-iron-DEG
Heat-iron |  |  |  | 10: Center; 12: N/A; 14: ℃; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Heat-iron-DEG
Heat-iron |  |  |  | 10: Right; 12: N/A; 14: ℃; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Heat-iron-DEG
Heat-iron | [7 empty] | ….......................... | Left | N/A | * |  | ℃ | N/A | N/A | #VALUE! | #VALUE! | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Heat-iron-DEG
Heat-iron |  |  |  | 10: Center; 12: N/A; 14: ℃; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Heat-iron-DEG |  |  | Center
Heat-iron |  |  |  | 10: Right; 12: N/A; 14: ℃; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Heat-iron-DEG |  |  | Right
Heat-iron | [7 empty] | ….......................... | Left | N/A | * |  | ℃ | N/A | N/A | #VALUE! | #VALUE! | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Heat-iron-DEG
Heat-iron |  |  |  | 10: Center; 12: N/A; 14: ℃; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Heat-iron-DEG
Heat-iron |  |  |  | 10: Right; 12: N/A; 14: ℃; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Heat-iron-DEG
Heat-iron | [7 empty] | Speed (*) | 1 | ±2 | Speed (*) - rpm |  | rpm | 1 | 2 | -1 | 3 | RPM |  | 0.3 |  |  | 0.1 | 0.3 | Heat-iron-RPM
Heat-iron |  |  |  | 10: 2; 12: ±2; 14: rpm; 15: 2; 16: 2; 17: 0; 18: 4; 19: RPM; 21: 0.3 |  | 0.1 | 0.3 | Heat-iron-RPM
Heat-iron |  |  |  | 10: 3; 12: ±2; 14: rpm; 15: 3; 16: 2; 17: 1; 18: 5; 19: RPM; 21: 0.3 |  | 0.1 | 0.3 | Heat-iron-RPM
Heat-iron |  |  |  | 10: 4; 12: ±2; 14: rpm; 15: 4; 16: 2; 17: 2; 18: 6; 19: RPM; 21: 0.3 |  | 0.1 | 0.3 | Heat-iron-RPM
Heat-iron |  |  |  | 10: 5; 12: ±2; 14: rpm; 15: 5; 16: 2; 17: 3; 18: 7; 19: RPM; 21: 0.3 |  | 0.1 | 0.3 | Heat-iron-RPM
Heat-press |  | Straight Linear Fusing Press | LAB-T-01 LAB-TIME-02 | 5: Straight Linear Fusing Press; 6: Straight Linear Fusing Press; 7: Máy ép keo, Máy ép nhiệt; 9: Temperature (*); 11: Temperature (*) - ℃; 13: Nhiệt độ; 14: ℃; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Heat-press-
Heat-press |  |  |  | 9: Min …..........; 12: ±2; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Heat-press-DEG
Heat-press |  |  |  | 9: Middle …..................; 12: ±2; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Heat-press-DEG
Heat-press |  |  |  | 9: Max …............; 12: ±2; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Heat-press-DEG
Heat-press |  |  |  | 21: #N/A |  | #N/A | #N/A | Heat-press-
Heat-press |  |  |  | 9: Press (*); 11: Press (*) - kg/cm2; 13: Áp suất; 14: kg/cm2; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Heat-press-PRESS
Heat-press |  |  |  | 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Heat-press-PRESS
Heat-press |  |  |  | 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Heat-press-PRESS
Heat-press | [7 empty] | Time | 15 | N/A | Time - sec | Bộ đếm thời gian | sec | 15 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Heat-press-TIME
Heat-press |  |  |  | 10: 30; 12: N/A; 14: sec; 15: 30; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Heat-press-TIME
Heat-press |  |  |  | 10: 60; 12: N/A; 14: sec; 15: 60; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Heat-press-TIME
Heat-press |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Heat-press-
Heat-press |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Heat-press-
Heat-press |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Heat-press-
Steam |  |  | LAB-T-01 LAB-TIME-02 | Steam Press | Steam Press | Máy ép hơi |  | Temperature (*) |  | 1 - | Temperature (*) - ℃ | Nhiệt độ | ℃ |  |  | 0 | 0 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Steam-DEG
Steam |  |  |  | 12: 2 -; 14: ℃; 17: 0; 18: 0; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Steam-DEG
Steam |  |  |  | 12: 3 -; 14: ℃; 17: 0; 18: 0; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Steam-DEG
Crock | AATCC 8, ISO 105 | Colour Fastness to Rubbing | LAB-LN-03 LAB-M-01 LAB-TIME-02 | Crockmaster 670 | Crockmeter | Thiết bị thử độ bền màu ma sát | Stroke length Finger diameter Downward force on finger Counter | Stroke length | 104 | ±3 | Stroke length - mm | Độ dài hành trình | mm | 104 | 3 | 101 | 107 | DIM |  | 0.02 | 1000000 | Service & Calibration | 0.01 | 0.02 | Crock-DIM |  |  | 104 - Stroke length - mm
Crock | [7 empty] | Finger diameter | 16 | ± 0.1 / ± 0.3 | Finger diameter - mm | Đường kính đầu tròn ma sát | mm | 16 | 0.3 | 15.7 | 16.3 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Crock-DIM |  |  | 16 - Finger diameter - mm
Crock | [7 empty] | Downward force on finger (*) | 9N - 917g | ±20 | Downward force on finger (*) - g | Tải trọng | g | 917 | 20 | 897 | 937 | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Crock-MASS |  |  | 9N - 917g - Downward force on finger (*) - g
Crock | [27 empty] | * | bar: 11
Crock | [7 empty] | Counter | 10 | N/A | Counter - Cycle | Bộ đếm | Cycle | 10 | N/A | N/A | N/A | TIM1 |  | 0.02 |  |  | 0.001 | 0.02 | Crock-TIM1 |  |  | 10 - Counter - Cycle
Crock |  |  |  | 9: Speed; 10: 10; 12: N/A; 11: Speed - Cycle/10 secs; 13: Tốc độ; 14: Cycle/10 secs; 19: CYCLE/TIME; 21: 0.1 |  | 0.1 | 0.1 | Crock-CYCLE/TIME |  |  | 10 - Speed - Cycle/10 secs
Crock | [27 empty] | * | bar: 11
Crock-rotary | AATCC TM 116 | Rotary Crocking | LAB-LN-03 LAB-M-01 LAB-TIME-02 | Rotary Crockmeter | Rotary Crocking | Thiết bị thử độ bền màu ma sát | Stroke length Finger diameter Downward force on finger Counter | Finger diameter | 16 | ±0.3 | Finger diameter - mm | Đường kính đầu tròn ma sát | mm | 16 | 0.3 | 15.7 | 16.3 | DIM |  | 0.02 | 1000000 | Service & Calibration | 0.01 | 0.02 | Crock-rotary-DIM |  |  | 16 - Finger diameter - mm
Crock-rotary | [7 empty] | Downward force on finger (*) - N | 11.1 | ±1.11 | Downward force on finger (*) - N | Tải trọng | N | 11.1 | 1.11 | 9.99 | 12.21 | FORCES |  | 0.14 |  |  | 0.0001 | 0.14 | Crock-rotary-FORCES |  |  | 11.1 - Downward force on finger (*) - N
Crock-rotary | [7 empty] | Reciprocal turns - Cycle/min | 40 | N/A | Reciprocal turns - Cycle/min | Số vòng lắc | Cycle/min | 40 | N/A | #VALUE! | #VALUE! | TIM1 |  | 0.02 |  |  | 0.001 | 0.02 | Crock-rotary-TIM1 |  |  | 40 - Reciprocal turns - Cycle/min
Rub-fastness |  | Rub Fastness Tester | LAB-LN-03 LAB-M-01 LAB-TIME-02 | Rub Fastness Tester | Rub Fastness Tester | Thiết bị thử độ bền màu ma sát | Stroke length Finger diameter Downward force on finger Counter | Stroke length | 35 | ±3 | Stroke length - mm | Độ dài hành trình | mm | 35 | 3 | 32 | 38 | DIM |  | 0.02 | 1000000 | Service & Calibration | 0.01 | 0.02 | Rub-fastness-DIM |  |  | 35 - Stroke length - mm
Rub-fastness | [7 empty] | Speed (*) | 40 | ±2.5 | Speed (*) - rpm | Tốc độ | rpm | 40 | 2.5 | 37.5 | 42.5 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Rub-fastness-RPM |  |  | 40 - Speed (*) - rpm
Rub-fastness | [7 empty] | Counter | 100 | N/A | Counter - Cycle | Bộ đếm | Cycle | 1000 | N/A | N/A | N/A | TIM1 |  | 0.02 |  |  | 0.001 | 0.02 | Rub-fastness-TIM1 |  |  | 100 - Counter - Cycle
Rub-fastness | [7 empty] | Load (*) | 1000 - 1000 | ±20 | Load (*) - g | Tải trọng | g | 1000 | 20 | 980 | 1020 | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Rub-fastness-MASS |  |  | 1000 - 1000 - Load (*) - g
Rub-fastness |  |  |  | 10: 1000 - 1000; 12: ±20; 14: g; 15: 1000; 16: 20; 17: 980; 18: 1020; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Rub-fastness-MASS |  |  | 1000 - 1000
Rub-fastness |  |  |  | 10: 1000 - 1000; 12: ±20; 14: g; 15: 1000; 16: 20; 17: 980; 18: 1020; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Rub-fastness-MASS
Rub-fastness |  |  |  | 10: 1000 - 1000; 12: ±20; 14: g; 15: 1000; 16: 20; 17: 980; 18: 1020; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Rub-fastness-MASS
Rub-fastness |  |  |  | 10: 1000 - 1000; 12: ±20; 14: g; 15: 1000; 16: 20; 17: 980; 18: 1020; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Rub-fastness-MASS
Rub-fastness |  |  |  | 10: 1000 - 1000; 12: ±20; 14: g; 15: 1000; 16: 20; 17: 980; 18: 1020; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Rub-fastness-MASS
Rub-fastness |  |  |  | 10: 500 - 500; 12: ±20; 14: g; 15: 500; 16: 20; 17: 480; 18: 520; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Rub-fastness-MASS
Rub-fastness |  |  |  | 10: 500 - 500; 12: ±20; 14: g; 15: 500; 16: 20; 17: 480; 18: 520; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Rub-fastness-MASS
Rub-fastness |  |  |  | 10: 500 - 500; 12: ±20; 14: g; 15: 500; 16: 20; 17: 480; 18: 520; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Rub-fastness-MASS
Rub-fastness |  |  |  | 10: 500 - 500; 12: ±20; 14: g; 15: 500; 16: 20; 17: 480; 18: 520; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Rub-fastness-MASS |  |  | 500 - 500
Rub-fastness |  |  |  | 10: 500 - 500; 12: ±20; 14: g; 15: 500; 16: 20; 17: 480; 18: 520; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Rub-fastness-MASS |  |  | 500 - 500
Rub-fastness |  |  |  | 10: 500 - 500; 12: ±20; 14: g; 15: 500; 16: 20; 17: 480; 18: 520; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Rub-fastness-MASS |  |  | 500 - 500
Rub-fastness |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Rub-fastness- |  |  | *
Rub-fastness |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Rub-fastness- |  |  | *
Rub-fastness |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Rub-fastness-
Pill-jis |  |  | LAB-LN-03 LAB-M-01 LAB-TIME-02 | 5: ICI TYPE PILLING TESTER; 6: ICI TYPE PILLING TESTER; 7: Thiết bị thử độ xù lông; 9: Specimen Box Inner Dimension Inspection; 10: Box 1; 11: Specimen Box Inner Dimension Inspection - mm; 21: #N/A |  | #N/A | #N/A | Pill-jis-
Pill-jis |  |  |  | 10: W; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis |  |  |  | 10: D; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis |  |  |  | 10: H; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis |  |  |  | 10: Box 2; 12: ±5; 21: #N/A |  | #N/A | #N/A | Pill-jis-
Pill-jis |  |  |  | 10: W; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis |  |  |  | 10: D; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM |  |  | D
Pill-jis |  |  |  | 10: H; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM |  |  | H
Pill-jis |  |  |  | 10: Box 3; 12: ±5; 21: #N/A |  | #N/A | #N/A | Pill-jis- |  |  | Box 3
Pill-jis |  |  |  | 10: W; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM |  |  | W
Pill-jis |  |  |  | 10: D; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM |  |  | D
Pill-jis |  |  |  | 10: H; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM |  |  | H
Pill-jis |  |  |  | 10: Box 4; 12: ±5; 21: #N/A |  | #N/A | #N/A | Pill-jis- |  |  | Box 4
Pill-jis |  |  |  | 10: W; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM |  |  | W
Pill-jis |  |  |  | 10: D; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM |  |  | D
Pill-jis |  |  |  | 10: H; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM |  |  | H
Pill-jis |  |  |  | 10: Box 5; 12: ±5; 21: #N/A |  | #N/A | #N/A | Pill-jis- |  |  | Box 5
Pill-jis |  |  |  | 10: W; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM |  |  | W
Pill-jis |  |  |  | 10: D; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis |  |  |  | 10: H; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis |  |  |  | 10: Box 6; 12: ±5; 21: #N/A |  | #N/A | #N/A | Pill-jis- |  |  | Box 6
Pill-jis |  |  |  | 10: W; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM |  |  | W
Pill-jis |  |  |  | 10: D; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis |  |  |  | 10: H; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis |  |  |  | 10: Box 7; 12: ±5; 21: #N/A |  | #N/A | #N/A | Pill-jis-
Pill-jis |  |  |  | 10: W; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis |  |  |  | 10: D; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis |  |  |  | 10: H; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis |  |  |  | 10: Box 8; 12: ±5; 21: #N/A |  | #N/A | #N/A | Pill-jis- |  |  | Box 8
Pill-jis |  |  |  | 10: W; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis |  |  |  | 10: D; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis |  |  |  | 10: H; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis |  |  |  | 10: Box 9; 12: ±5; 21: #N/A |  | #N/A | #N/A | Pill-jis-
Pill-jis |  |  |  | 10: W; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis |  |  |  | 10: D; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis |  |  |  | 10: H; 12: ±5; 14: mm; 15: 230; 16: 5; 17: 225; 18: 235; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-jis-DIM
Pill-jis | [7 empty] | Rubber Tube Weight Inspection | 51 | ±1 | Rubber Tube Weight Inspection - g |  | g | 51 | 1 | 50 | 52 | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Pill-jis-MASS |  |  | 51 - Rubber Tube Weight Inspection - g
Pill-jis |  |  |  | 10: 51; 12: ±1; 14: g; 15: 51; 16: 1; 17: 50; 18: 52; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Pill-jis-MASS |  |  | 51
Pill-jis |  |  |  | 10: 51; 12: ±1; 14: g; 15: 51; 16: 1; 17: 50; 18: 52; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Pill-jis-MASS |  |  | 51
Pill-jis |  |  |  | 10: 51; 12: ±1; 14: g; 15: 51; 16: 1; 17: 50; 18: 52; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Pill-jis-MASS
Pill-jis |  |  |  | 10: 51; 12: ±1; 14: g; 15: 51; 16: 1; 17: 50; 18: 52; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Pill-jis-MASS
Pill-jis |  |  |  | 10: 51; 12: ±1; 14: g; 15: 51; 16: 1; 17: 50; 18: 52; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Pill-jis-MASS
Pill-jis |  |  |  | 10: 51; 12: ±1; 14: g; 15: 51; 16: 1; 17: 50; 18: 52; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Pill-jis-MASS
Pill-jis |  |  |  | 10: 51; 12: ±1; 14: g; 15: 51; 16: 1; 17: 50; 18: 52; 19: MASS; 21: 0.07 |  | 0 | 0.07 | Pill-jis-MASS
Pill-jis |  |  |  | 10: 51; 12: ±1; 14: g; 15: 51; 16: 1; 17: 50; 18: 52; 19: MASS; 21: 0.07 |  | 0 | 0.07 | Pill-jis-MASS
Pill-jis |  |  |  | 10: 51; 12: ±1; 14: g; 15: 51; 16: 1; 17: 50; 18: 52; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Pill-jis-MASS
Pill-jis |  |  |  | 10: 51; 12: ±1; 14: g; 15: 51; 16: 1; 17: 50; 18: 52; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Pill-jis-MASS
Pill-jis |  |  |  | 10: 51; 12: ±1; 14: g; 15: 51; 16: 1; 17: 50; 18: 52; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Pill-jis-MASS
Pill-jis |  |  |  | 10: 51; 12: ±1; 14: g; 15: 51; 16: 1; 17: 50; 18: 52; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Pill-jis-MASS
Pill-jis |  |  |  | 10: 51; 12: ±1; 14: g; 15: 51; 16: 1; 17: 50; 18: 52; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Pill-jis-MASS
Pill-jis |  |  |  | 10: 51; 12: ±1; 14: g; 15: 51; 16: 1; 17: 50; 18: 52; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Pill-jis-MASS
Pill-jis |  |  |  | 10: 51; 12: ±1; 14: g; 15: 51; 16: 1; 17: 50; 18: 52; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Pill-jis-MASS
Pill-jis |  |  |  | 11: *; 14: g; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Pill-jis-MASS
Pill-jis |  |  |  | 11: *; 14: g; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Pill-jis-MASS
Pill-jis | [7 empty] | Specimen Box Rotation Speed Inspection (*) | 60 | ±2 | Specimen Box Rotation Speed Inspection (*) - rpm |  | rpm | 60 | 2 | 58 | 62 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Pill-jis-RPM
Pill-jis |  |  |  | 10: 60; 12: ±2; 14: rpm; 15: 60; 16: 2; 17: 58; 18: 62; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Pill-jis-RPM
Pill-jis |  |  |  | 10: 60; 12: ±2; 14: rpm; 15: 60; 16: 2; 17: 58; 18: 62; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Pill-jis-RPM
Pill-jis |  |  |  | 11: *; 14: rpm; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Pill-jis-RPM
Pill-jis | [7 empty] | Timer Measurement Inspection | 3600 | ±12 | Timer Measurement Inspection - sec |  | sec | 3600 | 12 | 3588 | 3612 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Pill-jis-TIME |  |  | 3600 - Timer Measurement Inspection - sec
Pill-jis |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Pill-jis- |  |  | *
Pill-jis |  |  |  | 9: 1…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Pill-jis- |  |  | *
Pill-jis |  |  |  | 9: 2…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Pill-jis-
Pill-jis |  |  |  | 9: 3…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Pill-jis-
Pill-jis |  |  |  | 9: 4…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Pill-jis-
Pill-jis |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Pill-jis-
Elastic-band |  |  | LAB-T-01 LAB-TIME-02 | HBI Flex Tester | HBI Flex Tester |  |  | Stroke | 2.5 | N/A | Stroke - inch | Hành trình | inch | 2.5 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 | 1000000 | Service & Calibration | 0.01 | 0.02 | Elastic-band-DIM
Elastic-band |  | HBI Flex Tester | TAV-QTHTD |  |  |  |  | Speed (*) | 20 | ±1 | Speed (*) - rpm | Tốc độ | rpm | 20 | 1 | 19 | 21 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Elastic-band-RPM
Elastic-band | [7 empty] | Cycle | 20 | ±1 | Cycle - cycle/min | Chu kỳ | cycle/min | 20 | 1 | 19 | 21 | TIM1 |  | 0.02 |  |  | 0.001 | 0.02 | Elastic-band-TIM1
Elastic-band |  |  |  | 10: 100; 12: ±2; 14: cycle/min; 15: 100; 16: 2; 17: 98; 18: 102; 19: TIM1; 21: 0.02 |  | 0.001 | 0.02 | Elastic-band-TIM1
Elastic-band | [7 empty] | Elongation | 5 | N/A | Elongation - inch |  | inch | 5 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Elastic-band-DIM |  |  | 5 - Elongation - inch
Elastic-band |  |  |  | 10: 7.5; 12: N/A; 14: inch; 15: 7.5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Elastic-band-DIM |  |  | 7.5
Elastic-band |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Elastic-band- |  |  | *
Elastic-band |  |  |  | 9: 1…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Elastic-band- |  |  | *
Elastic-band |  |  |  | 9: 2…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Elastic-band-
Elastic-band |  |  |  | 9: 3…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Elastic-band- |  |  | *
Elastic-band |  |  |  | 9: 4…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Elastic-band- |  |  | *
Elastic-band |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Elastic-band- |  |  | *
Cut-blade |  | Circular Blade Cut | LAB-LN-03 | Circular Blade Cut | Circular Blade Cut |  | Stroke length Finger diameter Downward force on finger Counter | Stroke length - mm | 50 | N/A | Stroke length - mm | Độ dài hành trình | mm | 50 | N/A | #VALUE! | #VALUE! | DIM |  | 0.07 | 1000000 | Service & Calibration | 0.01 | 0.02 | Cut-blade-DIM |  |  | 50 - Stroke length - mm
Cut-blade | [7 empty] | Blade diameter - mm | 45 | ±0.5 | Blade diameter - mm | Đường kính dao | mm | 45 | 0.5 | 44.5 | 45.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-blade-DIM |  |  | 45 - Blade diameter - mm
Cut-blade | [7 empty] | Blade thickness - mm | 0.3 | ±0.3 | Blade thickness - mm | Bề dày dao | mm | 0.3 | 0.3 | 0 | 0.6 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-blade-DIM |  |  | 0.3 - Blade thickness - mm
Cut-blade | [7 empty] | Downward force (*) - N | 5 | ±0.005 | Downward force (*) - N | Tải trọng | N | 5 | 0.005 | 4.995 | 5.005 | FORCES |  | 0.14 |  |  | 0.0001 | 0.14 | Cut-blade-FORCES |  |  | 5 - Downward force (*) - N
Cut-blade | [7 empty] | Counter - lần/min | N/A | N/A | Counter - lần/min | Bộ đếm | lần/min | N/A | N/A | #VALUE! | #VALUE! | TIM1 |  | 0.02 |  |  | 0.001 | 0.02 | Cut-blade-TIM1
Cut-blade | [7 empty] | Cutting Speed - m/min | N/A | N/A | Cutting Speed - m/min | Tốc độ cắt | m | N/A | N/A | #VALUE! | #VALUE! | DIM/TIME |  | 0.02 |  |  | 0.01 | 0.02 | Cut-blade-DIM/TIME
Cut-blade | [7 empty] | Speed (*) - rpm | N/A | N/A | Speed (*) - rpm | Tốc độ | rpm | N/A | N/A | #VALUE! | #VALUE! | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Cut-blade-RPM
Cut-blade |  |  |  | 11: *; 17: 0; 18: 0 |  | #N/A | #N/A | Cut-blade-
Cut-blade |  |  |  | 11: *; 17: 0; 18: 0 |  | #N/A | #N/A | Cut-blade-
Flex-bally-adidas | Adidas GE-24 | Cyclic fold durability, of leathers and their top layers | LAB-LN-03 LAB-M-01 LAB-F01:2023 LAB-TIME-02 LAB-A-01 | Bally Flexometer | Flexometer test | Thiết bị thử uốn gấp | Flex counter Flexing angle Clamp distance | Counter | 100 | N/A | Counter - time | Số lần gập | time | 100 | N/A | #VALUE! | #VALUE! | TIM1 |  | 0.02 | 1500000 | Service & Calibration | 0.001 | 0.02 | Flex-bally-adidas-TIM1 |  |  | 100 - Counter - time
Flex-bally-adidas | [7 empty] | Speed (*) | 100 | ±5 | Speed (*) - rpm | Tốc độ | rpm | 100 | 5 | 95 | 105 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Flex-bally-adidas-RPM |  |  | 100 - Speed (*) - rpm
Flex-bally-adidas | [7 empty] | Flexing angle 0 ⁰ | 1 | ±1 | Flexing angle 0 ⁰ - ⁰ | Góc | ⁰ | 0 | 1 | -1 | 1 | ANG |  | 0.03 |  |  | 0.05 | 0.03 | Flex-bally-adidas-ANG |  |  | 1 - Flexing angle 0 ⁰ - ⁰
Flex-bally-adidas |  |  |  | 10: 2; 12: ±1; 14: ⁰; 15: 0; 16: 1; 17: -1; 18: 1; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG |  |  | 2
Flex-bally-adidas |  |  |  | 10: 3; 12: ±1; 14: ⁰; 15: 0; 16: 1; 17: -1; 18: 1; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG |  |  | 3
Flex-bally-adidas |  |  |  | 10: 4; 12: ±1; 14: ⁰; 15: 0; 16: 1; 17: -1; 18: 1; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG |  |  | 4
Flex-bally-adidas |  |  |  | 10: 5; 12: ±1; 14: ⁰; 15: 0; 16: 1; 17: -1; 18: 1; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG |  |  | 5
Flex-bally-adidas |  |  |  | 10: 6; 12: ±1; 14: ⁰; 15: 0; 16: 1; 17: -1; 18: 1; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG
Flex-bally-adidas |  |  |  | 10: 7; 12: ±1; 14: ⁰; 15: 0; 16: 1; 17: -1; 18: 1; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG
Flex-bally-adidas |  |  |  | 10: 8; 12: ±1; 14: ⁰; 15: 0; 16: 1; 17: -1; 18: 1; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG
Flex-bally-adidas |  |  |  | 10: 9; 12: ±1; 14: ⁰; 15: 0; 16: 1; 17: -1; 18: 1; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG
Flex-bally-adidas |  |  |  | 10: 10; 12: ±1; 14: ⁰; 15: 0; 16: 1; 17: -1; 18: 1; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG
Flex-bally-adidas |  |  |  | 10: 11; 12: ±1; 14: ⁰; 15: 0; 16: 1; 17: -1; 18: 1; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG
Flex-bally-adidas |  |  |  | 10: 12; 12: ±1; 14: ⁰; 15: 0; 16: 1; 17: -1; 18: 1; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG
Flex-bally-adidas | [7 empty] | Flexing angle 22.5° | 1 | ±1 | Flexing angle 22.5° - ⁰ | Góc | ⁰ | 22.5 | 1 | 21.5 | 23.5 | ANG |  | 0.03 |  |  | 0.05 | 0.03 | Flex-bally-adidas-ANG
Flex-bally-adidas |  |  |  | 10: 2; 12: ±1; 14: ⁰; 15: 22.5; 16: 1; 17: 21.5; 18: 23.5; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG
Flex-bally-adidas |  |  |  | 10: 3; 12: ±1; 14: ⁰; 15: 22.5; 16: 1; 17: 21.5; 18: 23.5; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG
Flex-bally-adidas |  |  |  | 10: 4; 12: ±1; 14: ⁰; 15: 22.5; 16: 1; 17: 21.5; 18: 23.5; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG |  |  | 4
Flex-bally-adidas |  |  |  | 10: 5; 12: ±1; 14: ⁰; 15: 22.5; 16: 1; 17: 21.5; 18: 23.5; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG |  |  | 5
Flex-bally-adidas |  |  |  | 10: 6; 12: ±1; 14: ⁰; 15: 22.5; 16: 1; 17: 21.5; 18: 23.5; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG |  |  | 6
Flex-bally-adidas |  |  |  | 10: 7; 12: ±1; 14: ⁰; 15: 22.5; 16: 1; 17: 21.5; 18: 23.5; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG |  |  | 7
Flex-bally-adidas |  |  |  | 10: 8; 12: ±1; 14: ⁰; 15: 22.5; 16: 1; 17: 21.5; 18: 23.5; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG
Flex-bally-adidas |  |  |  | 10: 9; 12: ±1; 14: ⁰; 15: 22.5; 16: 1; 17: 21.5; 18: 23.5; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG
Flex-bally-adidas |  |  |  | 10: 10; 12: ±1; 14: ⁰; 15: 22.5; 16: 1; 17: 21.5; 18: 23.5; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG
Flex-bally-adidas |  |  |  | 10: 11; 12: ±1; 14: ⁰; 15: 22.5; 16: 1; 17: 21.5; 18: 23.5; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG
Flex-bally-adidas |  |  |  | 10: 12; 12: ±1; 14: ⁰; 15: 22.5; 16: 1; 17: 21.5; 18: 23.5; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-bally-adidas-ANG
Flex-bally-adidas | [7 empty] | Clamp distance | 1 | ±1 | Clamp distance - mm | Khoảng cách | mm | 25 | 1 | 24 | 26 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Flex-bally-adidas-DIM |  |  | 1 - Clamp distance - mm
Flex-bally-adidas |  |  |  | 10: 2; 12: ±1; 14: mm; 15: 25; 16: 1; 17: 24; 18: 26; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flex-bally-adidas-DIM |  |  | 2
Flex-bally-adidas |  |  |  | 10: 3; 12: ±1; 14: mm; 15: 25; 16: 1; 17: 24; 18: 26; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flex-bally-adidas-DIM |  |  | 3
Flex-bally-adidas |  |  |  | 10: 4; 12: ±1; 14: mm; 15: 25; 16: 1; 17: 24; 18: 26; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flex-bally-adidas-DIM
Flex-bally-adidas |  |  |  | 10: 5; 12: ±1; 14: mm; 15: 25; 16: 1; 17: 24; 18: 26; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flex-bally-adidas-DIM
Flex-bally-adidas |  |  |  | 10: 6; 12: ±1; 14: mm; 15: 25; 16: 1; 17: 24; 18: 26; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flex-bally-adidas-DIM |  |  | 6
Flex-bally-adidas |  |  |  | 10: 7; 12: ±1; 14: mm; 15: 25; 16: 1; 17: 24; 18: 26; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flex-bally-adidas-DIM |  |  | 7
Flex-bally-adidas |  |  |  | 10: 8; 12: ±1; 14: mm; 15: 25; 16: 1; 17: 24; 18: 26; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flex-bally-adidas-DIM
Flex-bally-adidas |  |  |  | 10: 9; 12: ±1; 14: mm; 15: 25; 16: 1; 17: 24; 18: 26; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flex-bally-adidas-DIM
Flex-bally-adidas |  |  |  | 10: 10; 12: ±1; 14: mm; 15: 25; 16: 1; 17: 24; 18: 26; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flex-bally-adidas-DIM
Flex-bally-adidas |  |  |  | 10: 11; 12: ±1; 14: mm; 15: 25; 16: 1; 17: 24; 18: 26; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flex-bally-adidas-DIM
Flex-bally-adidas |  |  |  | 10: 12; 12: ±1; 14: mm; 15: 25; 16: 1; 17: 24; 18: 26; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flex-bally-adidas-DIM |  |  | 12
Flex-bally-adidas | bar: 11
Flex-bally-adidas | bar: 11
Flex-bally-adidas | bar: 11
Flex-bally-satra55 | SATRA TM55 ISO 5402 ISO 17694 GB/T 20991 |  | LAB-LN-03 LAB-M-01 LAB-F01:2023 LAB-TIME-02 LAB-A-01 | Bally Flexometer | Bally Flexometer |  | Distance from the back of clamp to intersection with line CF Distance GJ DistanceEF to lower clamp measurement AD Radius of clamping plate corner Distance J-EF Perpendicular distance between poin C and line AD Angle adjacent to distance CF and AD Flexing angle Flexing speed | Distance from the back of clamp to intersection with line CF | 55 | ±2 | Distance from the back of clamp to intersection with line CF - mm | 55 | mm | 55 | 2 | 53 | 57 | DIM |  | 0.02 | 1200000 | Service & Calibration | 0.01 | 0.02 | Flex-bally-satra55-DIM
Flex-bally-satra55 | [7 empty] | Distance GJ | 15 | ±1 | Distance GJ - mm | 15 | mm | 15 | 1 | 14 | 16 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Flex-bally-satra55-DIM
Flex-bally-satra55 | [7 empty] | DistanceEF to lower clamp measurement AD | 4 | ±0.2 | DistanceEF to lower clamp measurement AD - mm | 4 | mm | 4 | 0.2 | 3.8 | 4.2 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Flex-bally-satra55-DIM |  |  | 4 - DistanceEF to lower clamp measurement AD - mm
Flex-bally-satra55 | [7 empty] | Radius of clamping plate corner | 2 | ±0.2 | Radius of clamping plate corner - mm | 2 | mm | 2 | 0.2 | 1.8 | 2.2 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Flex-bally-satra55-DIM |  |  | 2 - Radius of clamping plate corner - mm
Flex-bally-satra55 | [7 empty] | Distance J-EF | 50 | ±2 | Distance J-EF - mm | 50 | mm | 50 | 2 | 48 | 52 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Flex-bally-satra55-DIM |  |  | 50 - Distance J-EF - mm
Flex-bally-satra55 | [7 empty] | Perpendicular distance between poin C and line AD | 25 | ±1 | Perpendicular distance between poin C and line AD - mm | 25 | mm | 25 | 1 | 24 | 26 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Flex-bally-satra55-DIM |  |  | 25 - Perpendicular distance between poin C and line AD - mm
Flex-bally-satra55 | [7 empty] | Angle adjacent to distance CF and AD | 45 | ±5 | Angle adjacent to distance CF and AD - ⁰ | 45 | ⁰ | 45 | 5 | 40 | 50 | ANG |  | 0.03 |  |  | 0.05 | 0.03 | Flex-bally-satra55-ANG |  |  | 45 - Angle adjacent to distance CF and AD - ⁰
Flex-bally-satra55 | [7 empty] | Flexing angle | 22.5 | ±0.5 | Flexing angle - ⁰ | 22.5 | ⁰ | 22.5 | 0.5 | 22 | 23 | ANG |  | 0.03 |  |  | 0.05 | 0.03 | Flex-bally-satra55-ANG
Flex-bally-satra55 | [7 empty] | Flexing speed (*) | 100 | ±5 | Flexing speed (*) - rpm | 100 | rpm | 100 | 5 | 95 | 105 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Flex-bally-satra55-RPM
Flex-bally-satra55 | bar: 11
Flex-bally-satra55 | [27 empty] | * | bar: 11
Flex-bally-satra55 | [27 empty] | * | bar: 11
Flex-ross | SATRA TM60 | Ross Flexing | LAB-LN-03 LAB-M-01 LAB-F01:2023 LAB-TIME-02 LAB-A-01 | Ross Flexing Machine | Ross Flexing Machine | Thiết bị thử uốn gấp | Diameter of Flexing mandrel Flexing angle Flexing speed Frezzer compartment | Diameter of Flexing mandrel | 10 | ±0.6 | Diameter of Flexing mandrel - mm | Đường kính | mm | 10 | 0.6 | 9.4 | 10.6 | DIM |  | 0.02 | 800000 | Service & Calibration | 0.01 | 0.02 | Flex-ross-DIM |  |  | 10 - Diameter of Flexing mandrel - mm
Flex-ross | [7 empty] | Flexing angle | 90 | ±2 | Flexing angle - ⁰ | Góc | ⁰ | 90 | 2 | 88 | 92 | ANG |  | 0.03 |  |  | 0.05 | 0.03 | Flex-ross-ANG |  |  | 90 - Flexing angle - ⁰
Flex-ross | [7 empty] | Flexing speed (*) | 60 | ±5 | Flexing speed (*) - rpm | Tốc độ | rpm | 60 | 5 | 55 | 65 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Flex-ross-RPM |  |  | 60 - Flexing speed (*) - rpm
Flex-ross | [7 empty] | Frezzer compartment (*) | -5 | ±2 | Frezzer compartment (*) - ⁰C | Nhiệt độ | ℃ | -5 | 2 | -7 | -3 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Flex-ross-DEG |  |  | -5 - Frezzer compartment (*) - ⁰C
Flex-ross |  |  |  | 10: -10; 12: ±1; 14: ℃; 15: -10; 16: 1; 17: -11; 18: -9; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Flex-ross-DEG |  |  | -10
Flex-ross |  |  |  | 10: 20; 12: ±2; 14: ℃; 15: 20; 16: 2; 17: 18; 18: 22; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Flex-ross-DEG
Flex-ross | [7 empty] | Bally - Counter | 100 | ±5 | Bally - Counter - time |  | time | 100 | 5 | 95 | 105 | TIM1 |  | 0.02 |  |  | 0.001 | 0.02 | Flex-ross-TIM1
Flex-ross |  |  |  | 9: Bally - Angle; 10: 0; 12: ±1; 11: Bally - Angle - ⁰; 14: ⁰; 17: 0; 18: 0; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-ross-ANG
Flex-ross |  |  |  | 9: Bally - Angle; 10: 22.5; 12: ±1; 11: Bally - Angle - ⁰; 14: ⁰; 17: 0; 18: 0; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-ross-ANG
Flex-ross |  |  |  | 9: Bally - Angle; 10: 60; 12: ±1; 11: Bally - Angle - ⁰; 14: ⁰; 17: 0; 18: 0; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-ross-ANG |  |  | 60 - Bally - Angle - ⁰
Flex-ross |  |  |  | 9: Bally - Clamp distance; 10: 25; 12: ±1; 11: Bally - Clamp distance - mm; 14: mm; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flex-ross-DIM |  |  | 25 - Bally - Clamp distance - mm
Flex-ross | [7 empty] | Ross - Counter | 100 | ±5 | Ross - Counter - time |  | time | 100 | 5 | 95 | 105 | TIM1 |  | 0.02 |  |  | 0.001 | 0.02 | Flex-ross-TIM1 |  |  | 100 - Ross - Counter - time
Flex-ross |  |  |  | 9: Ross - Angle; 10: 60; 12: ±1; 11: Ross - Angle - ⁰; 14: ⁰; 17: 0; 18: 0; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Flex-ross-ANG |  |  | 60 - Ross - Angle - ⁰
Flex-ross | bar: 11
Flex-ross | bar: 11
Flex-ross | bar: 11
Flex-ross | [27 empty] | * | bar: 11
Flex-shoe | SATRA TM161 | Whole Shoe Flexing | LAB-LN-03 LAB-M-01 LAB-F01:2023 LAB-TIME-02 LAB-A-01 | Whole Shoe Flexing Machine | Bally Flexing, Shoe Flexing, Upper Materials Flexing, Outsole Flexing, Whole sole flexing machine | Thiết bị thử uốn gấp | Flexing angle Dimensions of flexing mandrels Length of clamps Distance between clamps Flexing speed | Flexing angle | 90 | ±2 | Flexing angle - ⁰ | Góc | ⁰ | 90 | 2 | 88 | 92 | ANG |  | 0.23 | 1200000 | Service & Calibration | 0.1 | 0.23 | Flex-shoe-ANG |  |  | 90 - Flexing angle - ⁰
Flex-shoe |  |  |  | 10: 10; 12: ±1; 13: 10; 14: ⁰; 15: 10; 16: 1; 17: 9; 18: 11; 19: ANG; 21: 0.23 |  | 0.1 | 0.23 | Flex-shoe-ANG |  |  | 10
Flex-shoe |  |  |  | 10: 15; 12: ±1; 13: 15; 14: ⁰; 15: 15; 16: 1; 17: 14; 18: 16; 19: ANG; 21: 0.23 |  | 0.1 | 0.23 | Flex-shoe-ANG |  |  | 15
Flex-shoe |  |  |  | 10: 20; 12: ±1; 13: 20; 14: ⁰; 15: 20; 16: 1; 17: 19; 18: 21; 19: ANG; 21: 0.23 |  | 0.1 | 0.23 | Flex-shoe-ANG |  |  | 20
Flex-shoe |  |  |  | 10: 25; 12: ±1; 13: 25; 14: ⁰; 15: 25; 16: 1; 17: 24; 18: 26; 19: ANG; 21: 0.23 |  | 0.1 | 0.23 | Flex-shoe-ANG |  |  | 25
Flex-shoe |  |  |  | 10: 30; 12: ±1; 13: 30; 14: ⁰; 15: 30; 16: 1; 17: 29; 18: 31; 19: ANG; 21: 0.23 |  | 0.1 | 0.23 | Flex-shoe-ANG |  |  | 30
Flex-shoe |  |  |  | 10: 35; 12: ±1; 13: 35; 14: ⁰; 15: 35; 16: 1; 17: 34; 18: 36; 19: ANG; 21: 0.23 |  | 0.1 | 0.23 | Flex-shoe-ANG
Flex-shoe |  |  |  | 10: 45; 12: ±1; 13: 45; 14: ⁰; 15: 45; 16: 1; 17: 44; 18: 46; 19: ANG; 21: 0.23 |  | 0.1 | 0.23 | Flex-shoe-ANG
Flex-shoe |  |  |  |  | Bally Flexing, Shoe Flexing, Upper Materials Flexing, Outsole Flexing, Whole sole flexing machine | Thiết bị thử uốn gấp |  | Dimensions of flexing mandrels | Diameter 30 | ±1 | Dimensions of flexing mandrels - mm | Đường kính | mm | 30 | 1 | 29 | 31 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-shoe-DIM
Flex-shoe |  |  |  |  | Bally Flexing, Shoe Flexing, Upper Materials Flexing, Outsole Flexing, Whole sole flexing machine | Thiết bị thử uốn gấp |  |  | Length 150 | ±1 |  | Chiều dài | mm | 150 | 1 | 149 | 151 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-shoe-DIM |  |  | Length 150
Flex-shoe |  |  |  |  | Bally Flexing, Shoe Flexing, Upper Materials Flexing, Outsole Flexing, Whole sole flexing machine | Thiết bị thử uốn gấp |  | Length of clamps | 155 | ±1 | Length of clamps - mm | Chiều dài ngàm | mm | 155 | 1 | 154 | 156 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-shoe-DIM |  |  | 155 - Length of clamps - mm
Flex-shoe |  |  |  |  | Bally Flexing, Shoe Flexing, Upper Materials Flexing, Outsole Flexing, Whole sole flexing machine | Thiết bị thử uốn gấp |  | Distance between clamps | D1* : 110 | ±1 | Distance between clamps - mm | Khoảng cách giữa các ngàm | mm | 110 | 1 | 109 | 111 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-shoe-DIM |  |  | D1* : 110 - Distance between clamps - mm
Flex-shoe |  |  |  |  | Bally Flexing, Shoe Flexing, Upper Materials Flexing, Outsole Flexing, Whole sole flexing machine | Thiết bị thử uốn gấp |  |  | D2* : 41.5 | ±0.5 |  |  | mm | 41.5 | 0.5 | 41 | 42 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-shoe-DIM |  |  | D2* : 41.5
Flex-shoe |  |  |  |  | Bally Flexing, Shoe Flexing, Upper Materials Flexing, Outsole Flexing, Whole sole flexing machine | Thiết bị thử uốn gấp |  | Flexing speed (*) | 125-150 | 125-150 | Flexing speed (*) - rpm | Tốc độ | rpm | 137.5 | 12.5 | 125 | 150 | RPM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-shoe-RPM
Flex-shoe | [27 empty] | * | bar: 11
Flex-shoe | [27 empty] | * | bar: 11
Flex-shoe | [27 empty] | * | bar: 11
Flex-shoe | [27 empty] | * | bar: 11
Flex-upper | Adidas SATRA TM25 ISO 4643 ISO 5423 | Upper Material Flexing | LAB-LN-03 LAB-M-01 LAB-F01:2023 LAB-TIME-02 LAB-A-01 | Upper Material Flexing | Upper Material Flexing |  | Speed Stroke length Angle Counter | Speed (*) | 300 | ±30 | Speed (*) - rpm |  | rpm | 300 | 30 | 270 | 330 | RPM |  | 0.1 | 800000 | Service & Calibration | 0.1 | 0.1 | Flex-upper-RPM |  |  | 300 - Speed (*) - rpm
Flex-upper | [7 empty] | Stroke length | 19 | N/A | Stroke length - mm |  | mm | 19 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Flex-upper-DIM
Flex-upper | [7 empty] | Angle - ⁰ | 40 | N/A | Angle - ⁰ |  | ⁰ | 40 | N/A | #VALUE! | #VALUE! | ANG |  | 0.03 |  |  | 0.05 | 0.03 | Flex-upper-ANG
Flex-upper | [7 empty] | Counter | 100 | ±5 | Counter - cycle |  | cycle | 100 | 5 | 95 | 105 | TIM1 |  | 0.02 |  |  | 0.001 | 0.02 | Flex-upper-TIM1
Flex-upper | [27 empty] | * | bar: 11
Flex-upper | [27 empty] | * | bar: 11
Flex-upper | [27 empty] | * | bar: 11
Flex-upper | [27 empty] | * | bar: 11
Flex-upper | [27 empty] | * | bar: 11
Flex-outsole |  |  | LAB-LN-03 LAB-M-01 LAB-F01:2023 LAB-TIME-02 LAB-A-01 | Outsole Flexing Tester | Outsole Flexing Tester |  | Test station Distance between clamps Travel distance of movable clamp Speed Flexing angle Stroke length Counter | Test station | 110 | ±1 | Test station - mm |  | mm | 110 | 1 | 109 | 111 | DIM |  | 0.23 | 800000 | Service & Calibration | 0.1 | 0.23 | Flex-outsole-DIM |  |  | 110 - Test station - mm
Flex-outsole | [7 empty] | Distance between clamps | 30 | ±0.1 | Distance between clamps - mm | Khoảng cách | mm | 30 | 0.1 | 29.9 | 30.1 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-outsole-DIM |  |  | 30 - Distance between clamps - mm
Flex-outsole | [7 empty] | Travel distance of movable clamp | 41.5 | ±0.5 | Travel distance of movable clamp - mm |  | mm | 41.5 | 0.5 | 41 | 42 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-outsole-DIM |  |  | 41.5 - Travel distance of movable clamp - mm
Flex-outsole | [7 empty] | Speed (*) | 125 | N/A | Speed (*) - rpm | Tốc độ | rpm | 125 | N/A | #VALUE! | #VALUE! | RPM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-outsole-RPM |  |  | 125 - Speed (*) - rpm
Flex-outsole |  |  |  | 10: 135; 12: N/A; 14: rpm; 15: 135; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RPM; 21: 0.23 |  | 0.1 | 0.23 | Flex-outsole-RPM |  |  | 135
Flex-outsole |  |  |  | 10: 140; 12: N/A; 14: rpm; 15: 140; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RPM; 21: 0.23 |  | 0.1 | 0.23 | Flex-outsole-RPM |  |  | 140
Flex-outsole | [7 empty] | Flexing angle | 90 | N/A | Flexing angle - ⁰ | Góc | ⁰ | 90 | N/A | #VALUE! | #VALUE! | ANG |  | 0.23 |  |  | 0.1 | 0.23 | Flex-outsole-ANG |  |  | 90 - Flexing angle - ⁰
Flex-outsole | [7 empty] | Flexing angle | Left | N/A | Flexing angle - ⁰ |  | ⁰ | N/A | N/A | #VALUE! | #VALUE! | ANG |  | 0.23 |  |  | 0.1 | 0.23 | Flex-outsole-ANG
Flex-outsole |  |  |  | 10: Middle; 12: N/A; 14: ⁰; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: ANG; 21: 0.23 |  | 0.1 | 0.23 | Flex-outsole-ANG
Flex-outsole |  |  |  | 10: Right; 12: N/A; 14: ⁰; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: ANG; 21: 0.23 |  | 0.1 | 0.23 | Flex-outsole-ANG
Flex-outsole | [7 empty] | Stroke length | N/A | N/A | Stroke length - mm | Hành trình | mm | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-outsole-DIM |  |  | N/A - Stroke length - mm
Flex-outsole | [7 empty] | Counter |  | N/A | Counter - cycle | Bộ đém | cycle | N/A | N/A | #VALUE! | #VALUE! | TIM1 |  | 0.23 |  |  | 0.1 | 0.23 | Flex-outsole-TIM1 |  |  | Counter - cycle
Flex-outsole | bar: 11
Flex-outsole | bar: 11
Flex-outsole | bar: 11
Flex-outsole | bar: 11
Insock | Adidas ISO 20344, ISO 22649, EN 12746 GB/T 20991 | INSOCK ABSORPTION AND DESORPTION TESTER | LAB-LN-03 LAB-M-01 LAB-F01:2023 LAB-TIME-02 LAB-A-01 | INSOCK ABSORPTION AND DESORPTION TESTER |  |  | Diameter Compression force Stroke length Speed Counter | Diameter - mm | 120 | N/A | Diameter - mm | Đường kính roller | mm |  |  | 0 | 0 | DIM |  | 0.23 | 800000 | Service & Calibration | 0.1 | 0.23 | Insock-DIM |  |  | 120 - Diameter - mm
Insock |  |  |  | 10: 50; 12: ±5; 13: Bề rộng roller; 14: dim; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Insock-DIM |  |  | 50
Insock | [7 empty] | Compression force (*) - N | 80 | ±5 | Compression force (*) - N |  | N | 80 | 5 | 75 | 85 | FORCES |  | 0.14 |  |  | 0.0001 | 0.14 | Insock-FORCES |  |  | 80 - Compression force (*) - N
Insock |  |  |  | 9: Stroke length; 10: 50; 12: ±1; 11: Stroke length - mm; 13: Hành trình; 14: mm; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Insock-DIM |  |  | 50 - Stroke length - mm
Insock |  |  |  | 9: Speed (*); 10: 20; 12: N/A; 11: Speed (*) - cycle/min; 13: Tốc độ; 14: cycle/min; 17: 0; 18: 0; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Insock-RPM |  |  | 20 - Speed (*) - cycle/min
Insock |  |  |  | 9: Counter; 12: N/A; 11: Counter; 13: Bộ đếm; 17: 0; 18: 0; 19: TIME; 21: #N/A |  | #N/A | #N/A | Insock-TIME |  |  | Counter
Insock | [27 empty] | * | bar: 11
Insock | [27 empty] | * | bar: 11
Insock | [27 empty] | * | bar: 11
Flex-water | SATRA TM34/type 2 machine | Water resistance tester for upper leather | LAB-LN-03 LAB-M-01 LAB-F01:2023 LAB-TIME-02 LAB-A-01 | Water resistance tester for upper leather | Water resistance tester for upper leather |  | Distance between clamps Distance from centre of pivot point to base of "V" Distance from pivot point to clamp face Flexing angle Internal angle of "V" clamps Internal radius of "V" clamps Flexing speed | Distance between clamps | 63 | ±3 | Distance between clamps - mm |  | mm | 63 | 3 | 60 | 66 | DIM |  | 0.23 | 800000 | Service & Calibration | 0.1 | 0.23 | Flex-water-DIM |  |  | 63 - Distance between clamps - mm
Flex-water | [7 empty] | Distance from centre of pivot point to base of "V" | 54 | ±0.5 | Distance from centre of pivot point to base of "V" - mm |  | mm | 54 | 0.5 | 53.5 | 54.5 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-water-DIM |  |  | 54 - Distance from centre of pivot point to base of "V" - mm
Flex-water | [7 empty] | Distance from pivot point to clamp face | 31.5 | ±1.5 | Distance from pivot point to clamp face - mm |  | mm | 31.5 | 1.5 | 30 | 33 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-water-DIM |  |  | 31.5 - Distance from pivot point to clamp face - mm
Flex-water | [7 empty] | Flexing angle | 30 | ±1 | Flexing angle - ⁰ |  | ⁰ | 30 | 1 | 29 | 31 | ANG |  | 0.23 |  |  | 0.1 | 0.23 | Flex-water-ANG |  |  | 30 - Flexing angle - ⁰
Flex-water | [7 empty] | Internal angle of "V" clamps | 31 | ±1 | Internal angle of "V" clamps - mm |  | mm | 31 | 1 | 30 | 32 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-water-DIM |  |  | 31 - Internal angle of "V" clamps - mm
Flex-water | [7 empty] | Internal radius of "V" clamps | 7.5 | ±0.5 | Internal radius of "V" clamps - mm |  | mm | 7.5 | 0.5 | 7 | 8 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-water-DIM |  |  | 7.5 - Internal radius of "V" clamps - mm
Flex-water | [7 empty] | Flexing speed | 90 | ±5 | Flexing speed - cycle/min |  | cycle/min | 90 | 5 | 85 | 95 | TIM1 |  | 0.23 |  |  | 0.1 | 0.23 | Flex-water-TIM1 |  |  | 90 - Flexing speed - cycle/min
Flex-water | [27 empty] | * | bar: 11
Flex-water | [27 empty] | * | bar: 11
Flex-water | [27 empty] | * | bar: 11
Flex-water | bar: 11
Flex-satra47 | SATRA TM47 | Permeability absorption apparatus | LAB-LN-03 LAB-M-01 LAB-F01:2023 LAB-TIME-02 LAB-A-01 | Permeability absorption apparatus | Permeability absorption apparatus |  | Temperature of water bath Diameter of jar lid openings Internal height of test pots Air speed | Temperature of water bath (*) | 32 | ±2 | Temperature of water bath (*) - ⁰C |  | ℃ | 32 | 2 | 30 | 34 | DEG |  | 0.23 | 800000 | Service & Calibration | 0.1 | 0.23 | Flex-satra47-DEG
Flex-satra47 | [7 empty] | Diameter of jar lid openings | 40-60 | 40-60 | Diameter of jar lid openings - mm |  | mm | 50 | 10 | 40 | 60 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-satra47-DIM
Flex-satra47 | [7 empty] | Internal height of test pots | 35 | ±5 | Internal height of test pots - mm |  | mm | 35 | 5 | 30 | 40 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-satra47-DIM |  |  | 35 - Internal height of test pots - mm
Flex-satra47 | [7 empty] | Air speed | 2.2 | ±0.5 | Air speed - min:sec |  | min:sec | 2.2 | 0.5 | 1.7 | 2.7 | TIME |  | #N/A |  |  | #N/A | #N/A | Flex-satra47-TIME |  |  | 2.2 - Air speed - min:sec
Flex-satra47 | bar: 11
Flex-satra47 | bar: 11
Flex-satra47 | bar: 11
Flex-vamp | SATRA TM25 |  | TAV-QTHTD | Vamp flexing machine | Vamp flexing machine |  | Height of "V" blocks External angle of "V" block Radius of curvature of "V" blocks Minimum spacing between "V" blocks Moving stroke of clamps Flexing speed Coaxial alignment of V blocks | Height of "V" blocks | >=30 | >=30 | Height of "V" blocks - mm | >=30 | mm | 35 | 5 | 30 | 40 | DIM |  | 0.23 | 800000 | Service & Calibration | 0.1 | 0.23 | Flex-vamp-DIM
Flex-vamp | [7 empty] | External angle of "V" block | 40 | ±2 | External angle of "V" block - ⁰ | 40 | ⁰ | 40 | 2 | 38 | 42 | ANG |  | 0.23 |  |  | 0.1 | 0.23 | Flex-vamp-ANG
Flex-vamp | [7 empty] | Radius of curvature of "V" blocks | 6.4 | ±0.5 | Radius of curvature of "V" blocks - mm | 6.4 | mm | 6.4 | 0.5 | 5.9 | 6.9 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-vamp-DIM
Flex-vamp | [7 empty] | Minimum spacing between "V" blocks | 9.5 | ±1 | Minimum spacing between "V" blocks - mm | 9.5 | mm | 9.5 | 1 | 8.5 | 10.5 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-vamp-DIM
Flex-vamp | [7 empty] | Moving stroke of clamps | 19 | ±1.5 | Moving stroke of clamps - mm | 19 | mm | 19 | 1.5 | 17.5 | 20.5 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-vamp-DIM
Flex-vamp | [7 empty] | Flexing speed | 300 | ±30 | Flexing speed - flexes:min | 300 | flexes:min | 300 | 30 | 270 | 330 | TIM1 |  | 0.23 |  |  | 0.1 | 0.23 | Flex-vamp-TIM1
Flex-vamp | [7 empty] | Coaxial alignment of V blocks | not more than 2.0 | not more than 2.0 | Coaxial alignment of V blocks - mm | not more than 2.0 | mm | 1 | 1 | 0 | 2 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-vamp-DIM
Flex-vamp | bar: 11
Flex-vamp | bar: 11
Flex-vamp | bar: 11
Sole | SATRA TM404 |  | LAB-LN-03 LAB-M-01 LAB-F01:2023 LAB-TIME-02 LAB-A-01 | Sole adhension tester | Sole adhension tester |  | Accuracy of load measuring system Dimension of toe piece Angle of profile of toe piece Width of anvil Thickness of toe piece Length of arc | Accuracy of load measuring system - kgf | 10 | N/A | Accuracy of load measuring system - kgf | 10 | kgf | 10 | N/A | #VALUE! | #VALUE! | FORCE |  | 0.23 | 1000000 | Service & Calibration | 0.1 | 0.23 | Sole-FORCE
Sole |  |  |  | 10: 20; 12: N/A; 13: 20; 14: kgf; 15: 20; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCE; 21: 0.23 |  | 0.1 | 0.23 | Sole-FORCE
Sole |  |  |  | 10: 30; 12: N/A; 13: 30; 14: kgf; 15: 30; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCE; 21: 0.23 |  | 0.1 | 0.23 | Sole-FORCE
Sole |  |  |  | 10: 40; 12: N/A; 13: 40; 14: kgf; 15: 40; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCE; 21: 0.23 |  | 0.1 | 0.23 | Sole-FORCE
Sole |  |  |  | 10: 50; 12: N/A; 13: 50; 14: kgf; 15: 50; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCE; 21: 0.23 |  | 0.1 | 0.23 | Sole-FORCE
Sole | [7 empty] | Dimension of toe piece | 10 | ±1 | Dimension of toe piece - mm | 10 | mm | 10 | 1 | 9 | 11 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Sole-DIM
Sole | [7 empty] | Angle of profile of toe piece | appoximately 45 | N/A | Angle of profile of toe piece - ⁰ | appoximately 45 | ⁰ | appoximately 45 | N/A | #VALUE! | #VALUE! | ANG |  | 0.23 |  |  | 0.1 | 0.23 | Sole-ANG
Sole | [7 empty] | Width of anvil | Greater than 75 | N/A | Width of anvil - mm | Greater than 75 | mm | Greater than 75 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Sole-DIM
Sole | [7 empty] | Thickness of toe piece | 10 | ±1 | Thickness of toe piece - mm | 10 | mm | 10 | 1 | 9 | 11 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Sole-DIM
Sole | [7 empty] | Length of arc | 22 | ±1 | Length of arc - mm | 22 | mm | 22 | 1 | 21 | 23 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Sole-DIM
Sole | bar: 11
Sole | bar: 11
Sole | bar: 11
Flex-bata | SATRA TM133 |  | LAB-LN-03 LAB-M-01 LAB-F01:2023 LAB-TIME-02 LAB-A-01 | Bata flexing machine | Bata flexing machine |  | Diameter of rollers Dimension of flexible belt Larger roller speed Rotation speed | Diameter of rollers Stadard: At ends: 87 | 87 | ±1 | Diameter of rollers Stadard: At ends: 87 - mm | 87 | mm | 87 | 1 | 86 | 88 | DIM |  | 0.23 | 1200000 | Service & Calibration | 0.1 | 0.23 | Flex-bata-DIM
Flex-bata | [7 empty] | Diameter of rollers Stadard: At centre: 90 | 90 | ±0.5 | Diameter of rollers Stadard: At centre: 90 - mm | 90 | mm | 90 | 0.5 | 89.5 | 90.5 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-bata-DIM
Flex-bata | [7 empty] | Diameter of rollers Option1: At end: 57 | 57 | ±1 | Diameter of rollers Option1: At end: 57 - mm | 57 | mm | 57 | 1 | 56 | 58 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-bata-DIM
Flex-bata | [7 empty] | Diameter of rollers Option1: At centre: 60 | 60 | ±0.5 | Diameter of rollers Option1: At centre: 60 - mm | 60 | mm | 60 | 0.5 | 59.5 | 60.5 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-bata-DIM
Flex-bata | [7 empty] | Diameter of rollers Option2: At end: 117 | 117 | ±1 | Diameter of rollers Option2: At end: 117 - mm | 117 | mm | 117 | 1 | 116 | 118 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-bata-DIM
Flex-bata | [7 empty] | Diameter of rollers Option2: At centre: 120 | 120 | ±0.5 | Diameter of rollers Option2: At centre: 120 - mm | 120 | mm | 120 | 0.5 | 119.5 | 120.5 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-bata-DIM
Flex-bata | [7 empty] | Diameter of rollers Diameter : 225 | 225 | ±5 | Diameter of rollers Diameter : 225 - mm | 225 | mm | 225 | 5 | 220 | 230 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-bata-DIM
Flex-bata | [7 empty] | Diameter of rollers Width : 170 | 170 | ±20 | Diameter of rollers Width : 170 - mm | 170 | mm | 170 | 20 | 150 | 190 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-bata-DIM
Flex-bata | [7 empty] | Dimension of flexible belt Length : 1930 | 1930 | ±50 | Dimension of flexible belt Length : 1930 - mm | 1930 | mm | 1930 | 50 | 1880 | 1980 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-bata-DIM
Flex-bata | [7 empty] | Dimension of flexible belt Width : 140 | 140 | ±5 | Dimension of flexible belt Width : 140 - mm | 140 | mm | 140 | 5 | 135 | 145 | DIM |  | 0.23 |  |  | 0.1 | 0.23 | Flex-bata-DIM
Flex-bata | [7 empty] | Larger roller speed | 247 | ±20 | Larger roller speed - rev:min | 247 | rev:min | 247 | 20 | 227 | 267 | TIME |  | #N/A |  |  | #N/A | #N/A | Flex-bata-TIME |  |  | 247 - Larger roller speed - rev:min
Flex-bata | [7 empty] | Rotation speed | 90 | ±8 | Rotation speed - cycle/min | 90 | cycle/min | 90 | 8 | 82 | 98 | TIME |  | #N/A |  |  | #N/A | #N/A | Flex-bata-TIME |  |  | 90 - Rotation speed - cycle/min
Flex-bata | [27 empty] | * | bar: 11
Flex-bata | [27 empty] | * | bar: 11
Flex-bata | [27 empty] | * | bar: 11
Lastometer | SATRA TM24 |  | LAB-LN-03 LAB-M-01 LAB-F01:2023 LAB-TIME-02 LAB-A-01 | Lastometer | Lastometer |  | Diameter of the claming ring Diameter of the plunger ball Accuracy of load measuring system Accuracy of distension measuring system Speed of operation | Diameter of the claming ring | 25 | ±0.05 | Diameter of the claming ring - mm | 25 | mm | 25 | 0.05 | 24.95 | 25.05 | DIM |  | 0.02 | 1500000 | Service & Calibration | 0.01 | 0.02 | Lastometer-DIM |  |  | 25 - Diameter of the claming ring - mm
Lastometer | [7 empty] | Diameter of the plunger ball | 6.25 | ±0.05 | Diameter of the plunger ball - mm | 6.25 | mm | 6.25 | 0.05 | 6.2 | 6.3 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Lastometer-DIM |  |  | 6.25 - Diameter of the plunger ball - mm
Lastometer | [7 empty] | Accuracy of load measuring system maximum error 1kgf | 10 | N/A | Accuracy of load measuring system maximum error 1kgf - kgf | 10 | kgf | 10 | N/A | #VALUE! | #VALUE! | FORCE |  | 0.23 |  |  | 0.1 | 0.23 | Lastometer-FORCE |  |  | 10 - Accuracy of load measuring system maximum error 1kgf - kgf
Lastometer |  |  |  | 10: 20; 12: N/A; 13: 20; 14: kgf; 15: 20; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCE; 21: 0.23 |  | 0.1 | 0.23 | Lastometer-FORCE |  |  | 20
Lastometer |  |  |  | 10: 30; 12: N/A; 13: 30; 14: kgf; 15: 30; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCE; 21: 0.23 |  | 0.1 | 0.23 | Lastometer-FORCE |  |  | 30
Lastometer |  |  |  | 10: 40; 12: N/A; 13: 40; 14: kgf; 15: 40; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCE; 21: 0.23 |  | 0.1 | 0.23 | Lastometer-FORCE |  |  | 40
Lastometer |  |  |  | 10: 50; 12: N/A; 13: 50; 14: kgf; 15: 50; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCE; 21: 0.23 |  | 0.1 | 0.23 | Lastometer-FORCE |  |  | 50
Lastometer |  |  |  | 10: 60; 12: N/A; 13: 60; 14: kgf; 15: 60; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCE; 21: 0.23 |  | 0.1 | 0.23 | Lastometer-FORCE |  |  | 60
Lastometer |  |  |  | 10: 70; 12: N/A; 13: 70; 14: kgf; 15: 70; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCE; 21: 0.23 |  | 0.1 | 0.23 | Lastometer-FORCE |  |  | 70
Lastometer |  |  |  | 10: 80; 12: N/A; 13: 80; 14: kgf; 15: 80; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCE; 21: 0.23 |  | 0.1 | 0.23 | Lastometer-FORCE |  |  | 80
Lastometer | [7 empty] | Accuracy of distension measuring system | 1 | N/A | Accuracy of distension measuring system - mm | 1 | mm | 1 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Lastometer-DIM |  |  | 1 - Accuracy of distension measuring system - mm
Lastometer |  |  |  | 10: 2; 12: N/A; 13: 2; 14: mm; 15: 2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lastometer-DIM |  |  | 2
Lastometer |  |  |  | 10: 3; 12: N/A; 13: 3; 14: mm; 15: 3; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lastometer-DIM
Lastometer |  |  |  | 10: 4; 12: N/A; 13: 4; 14: mm; 15: 4; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lastometer-DIM
Lastometer |  |  |  | 10: 5; 12: N/A; 13: 5; 14: mm; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lastometer-DIM
Lastometer |  |  |  | 10: 6; 12: N/A; 13: 6; 14: mm; 15: 6; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lastometer-DIM
Lastometer |  |  |  | 10: 7; 12: N/A; 13: 7; 14: mm; 15: 7; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lastometer-DIM
Lastometer |  |  |  | 10: 8; 12: N/A; 13: 8; 14: mm; 15: 8; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lastometer-DIM
Lastometer |  |  |  | 10: 9; 12: N/A; 13: 9; 14: mm; 15: 9; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lastometer-DIM |  |  | 9
Lastometer |  |  |  | 10: 10; 12: N/A; 13: 10; 14: mm; 15: 10; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lastometer-DIM |  |  | 10
Lastometer |  |  |  | 10: 11; 12: N/A; 13: 11; 14: mm; 15: 11; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lastometer-DIM |  |  | 11
Lastometer |  |  |  | 10: 12; 12: N/A; 13: 12; 14: mm; 15: 12; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lastometer-DIM |  |  | 12
Lastometer |  |  |  | 10: 13; 12: N/A; 13: 13; 14: mm; 15: 13; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lastometer-DIM |  |  | 13
Lastometer |  |  |  | 10: 14; 12: N/A; 13: 14; 14: mm; 15: 14; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lastometer-DIM |  |  | 14
Lastometer | [7 empty] | Speed of operation | 0.2 | ±0.05 | Speed of operation - mm/sec | 0.2 | mm/sec | 0.2 | 0.05 | 0.15 | 0.25 | DIM/TIME |  | 0.23 |  |  | 0.1 | 0.23 | Lastometer-DIM/TIME |  |  | 0.2 - Speed of operation - mm/sec
Lastometer | [27 empty] | * | bar: 11
Lastometer | [27 empty] | * | bar: 11
Lastometer | [27 empty] | * | bar: 11
Taber |  | Taber Type Abrasion Tester | LAB-LN-03 LAB-M-01 LAB-F01:2023 |  | Taber Type Abrasion Tester | Thiết bị thử độ mài mòn |  | Counter - Cycle | 1000 | N/A | Counter - Cycle | Bộ đếm | Cycle | 1000 | N/A | #VALUE! | #VALUE! | TIM1 |  | 0.07 | 1000000 | Service & Calibration | 0.001 | 0.02 | Taber-TIM1 |  |  | 1000 - Counter - Cycle
Taber |  |  |  | 10: 2000; 12: N/A; 13: Bộ đếm; 14: Cycle; 15: 2000; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIM1; 21: 0.07 |  | 0.001 | 0.02 | Taber-TIM1 |  |  | 2000
Taber | [7 empty] | Speed (*) - rpm | 60rpm | ±2 | Speed (*) - rpm | Tốc độ | rpm | 60 | 2 | 58 | 62 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Taber-RPM |  |  | 60rpm - Speed (*) - rpm
Taber |  |  |  | 10: 72rpm; 12: ±2; 14: rpm; 15: 72; 16: 2; 17: 70; 18: 74; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Taber-RPM |  |  | 72rpm
Taber | [7 empty] | Load (*) - g | 750 - 750 | N/A | Load (*) - g | Tải trọng | g | 750 | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Taber-MASS |  |  | 750 - 750 - Load (*) - g
Taber |  |  |  | 10: 750 - 750; 12: N/A; 13: Tải trọng; 14: g; 15: 750; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Taber-MASS |  |  | 750 - 750
Taber |  |  |  | 10: 250 - 250; 12: N/A; 13: Tải trọng; 14: g; 15: 250; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Taber-MASS |  |  | 250 - 250
Taber |  |  |  | 10: 250 - 250; 12: N/A; 13: Tải trọng; 14: g; 15: 250; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Taber-MASS |  |  | 250 - 250
Taber | [7 empty] | Stroke length - mm |  | ±3 | Stroke length - mm | Độ dài | mm | N/A | 3 | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Taber-DIM |  |  | Stroke length - mm
Taber |  |  |  | 9: Pressure - psi; 12: ±1; 11: Pressure - psi; 13: Áp lực; 14: psi; 16: 1; 17: -1; 18: 1; 19: TIM1; 21: 0.02 |  | 0.001 | 0.02 | Taber-TIM1 |  |  | Pressure - psi
Taber | [7 empty] | Speed (*) - rpm |  | ±15 | Speed (*) - rpm | Tốc độ | rpm | 115 | 15 | 100 | 130 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Taber-RPM |  |  | Speed (*) - rpm
Taber | [7 empty] | Weight - g | 1/2 lb | N/A | Weight - g | Tải trọng | g | 1 | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Taber-MASS |  |  | 1/2 lb - Weight - g
Taber |  |  |  | 10: 1/2 lb; 12: N/A; 13: Tải trọng; 14: g; 15: 2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Taber-MASS |  |  | 1/2 lb
Taber |  |  |  | 9: Time - sec; 12: N/A; 11: Time - sec; 13: Thời gian; 14: sec; 17: 0; 18: 0; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Taber-TIME
Taber | bar: 11
Taber | bar: 11
Taber | bar: 11
Taber-jis |  | JIS DYEING RUBBER TESTER | LAB-LN-03 LAB-M-01 LAB-F01:2023 | JIS DYEING RUBBER TESTER | JIS DYEING RUBBER TESTER |  |  | Counter | 100 | N/A | Counter - Cycle | Bộ đếm | Cycle | 100 | N/A | #VALUE! | #VALUE! | TIM1 |  | 0.02 | 1000000 | Service & Calibration | 0.001 | 0.02 | Taber-jis-TIM1
Taber-jis | [7 empty] | Speed (*) | 30rpm | ±2 | Speed (*) - rpm | Tốc độ | rpm | 30 | 2 | 28 | 32 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Taber-jis-RPM
Taber-jis |  |  |  | 9: Force (*) - g; 10: Station 1; 12: N/A; 11: Force (*) - g; 13: Tải trọng; 14: g; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Taber-jis-FORCE
Taber-jis |  |  |  | 12: N/A; 14: g; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Taber-jis-FORCE |  |  | 0
Taber-jis |  |  |  | 10: Station 2; 12: N/A; 13: Tải trọng; 14: g; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Taber-jis-FORCE |  |  | Station 2
Taber-jis |  |  |  | 12: N/A; 14: g; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Taber-jis-FORCE |  |  | 0
Taber-jis |  |  |  | 10: Station 3; 12: N/A; 13: Tải trọng; 14: g; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Taber-jis-FORCE |  |  | Station 3
Taber-jis |  |  |  | 12: N/A; 14: g; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Taber-jis-FORCE |  |  | 0
Taber-jis | [7 empty] | Stroke length | Station 1 | N/A | Stroke length - mm | Hành trình | mm | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Taber-jis-DIM |  |  | Station 1 - Stroke length - mm
Taber-jis |  |  |  | 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Taber-jis-DIM |  |  | 0
Taber-jis |  |  |  | 10: Station 2; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Taber-jis-DIM |  |  | Station 2
Taber-jis |  |  |  | 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Taber-jis-DIM |  |  | 0
Taber-jis |  |  |  | 10: Station 3; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Taber-jis-DIM |  |  | Station 3
Taber-jis |  |  |  | 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Taber-jis-DIM |  |  | 0
Taber-jis | [7 empty] | Stroke length |  | ±3 | Stroke length - mm | Độ dài | mm | N/A | 3 | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Taber-jis-DIM |  |  | Stroke length - mm
Taber-jis | [7 empty] | Air Pressure (*) - psi | 5.5-6.5 | 5.5-6.5 | Air Pressure (*) - psi | Áp suất | psi | 6 | 0.5 | 5.5 | 6.5 | PRESS |  | 0.1 |  |  | 0.05 | 0.1 | Taber-jis-PRESS |  |  | 5.5-6.5 - Air Pressure (*) - psi
Taber-jis | [7 empty] | Pressure (*) - psi | 1 | ±1 | Pressure (*) - psi | Áp suất | psi | 1 | 1 | 0 | 2 | PRESS |  | 0.1 |  |  | 0.05 | 0.1 | Taber-jis-PRESS |  |  | 1 - Pressure (*) - psi
Taber-jis |  |  |  | 10: 2; 12: ±1; 14: psi; 15: 2; 16: 1; 17: 1; 18: 3; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Taber-jis-PRESS |  |  | 2
Taber-jis |  |  |  | 10: 3; 12: ±1; 14: psi; 15: 3; 16: 1; 17: 2; 18: 4; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Taber-jis-PRESS |  |  | 3
Taber-jis |  |  |  | 10: 4; 12: ±1; 14: psi; 15: 4; 16: 1; 17: 3; 18: 5; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Taber-jis-PRESS |  |  | 4
Taber-jis |  |  |  | 10: 5; 12: ±1; 14: psi; 15: 5; 16: 1; 17: 4; 18: 6; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Taber-jis-PRESS |  |  | 5
Taber-jis |  |  |  | 10: 6; 12: ±1; 14: psi; 15: 6; 16: 1; 17: 5; 18: 7; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Taber-jis-PRESS |  |  | 6
Taber-jis |  |  |  | 10: 7; 12: ±1; 14: psi; 15: 7; 16: 1; 17: 6; 18: 8; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Taber-jis-PRESS |  |  | 7
Taber-jis |  |  |  | 10: 8; 12: ±1; 14: psi; 15: 8; 16: 1; 17: 7; 18: 9; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Taber-jis-PRESS
Taber-jis |  |  |  | 10: 9; 12: ±1; 14: psi; 15: 9; 16: 1; 17: 8; 18: 10; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Taber-jis-PRESS
Taber-jis |  |  |  | 10: 10; 12: ±1; 14: psi; 15: 10; 16: 1; 17: 9; 18: 11; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Taber-jis-PRESS
Taber-jis |  |  |  | 10: 11; 12: ±1; 14: psi; 15: 11; 16: 1; 17: 10; 18: 12; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Taber-jis-PRESS
Taber-jis |  |  |  | 10: 12; 12: ±1; 14: psi; 15: 12; 16: 1; 17: 11; 18: 13; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Taber-jis-PRESS
Taber-jis | [7 empty] | Speed (*) |  | ±15 | Speed (*) - rpm | Tốc độ | rpm | 115 | 15 | 100 | 130 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Taber-jis-RPM |  |  | 33.5 - 33.5⁰C - 33.5⁰C
Taber-jis | [7 empty] | Weight | 0.5 (226.8g) | N/A | Weight - lb | Tải trọng | lb | 0.5 (226.8g) | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Taber-jis-MASS |  |  | 45 - 0.45 - 0.45
Taber-jis |  |  |  | 10: 1 (453.6g); 12: N/A; 14: lb; 15: 1 (453.6g); 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Taber-jis-MASS
Taber-jis |  |  |  | 10: 1 (453.6g); 12: N/A; 14: lb; 15: 1 (453.6g); 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Taber-jis-MASS
Taber-jis |  |  |  | 10: 1.5 (680.4g); 12: N/A; 14: lb; 15: 1.5 (680.4g); 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Taber-jis-MASS
Taber-jis |  |  |  | 10: 1.5 (680.4g); 12: N/A; 14: lb; 15: 1.5 (680.4g); 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Taber-jis-MASS
Taber-jis |  |  |  | 10: 2 (907.2g); 12: N/A; 14: lb; 15: 2 (907.2g); 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Taber-jis-MASS
Taber-jis |  |  |  | 9: Distension; 10: 5.5; 12: ±1; 11: Distension - mm; 13: Độ phồng; 14: mm; 17: 0; 18: 0; 19: DIM; 21: 0.1 |  | 0.01 | 0.02 | Taber-jis-DIM
Taber-jis | [7 empty] | Counter | 100 | N/A | Counter - min | Bộ đếm | min | 100 | N/A | #VALUE! | #VALUE! | TIME |  | 0.1 |  |  | 0.001 | 0.02 | Taber-jis-TIME
Taber-jis |  |  |  | 9: Time; 12: N/A; 11: Time - sec; 13: Thời gian; 14: sec; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Taber-jis-TIME
Taber-jis |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Taber-jis- |  |  | *
Taber-jis |  |  |  | 9: 1…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Taber-jis- |  |  | *
Taber-jis |  |  |  | 9: 2…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Taber-jis-
Taber-jis |  |  |  | 9: 3…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Taber-jis-
Taber-jis |  |  |  | 9: 4…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Taber-jis-
Taber-jis |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Taber-jis-
Tear-digital | EN ISO 13937‐1, ISO 1974, ASTM D1424 | Tear Strength | LAB-LN-03 LAB-M-01 LAB-A-01 | Elmatear 2, Elmatear 3, SDL Tear Tester | Tear Strength Tester | Thiết bị thử độ bền xé rách | Jaw separation Weight Template dimension Check weight | Jaw seperation | ISO 13937-1:2 - 3 | ±0.5 | Jaw seperation - mm | Ngàm kẹp | mm | 2.5 | 0.5 | 2 | 3 | DIM |  | 0.02 | 2000000 | Service & Calibration | 0.01 | 0.02 | Tear-digital-DIM
Tear-digital |  |  |  | 10: ASTM D1424 - 2.5; 12: ±0.25; 14: mm; 15: 2.5; 16: 0.25; 17: 2.25; 18: 2.75; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-digital-DIM
Tear-digital |  |  |  | 10: ISO 1974 - 2.8; 12: ±0.3; 14: mm; 15: 2.8; 16: 0.3; 17: 2.5; 18: 3.1; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-digital-DIM
Tear-digital | [7 empty] | Jaw width | 30 - 40 | 30-40 | Jaw width - mm | Chiều rộng ngàm kẹp | mm | 35 | 5 | 30 | 40 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-digital-DIM
Tear-digital | [7 empty] | Jaw height | 15 - 20 | 15-20 | Jaw height - mm | Chiều cao ngàm kẹp | mm | 17.5 | 2.5 | 15 | 20 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-digital-DIM
Tear-digital | [7 empty] | Check weight - N | Pendulum A - 8 | N/A | Check weight - N | Kiểm tra khối lượng | N | 4.0455 | 0.0895 | 3.956 | 4.135 | FORCE |  | 0.0064 |  |  | 0.01 | 0.1 | Tear-digital-FORCE
Tear-digital |  |  |  | 10: Pendulum B - 16; 12: N/A; 14: N; 15: 8.09; 16: 0.18; 17: 7.91; 18: 8.27; 19: FORCE; 21: 0.0128 |  | 0.01 | 0.1 | Tear-digital-FORCE |  |  | Pendulum B - 16
Tear-digital |  |  |  | 10: Pendulum C - 32; 12: N/A; 14: N; 15: 16.18; 16: 0.36; 17: 15.82; 18: 16.54; 19: FORCE; 21: 0.0256 |  | 0.01 | 0.1 | Tear-digital-FORCE |  |  | Pendulum C - 32
Tear-digital |  |  |  | 10: Pendulum D - 64; 12: N/A; 14: N; 15: 32.365; 16: 0.715; 17: 31.65; 18: 33.08; 19: FORCE; 21: 0.0512 |  | 0.01 | 0.1 | Tear-digital-FORCE |  |  | Pendulum D - 64
Tear-digital |  |  |  | 10: Pendulum E - 128; 12: N/A; 14: N; 15: 64.73; 16: 1.43; 17: 63.3; 18: 66.16; 19: FORCE; 21: 0.1024 |  | 0.01 | 0.1 | Tear-digital-FORCE |  |  | Pendulum E - 128
Tear-digital | [7 empty] | Template demension | 63 | ±1 | Template demension - mm | Kích thước tấm mẫu cắt | mm | 63 | 1 | 62 | 64 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-digital-DIM |  |  | 63 - Template demension - mm
Tear-digital |  |  |  | 10: 100; 12: ±2; 14: mm; 15: 100; 16: 2; 17: 98; 18: 102; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-digital-DIM |  |  | 100
Tear-digital | [7 empty] | 100 x 63 (Shaped) | 63 | ±1 | 100 x 63 (Shaped) - mm |  | mm | 63 | 1 | 62 | 64 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-digital-DIM |  |  | 63 - 100 x 63 (Shaped) - mm
Tear-digital |  |  |  | 10: 100; 12: ±2; 14: mm; 15: 100; 16: 2; 17: 98; 18: 102; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-digital-DIM |  |  | 100
Tear-digital | [7 empty] | Blade setting tool | 15 | ±0.5 | Blade setting tool - mm | Thước chỉnh dao | mm | 15 | 0.5 | 14.5 | 15.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-digital-DIM |  |  | 15 - Blade setting tool - mm
Tear-digital |  |  |  | 10: 20; 12: ±0.5; 14: mm; 15: 20; 16: 0.5; 17: 19.5; 18: 20.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-digital-DIM |  |  | 20
Tear-digital | [7 empty] | ISO 13937-1 | 100 | N/A | ISO 13937-1 - mm |  | mm | 100 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-digital-DIM |  |  | 100 - ISO 13937-1 - mm
Tear-digital |  |  |  | 10: 63; 12: N/A; 14: mm; 15: 63; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-digital-DIM |  |  | 63
Tear-digital |  |  |  | 10: 75; 12: N/A; 14: mm; 15: 75; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-digital-DIM |  |  | 75
Tear-digital | [7 empty] | ASTM D1424-96 | 102 | N/A | ASTM D1424-96 - mm |  | mm | 102 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-digital-DIM |  |  | 102 - ASTM D1424-96 - mm
Tear-digital |  |  |  | 10: 63; 12: N/A; 14: mm; 15: 63; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-digital-DIM |  |  | 63
Tear-digital |  |  |  | 10: 75; 12: N/A; 14: mm; 15: 75; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-digital-DIM |  |  | 75
Tear-digital | [7 empty] | Strip cut | 20 | N/A | Strip cut - mm | Strip cut | mm | 20 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-digital-DIM |  |  | 20 - Strip cut - mm
Tear-digital |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Tear-digital- |  |  | *
Tear-digital |  |  |  | 9: 1…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Tear-digital- |  |  | *
Tear-digital |  |  |  | 9: 2…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Tear-digital- |  |  | *
Tear-digital |  |  |  | 9: 3…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Tear-digital- |  |  | *
Tear-digital |  |  |  | 9: 4…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Tear-digital- |  |  | *
Tear-digital |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Tear-digital- |  |  | *
Tear-elmendorf |  | Tear Strength | LAB-LN-03 LAB-M-01 LAB-A-01 | Elmendorf Tearing Tester | Tear Strength Tester | Thiết bị thử độ bền xé rách | Slit Length Jaw Dimension Weight Mass Ruler Sample | Slit Length | 20 | ±0.15 | Slit Length - mm | Độ rộng khe hở | mm | 20 | 0.15 | 19.85 | 20.15 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-elmendorf-DIM |  |  | 20 - Slit Length - mm
Tear-elmendorf | [7 empty] | Jaw Dimension 1 | Width 30-40 | 30-40 | Jaw Dimension 1 - mm | Kích thước ngàm kẹp 1 | mm | 35 | 5 | 30 | 40 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-elmendorf-DIM |  |  | Width 30-40 - Jaw Dimension 1 - mm
Tear-elmendorf |  |  |  | 10: Height 15-20; 12: 15-20; 14: mm; 15: 17.5; 16: 2.5; 17: 15; 18: 20; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-elmendorf-DIM |  |  | Height 15-20
Tear-elmendorf | [7 empty] | Jaw Dimension 2 | Width 30-40 | 30-40 | Jaw Dimension 2 - mm | Kích thước ngàm kẹp 2 | mm | 35 | 5 | 30 | 40 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-elmendorf-DIM |  |  | Width 30-40 - Jaw Dimension 2 - mm
Tear-elmendorf |  |  |  | 10: Height 15-20; 12: 15-20; 14: mm; 15: 17.5; 16: 2.5; 17: 15; 18: 20; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-elmendorf-DIM |  |  | Height 15-20
Tear-elmendorf | [7 empty] | Weight Mass - N | Pendulum A | N/A | Weight Mass - N | Kiểm tra khối lượng | N | 4.0455 | 0.0895 | 3.956 | 4.135 | FORCE |  | 0.0064 |  |  | 0.01 | 0.1 | Tear-elmendorf-FORCE
Tear-elmendorf |  |  |  | 10: Pendulum B; 12: N/A; 14: N; 15: 8.09; 16: 0.18; 17: 7.91; 18: 8.27; 19: FORCE; 21: 0.0128 |  | 0.01 | 0.1 | Tear-elmendorf-FORCE
Tear-elmendorf |  |  |  | 10: Pendulum C; 12: N/A; 14: N; 15: 16.18; 16: 0.36; 17: 15.82; 18: 16.54; 19: FORCE; 21: 0.0256 |  | 0.01 | 0.1 | Tear-elmendorf-FORCE
Tear-elmendorf |  |  |  | 10: Pendulum D; 12: N/A; 14: N; 15: 32.365; 16: 0.715; 17: 31.65; 18: 33.08; 19: FORCE; 21: 0.0512 |  | 0.01 | 0.1 | Tear-elmendorf-FORCE
Tear-elmendorf |  |  |  | 10: Pendulum E; 12: N/A; 14: N; 15: 64.73; 16: 1.43; 17: 63.3; 18: 66.16; 19: FORCE; 21: 0.1024 |  | 0.01 | 0.1 | Tear-elmendorf-FORCE
Tear-elmendorf |  |  |  | 11: *; 19: FORCE; 21: 0.1 |  | 0.01 | 0.1 | Tear-elmendorf-FORCE
Tear-elmendorf |  |  |  | 11: *; 19: FORCE; 21: 0.1 |  | 0.01 | 0.1 | Tear-elmendorf-FORCE
Tear-elmendorf | [7 empty] | Ruler Sample | 63 | ±0.5 | Ruler Sample - mm | Thước mẫu | mm | 63 | 0.5 | 62.5 | 63.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-elmendorf-DIM
Tear-elmendorf |  |  |  | 10: 100; 12: ±0.5; 14: mm; 15: 100; 16: 0.5; 17: 99.5; 18: 100.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-elmendorf-DIM
Tear-elmendorf | [7 empty] | Ruler Sample | 68 | ±0.5 | Ruler Sample - mm | Thước mẫu | mm | 68 | 0.5 | 67.5 | 68.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-elmendorf-DIM
Tear-elmendorf |  |  |  | 10: 80; 12: ±0.5; 14: mm; 15: 80; 16: 0.5; 17: 79.5; 18: 80.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-elmendorf-DIM
Tear-elmendorf |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Tear-elmendorf-
Tear-elmendorf |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Tear-elmendorf- |  |  | *
Tear-elmendorf |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Tear-elmendorf- |  |  | *
Tear-1555 | EN ISO 13937‐1, ISO 1974, ASTM D1424 | Tear Strength | LAB-LN-03 LAB-M-01 LAB-A-01 | Elmatear 1555 | Tear Strength Tester | Thiết bị thử độ bền xé rách | Jaw separation Weight Template dimension Check weight | Jaw seperation | ISO 13937-1:2000 - 3 | ±0.5 | Jaw seperation - mm | Ngàm kẹp | mm | 3 | 0.5 | 2.5 | 3.5 | DIM |  | 0.02 | 2000000 | Service & Calibration | 0.01 | 0.02 | Tear-1555-DIM |  |  | ISO 13937-1:2000 - 3 - Jaw seperation - mm
Tear-1555 | [7 empty] | Jaw width | 30 - 40 | 30-40 | Jaw width - mm | Chiều rộng ngàm kẹp | mm | 35 | 5 | 30 | 40 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-1555-DIM |  |  | 30 - 40 - Jaw width - mm
Tear-1555 | [7 empty] | Jaw height | 15 - 20 | 15-20 | Jaw height - mm | Chiều cao ngàm kẹp | mm | 17.5 | 2.5 | 15 | 20 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-1555-DIM |  |  | 15 - 20 - Jaw height - mm
Tear-1555 | [7 empty] | Check weight - N | Pendulum 1/4 A - 2 | 0.989-1.034 | Check weight - N | Kiểm tra khối lượng | N | 1.0115 | 0.0225 | 0.989 | 1.034 | FORCE |  | 0.0016 |  |  | 0.01 | 0.1 | Tear-1555-FORCE |  |  | Pendulum 1/4 A - 2 - Check weight - N
Tear-1555 |  |  |  | 10: Pendulum 1/2 A - 4; 12: 1.978-2.068; 14: N; 15: 2.023; 16: 0.045; 17: 1.978; 18: 2.068; 19: FORCE; 21: 0.0032 |  | 0.01 | 0.1 | Tear-1555-FORCE |  |  | Pendulum 1/2 A - 4
Tear-1555 |  |  |  | 10: Pendulum A - 8; 12: 3.956 - 4.135; 14: N; 15: 4.0455; 16: 0.0895; 17: 3.956; 18: 4.135; 19: FORCE; 21: 0.0064 |  | 0.01 | 0.1 | Tear-1555-FORCE |  |  | Pendulum A - 8
Tear-1555 |  |  |  | 10: Pendulum B - 16; 12: 7.91 - 8.27; 14: N; 15: 8.09; 16: 0.18; 17: 7.91; 18: 8.27; 19: FORCE; 21: 0.0128 |  | 0.01 | 0.1 | Tear-1555-FORCE |  |  | Pendulum B - 16
Tear-1555 |  |  |  | 10: Pendulum C - 32; 12: 15.82 - 16.54; 14: N; 15: 16.18; 16: 0.36; 17: 15.82; 18: 16.54; 19: FORCE; 21: 0.0256 |  | 0.01 | 0.1 | Tear-1555-FORCE |  |  | Pendulum C - 32
Tear-1555 |  |  |  | 10: Pendulum D - 64; 12: 31.65 - 33.08; 14: N; 15: 32.365; 16: 0.715; 17: 31.65; 18: 33.08; 19: FORCE; 21: 0.0512 |  | 0.01 | 0.1 | Tear-1555-FORCE |  |  | Pendulum D - 64
Tear-1555 |  |  |  | 10: Pendulum E - 128; 12: 63.3 - 66.16; 14: N; 15: 64.73; 16: 1.43; 17: 63.3; 18: 66.16; 19: FORCE; 21: 0.1024 |  | 0.01 | 0.1 | Tear-1555-FORCE |  |  | Pendulum E - 128
Tear-1555 | [7 empty] | Template demension | 63 | ±1 | Template demension - mm | Kích thước tấm mẫu cắt | mm | 63 | 1 | 62 | 64 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-1555-DIM |  |  | 63 - Template demension - mm
Tear-1555 |  |  |  | 10: 100; 12: ±2; 14: mm; 15: 100; 16: 2; 17: 98; 18: 102; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-1555-DIM |  |  | 100
Tear-1555 | [7 empty] | 100 x 63 (Shaped) | 63 | ±1 | 100 x 63 (Shaped) - mm |  | mm | 63 | 1 | 62 | 64 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-1555-DIM |  |  | 63 - 100 x 63 (Shaped) - mm
Tear-1555 |  |  |  | 10: 100; 12: ±2; 14: mm; 15: 100; 16: 2; 17: 98; 18: 102; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-1555-DIM |  |  | 100
Tear-1555 | [7 empty] | Blade setting tool | 15 | ±0.5 | Blade setting tool - mm | Thước chỉnh dao | mm | 15 | 0.5 | 14.5 | 15.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-1555-DIM |  |  | 15 - Blade setting tool - mm
Tear-1555 |  |  |  | 10: 20; 12: ±0.5; 14: mm; 15: 20; 16: 0.5; 17: 19.5; 18: 20.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-1555-DIM
Tear-1555 | [7 empty] | ISO 13937-1 | 100 | ±2 | ISO 13937-1 - mm |  | mm | 100 | 2 | 98 | 102 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-1555-DIM
Tear-1555 |  |  |  | 10: 63; 12: ±1; 14: mm; 15: 63; 16: 1; 17: 62; 18: 64; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-1555-DIM
Tear-1555 | [7 empty] | ASTM D1424-96 | 102 | ±2 | ASTM D1424-96 - mm |  | mm | 102 | 2 | 100 | 104 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-1555-DIM
Tear-1555 |  |  |  | 10: 63; 12: ±1; 14: mm; 15: 63; 16: 1; 17: 62; 18: 64; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-1555-DIM
Tear-1555 | [7 empty] | Strip cut | 15 | N/A | Strip cut - mm |  | mm | 15 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tear-1555-DIM
Tear-1555 |  |  |  | 10: 20; 12: N/A; 14: mm; 15: 20; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tear-1555-DIM
Tear-1555 |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Tear-1555-
Tear-1555 |  |  |  | 9: 1…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Tear-1555-
Tear-1555 |  |  |  | 9: 2…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Tear-1555-
Tear-1555 |  |  |  | 9: 3…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Tear-1555-
Tear-1555 |  |  |  | 9: 4…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Tear-1555-
Tear-1555 |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Tear-1555-
Burst-pneumatic | Adidas 4.09, ISO 13938-2, ASTM D 3786-6 | Pneumatic Bursting Strength | LAB-TIME-02 LAB-P-01 LAB-LN-03 | TruBurst, AutoBurst, PnuBurst | Pneumatic Bursting Strength Tester | Thiết bị thử độ bục | Time Pressure Dome Diameter Distension | 'Time | 10 | ±0.5 | Time - sec | Thời gian | sec | 10 | 0.5 | 9.5 | 10.5 | TIME |  | 0.02 | 2500000 | Service & Calibration | 0.001 | 0.02 | Burst-pneumatic-TIME
Burst-pneumatic |  |  |  | 10: 20; 12: ±0.5; 14: sec; 15: 20; 16: 0.5; 17: 19.5; 18: 20.5; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Burst-pneumatic-TIME
Burst-pneumatic |  |  |  | 10: 30; 12: ±0.5; 14: sec; 15: 30; 16: 0.5; 17: 29.5; 18: 30.5; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Burst-pneumatic-TIME
Burst-pneumatic | [7 empty] | Pressure (*) - kpa | 100 | ±20 | Pressure (*) - kpa | Áp suất | kpa | 100 | 20 | 80 | 120 | PRESS |  | 0.1 |  |  | 0.05 | 0.1 | Burst-pneumatic-PRESS
Burst-pneumatic |  |  |  | 10: 200; 12: ±20; 14: kpa; 15: 200; 16: 20; 17: 180; 18: 220; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-pneumatic-PRESS
Burst-pneumatic |  |  |  | 10: 300; 12: ±20; 14: kpa; 15: 300; 16: 20; 17: 280; 18: 320; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-pneumatic-PRESS
Burst-pneumatic |  |  |  | 10: 400; 12: ±20; 14: kpa; 15: 400; 16: 20; 17: 380; 18: 420; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-pneumatic-PRESS |  |  | 400
Burst-pneumatic |  |  |  | 10: 500; 12: ±20; 14: kpa; 15: 500; 16: 20; 17: 480; 18: 520; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-pneumatic-PRESS |  |  | 500
Burst-pneumatic |  |  |  | 10: 600; 12: ±20; 14: kpa; 15: 600; 16: 20; 17: 580; 18: 620; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-pneumatic-PRESS |  |  | 600
Burst-pneumatic |  |  |  | 10: 700; 12: ±20; 14: kpa; 15: 700; 16: 20; 17: 680; 18: 720; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-pneumatic-PRESS |  |  | 700
Burst-pneumatic |  |  |  | 10: 800; 12: ±20; 14: kpa; 15: 800; 16: 20; 17: 780; 18: 820; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-pneumatic-PRESS |  |  | 800
Burst-pneumatic |  |  |  | 10: 900; 12: ±20; 14: kpa; 15: 900; 16: 20; 17: 880; 18: 920; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-pneumatic-PRESS |  |  | 900
Burst-pneumatic |  |  |  | 10: 1000; 12: ±20; 14: kpa; 15: 1000; 16: 20; 17: 980; 18: 1020; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-pneumatic-PRESS |  |  | 1000
Burst-pneumatic | [7 empty] | Dome Diameter | Dome 7.3 cm² - 30.5mm | ±0.2 | Dome Diameter - mm | Đường kính mặt dưới của vòm | mm | 30.5 | 0.2 | 30.3 | 30.7 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Burst-pneumatic-DIM
Burst-pneumatic |  |  |  | 10: Dome 10 cm² - 37.5mm; 12: ±0.2; 14: mm; 15: 37.5; 16: 0.2; 17: 37.3; 18: 37.7; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM
Burst-pneumatic |  |  |  | 10: Dome 50 cm² - 79.8mm; 12: ±0.2; 14: mm; 15: 79.8; 16: 0.2; 17: 79.6; 18: 80; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM
Burst-pneumatic |  |  |  | 10: Dome 100cm² - 112.8mm; 12: ±0.2; 14: mm; 15: 112.8; 16: 0.2; 17: 112.6; 18: 113; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | 33.5 - 33.5⁰C - 33.5⁰C
Burst-pneumatic |  |  |  | 11: *; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | *
Burst-pneumatic |  |  |  | 11: *; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | *
Burst-pneumatic |  |  |  | 11: *; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | *
Burst-pneumatic | [7 empty] | Ring Diameter | Ring 7.3 cm² - 30.5mm | ±0.2 | Ring Diameter - mm | Đường kính vòng | mm | 30.5 | 0.2 | 30.3 | 30.7 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | Ring 7.3 cm² - 30.5mm - Ring Diameter - mm
Burst-pneumatic |  |  |  | 10: Ring 10 cm² - 37.5mm; 12: ±0.2; 14: mm; 15: 37.5; 16: 0.2; 17: 37.3; 18: 37.7; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | Ring 10 cm² - 37.5mm
Burst-pneumatic |  |  |  | 10: Ring 50 cm² - 79.8mm; 12: ±0.2; 14: mm; 15: 79.8; 16: 0.2; 17: 79.6; 18: 80; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | Ring 50 cm² - 79.8mm
Burst-pneumatic |  |  |  | 10: Ring 100cm² - 112.8mm; 12: ±0.2; 14: mm; 15: 112.8; 16: 0.2; 17: 112.6; 18: 113; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | Ring 100cm² - 112.8mm
Burst-pneumatic |  |  |  | 11: *; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | *
Burst-pneumatic |  |  |  | 11: *; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | *
Burst-pneumatic |  |  |  | 11: *; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | *
Burst-pneumatic | [7 empty] | Distension | 10 | ±1 | Distension - mm | Độ phồng | mm | 10 | 1 | 9 | 11 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | 10 - Distension - mm
Burst-pneumatic |  |  |  | 10: 20; 12: ±1; 14: mm; 15: 20; 16: 1; 17: 19; 18: 21; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | 20
Burst-pneumatic |  |  |  | 10: 30; 12: ±1; 14: mm; 15: 30; 16: 1; 17: 29; 18: 31; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM
Burst-pneumatic |  |  |  | 10: 40; 12: ±1; 14: mm; 15: 40; 16: 1; 17: 39; 18: 41; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | 40
Burst-pneumatic |  |  |  | 10: 50; 12: ±1; 14: mm; 15: 50; 16: 1; 17: 49; 18: 51; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | 50
Burst-pneumatic |  |  |  | 10: 60; 12: ±1; 14: mm; 15: 60; 16: 1; 17: 59; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | 60
Burst-pneumatic |  |  |  | 10: 70; 12: ±1; 14: mm; 15: 70; 16: 1; 17: 69; 18: 71; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Burst-pneumatic-DIM |  |  | 70
Burst-pneumatic |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Burst-pneumatic- |  |  | *
Burst-pneumatic |  |  |  | 9: 1…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Burst-pneumatic- |  |  | *
Burst-pneumatic |  |  |  | 9: 2…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Burst-pneumatic- |  |  | *
Burst-pneumatic |  |  |  | 9: 3…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Burst-pneumatic- |  |  | *
Burst-pneumatic |  |  |  | 9: 4…; 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Burst-pneumatic- |  |  | *
Burst-pneumatic |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Burst-pneumatic- |  |  | *
Burst-mullen | ASTM D3786-1, ISO 13938-1 | Hydraulic Bursting Strength | LAB-TIME-02 LAB-P-01 LAB-LN-03 | Mullen | Hydraulic Bursting Strength Tester | Thiết bị thử độ bục | Pressure Foil | Pressure (*) - psi/bar/(kg/cm²)/mpa/…... Áp suất | Gauge 1 ….... psi/bar/(kg/cm²)/mpa/…... |  | * | Áp suất |  |  |  | 0 | 0 | PRESS |  | 0.1 | 1800000 | Service & Calibration | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | Gauge 1 ….... psi/bar/(kg/cm²)/mpa/…... - *
Burst-mullen |  |  |  | 10: 1 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 1 -
Burst-mullen |  |  |  | 10: 2 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 2 -
Burst-mullen |  |  |  | 10: 3 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 3 -
Burst-mullen |  |  |  | 10: 4 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 4 -
Burst-mullen |  |  |  | 10: 5 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 5 -
Burst-mullen |  |  |  | 10: 6 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 6 -
Burst-mullen |  |  |  | 10: 7 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 7 -
Burst-mullen |  |  |  | 10: 8 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 8 -
Burst-mullen |  |  |  | 10: 9 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 9 -
Burst-mullen |  |  |  | 10: 10 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 10 -
Burst-mullen |  |  |  | 10: Gauge 2 ….... psi/bar/(kg/cm²)/mpa/…...; 21: #N/A |  | #N/A | #N/A | Burst-mullen- |  |  | Gauge 2 ….... psi/bar/(kg/cm²)/mpa/…...
Burst-mullen |  |  |  | 10: 1 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 1 -
Burst-mullen |  |  |  | 10: 2 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 2 -
Burst-mullen |  |  |  | 10: 3 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 3 -
Burst-mullen |  |  |  | 10: 4 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 4 -
Burst-mullen |  |  |  | 10: 5 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 5 -
Burst-mullen |  |  |  | 10: 6 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 6 -
Burst-mullen |  |  |  | 10: 7 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 7 -
Burst-mullen |  |  |  | 10: 8 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 8 -
Burst-mullen |  |  |  | 10: 9 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 9 -
Burst-mullen |  |  |  | 10: 10 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 10 -
Burst-mullen |  |  |  | 10: Gauge 3 ….... psi/bar/(kg/cm²)/mpa/…...; 21: #N/A |  | #N/A | #N/A | Burst-mullen- |  |  | Gauge 3 ….... psi/bar/(kg/cm²)/mpa/…...
Burst-mullen |  |  |  | 10: 1 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 1 -
Burst-mullen |  |  |  | 10: 2 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 2 -
Burst-mullen |  |  |  | 10: 3 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 3 -
Burst-mullen |  |  |  | 10: 4 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 4 -
Burst-mullen |  |  |  | 10: 5 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 5 -
Burst-mullen |  |  |  | 10: 6 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 6 -
Burst-mullen |  |  |  | 10: 7 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 7 -
Burst-mullen |  |  |  | 10: 8 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 8 -
Burst-mullen |  |  |  | 10: 9 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS |  |  | 9 -
Burst-mullen |  |  |  | 10: 10 -; 12: ±; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS
Burst-mullen |  |  |  | 9: Check foil; 12: ±; 11: Check foil - psi; 13: Kiểm tra màng nhôm; 14: psi; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Burst-mullen-PRESS
Burst-mullen |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Burst-mullen-
Burst-mullen |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Burst-mullen-
Stoll |  | Surface Abrasion Tester | LAB-TIME-02 LAB-P-01 LAB-LN-03 | CSI | Surface Abrasion Tester | Thiết bị thử độ bục | Pressure Foil | Pressure (*) |  | 5.5 - 6.5 | Pressure (*) - psi | Áp suất | psi |  | 5.5 - 6.5 | 5.5 | 6.5 | PRESS |  | 0.1 | 1800000 | Service & Calibration | 0.05 | 0.1 | Stoll-PRESS |  |  | Pressure (*) - psi
Stoll |  |  |  | 11: *; 14: psi; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Stoll-PRESS |  |  | *
Stoll |  |  |  | 11: *; 14: psi; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Stoll-PRESS |  |  | *
Stoll |  |  |  | 11: *; 14: psi; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Stoll-PRESS |  |  | *
Stoll |  |  |  | 11: *; 14: psi; 17: 0; 18: 0; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Stoll-PRESS |  |  | *
Stoll | [7 empty] | Height | 15 | N/A | Height - mm | Chiều cao màng | mm | 15 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Stoll-DIM |  |  | 15 - Height - mm
Stoll | [7 empty] | Diameter - mm | 15 | N/A | Diameter - mm | Đường kính thử mẫu | mm | 15 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Stoll-DIM |  |  | 15 - Diameter - mm
Stoll | [7 empty] | Speed (*) | 120 | ±10 | Speed (*) - rpm | Tốc độ | rpm | 120 | 10 | 110 | 130 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Stoll-RPM |  |  | 120 - Speed (*) - rpm
Stoll |  |  |  | 9: Stroke length; 10: N/A; 12: N/A; 11: Stroke length - mm; 13: Độ dài hành trình; 14: mm; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Stoll-DIM |  |  | N/A - Stroke length - mm
Stoll | [7 empty] | Weight | 1 | N/A | Weight - lb | Tạ | lb | 1 | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Stoll-MASS
Stoll |  |  |  | 10: 2; 12: N/A; 14: lb; 15: 2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Stoll-MASS
Stoll |  |  |  | 10: 1; 12: N/A; 14: lb; 15: 1; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Stoll-MASS | 1
Stoll |  |  |  | 10: 2; 12: N/A; 14: lb; 15: 2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Stoll-MASS | 1
Stoll |  |  |  | 11: *; 14: lb; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Stoll- | 1
Light-xrite | NIST X-Rite | Light box | LAB-L-01 | 5: SPL QC, SPL III; 6: Light box; 7: Tủ đèn so màu; 8: Colour temperature Illuminance CALIBRATION; 21: #N/A; 22: 4000000 | Service & Calibration by X-rite Authorized Service Partner | #N/A | #N/A | Light-xrite-
Light-xrite | [7 empty] | Colour temperature (*) | D65 - 6500 | ±200 | Colour temperature (*) - K | Nhiệt độ màu | K | 6500 | 200 | 6300 | 6700 | KEL |  | #N/A |  |  | #N/A | #N/A | Light-xrite-KEL
Light-xrite |  |  |  | 10: Hz - 2300; 12: ±200; 14: K; 15: 2300; 16: 200; 17: 2100; 18: 2500; 19: KEL; 21: #N/A |  | #N/A | #N/A | Light-xrite-KEL
Light-xrite |  |  |  | 10: A - 2856; 12: ±200; 14: K; 15: 2856; 16: 200; 17: 2656; 18: 3056; 19: KEL; 21: #N/A |  | #N/A | #N/A | Light-xrite-KEL
Light-xrite |  |  |  | 10: CWF - 4150; 12: ±200; 14: K; 15: 4150; 16: 200; 17: 3950; 18: 4350; 19: KEL; 21: #N/A |  | #N/A | #N/A | Light-xrite-KEL |  |  | CWF - 4150
Light-xrite |  |  |  | 10: TL84 - 4000; 12: ±200; 14: K; 15: 4000; 16: 200; 17: 3800; 18: 4200; 19: KEL; 21: #N/A |  | #N/A | #N/A | Light-xrite-KEL |  |  | TL84 - 4000
Light-xrite |  |  |  | 10: TL83 - 3000; 12: ±200; 14: K; 15: 3000; 16: 200; 17: 2800; 18: 3200; 19: KEL; 21: #N/A |  | #N/A | #N/A | Light-xrite-KEL |  |  | TL83 - 3000
Light-xrite |  |  |  | 10: U30 - 3000; 12: ±200; 14: K; 15: 3000; 16: 200; 17: 2800; 18: 3200; 19: KEL; 21: #N/A |  | #N/A | #N/A | Light-xrite-KEL |  |  | U30 - 3000
Light-xrite |  |  |  | 10: U35 - 3500; 12: ±200; 14: K; 15: 3500; 16: 200; 17: 3300; 18: 3700; 19: KEL; 21: #N/A |  | #N/A | #N/A | Light-xrite-KEL |  |  | U35 - 3500
Light-xrite |  |  |  | 9: Illuminance (*); 10: D65; 12: N/A; 11: Illuminance (*) - Lux; 13: Độ rọi; 14: Lux; 19: LUX; 21: #N/A |  | #N/A | #N/A | Light-xrite-LUX |  |  | D65 - Illuminance (*) - Lux
Light-xrite |  |  |  | 10: Hz; 12: N/A; 14: Lux; 19: LUX; 21: #N/A |  | #N/A | #N/A | Light-xrite-LUX |  |  | Hz
Light-xrite |  |  |  | 10: A; 12: N/A; 14: Lux; 19: LUX; 21: #N/A |  | #N/A | #N/A | Light-xrite-LUX |  |  | A
Light-xrite |  |  |  | 10: CWF; 12: N/A; 14: Lux; 19: LUX; 21: #N/A |  | #N/A | #N/A | Light-xrite-LUX |  |  | CWF
Light-xrite |  |  |  | 10: TL84; 12: N/A; 14: Lux; 19: LUX; 21: #N/A |  | #N/A | #N/A | Light-xrite-LUX |  |  | TL84
Light-xrite |  |  |  | 10: TL83; 12: N/A; 14: Lux; 19: LUX; 21: #N/A |  | #N/A | #N/A | Light-xrite-LUX |  |  | TL83
Light-xrite |  |  |  | 10: U30; 12: N/A; 14: Lux; 19: LUX; 21: #N/A |  | #N/A | #N/A | Light-xrite-LUX |  |  | U30
Light-xrite |  |  |  | 10: U35; 12: N/A; 14: Lux; 19: LUX; 21: #N/A |  | #N/A | #N/A | Light-xrite-LUX |  |  | U35
Light-xrite |  |  |  | 10: UV; 12: N/A; 14: Lux; 19: LUX; 21: #N/A |  | #N/A | #N/A | Light-xrite-LUX |  |  | UV
Light-uv |  | Lightfastness machine | TAV-QTHND TAV-QT.17-RAD | QUV | Lightfastness machine | Thiết bị thử độ bền màu ánh sáng | Time Pressure Dome Diameter Distension | Irradiance - W/m² | 0.78 | N/A | Irradiance - W/m² | Năng lượng bức xạ | W/m2 | 0.78 | N/A | #VALUE! | #VALUE! | RAD |  | 0.035 |  |  | 1 | 0.035 | Light-uv-RAD |  |  | 0.78 - Irradiance - W/m²
Light-uv |  |  |  | 12: N/A; 11: *; 14: W/m2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RAD; 21: 0.035 |  | 1 | 0.035 | Light-uv-RAD |  |  | *
Light-uv |  |  |  | 12: N/A; 11: *; 14: W/m2; 16: N/A; 19: RAD; 21: 0.035 |  | 1 | 0.035 | Light-uv-RAD
Light-uv |  |  |  | 12: N/A; 11: *; 14: W/m2; 16: N/A; 19: RAD; 21: 0.035 |  | 1 | 0.035 | Light-uv-RAD
Light-uv |  |  |  | 12: N/A; 11: *; 14: W/m2; 16: N/A; 19: RAD; 21: 0.035 |  | 1 | 0.035 | Light-uv-RAD
Light-uv |  |  |  | 12: N/A; 11: *; 14: W/m2; 16: N/A; 19: RAD; 21: 0.035 |  | 1 | 0.035 | Light-uv-RAD
Light-uv |  |  |  | 12: N/A; 11: *; 14: W/m2; 16: N/A; 19: RAD; 21: 0.035 |  | 1 | 0.035 | Light-uv-RAD
Light-uv |  |  |  | 12: N/A; 11: *; 14: W/m2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RAD; 21: 0.035 |  | 1 | 0.035 | Light-uv-RAD
Light-uv |  |  |  | 12: N/A; 11: *; 14: W/m2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RAD; 21: 0.035 |  | 1 | 0.035 | Light-uv-RAD
Light-uv |  |  |  | 9: Temperature (*) - ⁰C; 12: N/A; 11: Temperature (*) - ⁰C; 13: Nhiệt độ; 14: ℃; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Light-uv-DEG
Light-uv |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Light-uv-
Light-uv |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Light-uv-
Light-uv |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Light-uv-
Light-uv |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Light-uv-
Light-meter |  |  | TAV-QTHAS | Lightmeter | Lightmeter |  | Colour temperature | Colour temperature (*) | 6500 | ±300 | Colour temperature (*) - K | Nhiệt độ màu | K | N/A |  | #VALUE! | #VALUE! | KEL |  | 40 | 1500000 | Service & Verification | 1 | 40 | Light-meter-KEL
Light-meter |  |  | TAV-QTHAS | Lightmeter |  |  | Illuminance | Illuminance (*) |  | N/A | Illuminance (*) - Lux | Độ rọi | Lux |  | N/A | #VALUE! | #VALUE! | LUX |  | 0.012 |  |  | 1 | 0.012 | Light-meter-LUX
Light-meter |  |  |  | 12: N/A; 14: Lux; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-meter-LUX
Light-meter |  |  |  | 12: N/A; 14: Lux; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-meter-LUX
Light-meter |  |  |  | 12: N/A; 14: Lux; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-meter-LUX
Light-meter |  |  |  | 12: N/A; 14: Lux; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-meter-LUX
Light-meter |  |  |  | 12: N/A; 14: Lux; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-meter-LUX
Light-meter |  |  |  | 12: N/A; 14: Lux; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-meter-LUX
Light-meter |  |  |  | 12: N/A; 14: Lux; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-meter-LUX
Light-meter |  |  |  | 12: N/A; 14: Lux; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-meter-LUX
Light-meter |  |  |  | 12: N/A; 14: Lux; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-meter-LUX
Light-meter |  |  |  | 12: N/A; 14: Lux; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-meter-LUX
Light-meter |  |  |  | 12: N/A; 14: Lux; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-meter-LUX
Light-fw |  | Lightfastness machine | TAV-QTHND TAV-QT.17-RAD | Lightfastness machine | Lightfastness machine | Thiết bị thử bền màu ánh sáng | Time Pressure Dome Diameter Distension | Irradiance- W/m² | 550 | N/A | Irradiance- W/m² | Năng lượng bức xạ | W/m2 | 550 | N/A | #VALUE! | #VALUE! | RAD |  | 0.035 |  |  | 1 | 0.035 | Light-fw-RAD
Light-fw |  |  |  | 12: N/A; 14: W/m2; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RAD; 21: 0.035 |  | 1 | 0.035 | Light-fw-RAD
Light-fw |  |  |  | 11: *; 14: W/m2; 19: RAD; 21: 0.035 |  | 1 | 0.035 | Light-fw-RAD
Light-fw |  |  |  | 9: Black Standard Temp (BST) (*) - ⁰C; 12: ±2; 11: Black Standard Temp (BST) (*) - ⁰C; 13: Nhiệt độ bảng đen; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 21: 0.15 |  | 0.01 | 0.15 | Light-fw-DEG
Light-fw |  |  |  | 12: ±2; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 21: 0.15 |  | 0.01 | 0.15 | Light-fw-DEG
Light-fw |  |  |  | 11: *; 14: ℃; 17: 0; 18: 0; 19: DEG; 21: 0.15 |  | 0.01 | 0.15 | Light-fw-DEG
Light-fw | [7 empty] | Temperature (*) - ⁰C | 70 | ±2 | Temperature (*) - ⁰C | Nhiệt độ | ℃ | 70 | 2 | 68 | 72 | DEG |  | 0.15 |  |  | 0.01 | 0.15 | Light-fw-DEG
Light-fw |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Light-fw- |  |  | *
Light-fw |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Light-fw- |  |  | *
Light-fastness | AATCC TM 16, ISO B02 | Colourfastness to Light, Weathering | TAV-QTHND TAV-QT.17-RAD | 5: TruFade, Suntest; 6: Colourfastness to Light, Weathering; 7: Thiết bị thử độ bền màu ánh sáng; 17: 0; 18: 0; 21: #N/A; 22: 3000000 | Service & Verification | #N/A | #N/A | Light-fastness- |  |  | 0
Light-fastness | [6 empty] | Temperature of Oven | Black Standard Temp (BST) (*) - ⁰C | 45 | ±3 | Black Standard Temp (BST) (*) - ⁰C | Nhiệt độ bảng đen chuẩn | ℃ | 45 | 3 | 42 | 48 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Light-fastness-DEG |  |  | 45 - Black Standard Temp (BST) (*) - ⁰C
Light-fastness |  |  |  | 8: Temperature of Oven - Adidas; 10: 47; 12: ±3; 14: ℃; 15: 47; 16: 3; 17: 44; 18: 50; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Light-fastness-DEG |  |  | 47
Light-fastness |  |  |  | 12: ±; 11: *; 14: ℃; 17: 0; 18: 0; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Light-fastness-DEG |  |  | *
Light-fastness | [7 empty] | Black Panel Temp (BPT) (*) - ⁰C | 60 | ±3 | Black Panel Temp (BPT) (*) - ⁰C | Nhiệt độ bảng đen | ℃ | 60 | 3 | 57 | 63 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Light-fastness-DEG
Light-fastness |  |  |  | 10: 63; 12: ±3; 11: *; 14: G; 15: 63; 16: 3; 17: 60; 18: 66; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Light-fastness-DEG
Light-fastness |  |  |  | 12: ±; 11: *; 14: G; 17: 0; 18: 0; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Light-fastness-DEG
Light-fastness | [7 empty] | Chamber temp (CHT) (*) - ⁰C | 33 | ±2 | Chamber temp (CHT) (*) - ⁰C - ⁰C | Nhiệt độ buồng | ℃ | 33 | 2 | 31 | 35 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Light-fastness-DEG
Light-fastness |  |  |  | 10: 43; 12: ±2; 11: *; 14: G; 15: 43; 16: 2; 17: 41; 18: 45; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Light-fastness-DEG
Light-fastness |  |  |  | 12: ±; 11: *; 14: G; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Light-fastness-DEG
Light-fastness | [7 empty] | Humidity | 30 | ±5 | Humidity - % RH | Độ ẩm | % RH | 30 | 5 | 25 | 35 | RH |  | 0.4 |  |  | 0.1 | 0.4 | Light-fastness-RH
Light-fastness |  |  |  | 10: 40; 12: ±5; 14: % RH; 15: 40; 16: 5; 17: 35; 18: 45; 19: RH; 21: 0.4 |  | 0.1 | 0.4 | Light-fastness-RH
Light-fastness |  |  |  | 12: ±; 11: *; 19: RH; 21: 0.4 |  | 0.1 | 0.4 | Light-fastness-RH
Light-fastness | [7 empty] | Irradiance | 42 | ±2 | Irradiance - W/m2 | Năng lượng bức xạ | W/m2 | 42 | 2 | 40 | 44 | RAD |  | 0.035 |  |  | 1 | 0.035 | Light-fastness-RAD
Light-fastness |  |  |  | 10: 48; 12: ±2; 14: W/m2; 15: 48; 16: 2; 17: 46; 18: 50; 19: RAD; 21: 0.035 |  | 1 | 0.035 | Light-fastness-RAD
Light-fastness | bar: 12 to 11
Light-box | ASTM 1729, ISO 3664:2009 | Light box | TAV-QTHAS TAV-QT-LIGHT1 | 5: SPL QC, SPL III, Judge QC, Judge II, Verivide CAC, Other lightbox; 6: Light box; 7: Tủ đèn so màu; 8: Colour temperature Illuminance; 21: #N/A; 22: 1500000 | Service & Verification - Lightmeter | #N/A | #N/A | Light-box-
Light-box | [7 empty] | Colour temperature (*) | D65 - 6500 | ±200/±300 | Colour temperature (*) - K | Nhiệt độ màu | K | 6500 |  | 6500 | 6500 | KEL |  | 40 |  |  | 1 | 40 | Light-box-KEL
Light-box |  |  |  | 10: D50 - 5000; 12: ±200/±300; 14: K; 15: 5000; 17: 5000; 18: 5000; 19: KEL; 21: 40 |  | 1 | 40 | Light-box-KEL
Light-box |  |  |  | 10: D75 - 7500; 12: ±200/±300; 14: K; 15: 7500; 17: 7500; 18: 7500; 19: KEL; 21: 40 |  | 1 | 40 | Light-box-KEL
Light-box |  |  |  | 10: Hz - 2300; 12: ±200/±300; 14: K; 15: 2300; 17: 2300; 18: 2300; 19: KEL; 21: 40 |  | 1 | 40 | Light-box-KEL
Light-box |  |  |  | 10: A - 2856; 12: ±200/±300; 14: K; 15: 2856; 17: 2856; 18: 2856; 19: KEL; 21: 40 |  | 1 | 40 | Light-box-KEL
Light-box |  |  |  | 10: CWF - 4150; 12: ±200/±300; 14: K; 15: 4150; 17: 4150; 18: 4150; 19: KEL; 21: 40 |  | 1 | 40 | Light-box-KEL
Light-box |  |  |  | 10: TL84 - 4000; 12: ±200/±300; 14: K; 15: 4000; 17: 4000; 18: 4000; 19: KEL; 21: 40 |  | 1 | 40 | Light-box-KEL
Light-box |  |  |  | 10: TL83 - 3000; 12: ±200/±300; 14: K; 15: 3000; 17: 3000; 18: 3000; 19: KEL; 21: 40 |  | 1 | 40 | Light-box-KEL
Light-box |  |  |  | 10: U30 - 3000; 12: ±200/±300; 14: K; 15: 3000; 17: 3000; 18: 3000; 19: KEL; 21: 40 |  | 1 | 40 | Light-box-KEL
Light-box |  |  |  | 10: U35 - 3500; 12: ±200/±300; 14: K; 15: 3500; 17: 3500; 18: 3500; 19: KEL; 21: 40 |  | 1 | 40 | Light-box-KEL |  |  | U35 - 3500
Light-box | [7 empty] | Illuminance (*) | D65 | N/A | Illuminance (*) - Lux | Độ rọi | Lux | N/A |  | #VALUE! | #VALUE! | LUX |  | 0.012 |  |  | 0.01 | 0.012 | Light-box-LUX
Light-box |  |  |  | 10: D50; 12: N/A; 14: Lux; 15: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-box-LUX
Light-box |  |  |  | 10: D75; 12: N/A; 14: Lux; 15: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-box-LUX |  |  | D75
Light-box |  |  |  | 10: Hz; 12: N/A; 14: Lux; 15: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-box-LUX |  |  | Hz
Light-box |  |  |  | 10: A; 12: N/A; 14: Lux; 15: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-box-LUX |  |  | A
Light-box |  |  |  | 10: CWF; 12: N/A; 14: Lux; 15: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-box-LUX |  |  | CWF
Light-box |  |  |  | 10: TL84; 12: N/A; 14: Lux; 15: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-box-LUX |  |  | TL84
Light-box |  |  |  | 10: TL83; 12: N/A; 14: Lux; 15: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-box-LUX
Light-box |  |  |  | 10: U30; 12: N/A; 14: Lux; 15: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-box-LUX
Light-box |  |  |  | 10: U35; 12: N/A; 14: Lux; 15: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-box-LUX
Light-box |  |  |  | 10: UV; 12: N/A; 14: Lux; 15: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-box-LUX
Light-nike | AATCC EP1&EP2 | Light box | X-rite Lightspex Calibration Tool & Software | 5: SPL QC, SPL III; 6: Light box; 7: Tủ đèn so màu; 8: Colour temperature Illuminance CALIBRATION; 21: #N/A; 22: 4000000 | Service & Calibration by X-rite Authorized Service Partner | #N/A | #N/A | Light-nike-
Light-nike | [7 empty] | Colour temperature (*) | D65 - 6500 | ±200 | Colour temperature (*) - K | Nhiệt độ màu | K | 6500 | 200 | 6300 | 6700 | KEL |  | 40 |  |  | 1 | 40 | Light-nike-KEL
Light-nike |  |  |  | 10: Hz - 2300; 12: ±200; 14: K; 15: 2300; 16: 200; 17: 2100; 18: 2500; 19: KEL; 21: 40 |  | 1 | 40 | Light-nike-KEL
Light-nike |  |  |  | 10: A - 2856; 12: ±200; 14: K; 15: 2856; 16: 200; 17: 2656; 18: 3056; 19: KEL; 21: 40 |  | 1 | 40 | Light-nike-KEL
Light-nike |  |  |  | 10: CWF - 4150; 12: ±200; 14: K; 15: 4150; 16: 200; 17: 3950; 18: 4350; 19: KEL; 21: 40 |  | 1 | 40 | Light-nike-KEL |  |  | CWF - 4150
Light-nike |  |  |  | 10: TL84 - 4000; 12: ±200; 14: K; 15: 4000; 16: 200; 17: 3800; 18: 4200; 19: KEL; 21: 40 |  | 1 | 40 | Light-nike-KEL |  |  | TL84 - 4000
Light-nike |  |  |  | 10: TL83 - 3000; 12: ±200; 14: K; 15: 3000; 16: 200; 17: 2800; 18: 3200; 19: KEL; 21: 40 |  | 1 | 40 | Light-nike-KEL |  |  | TL83 - 3000
Light-nike |  |  |  | 10: U30 - 3000; 12: ±200; 14: K; 15: 3000; 16: 200; 17: 2800; 18: 3200; 19: KEL; 21: 40 |  | 1 | 40 | Light-nike-KEL |  |  | U30 - 3000
Light-nike |  |  |  | 10: U35 - 3500; 12: ±200; 14: K; 15: 3500; 16: 200; 17: 3300; 18: 3700; 19: KEL; 21: 40 |  | 1 | 40 | Light-nike-KEL |  |  | U35 - 3500
Light-nike | [7 empty] | Illuminance (*) | D65 | 1080-1340 | Illuminance (*) - Lux | Độ rọi | Lux | 1210 | 130 | 1080 | 1340 | LUX |  | 0.012 |  |  | 0.01 | 0.012 | Light-nike-LUX |  |  | D65 - Illuminance (*) - Lux
Light-nike |  |  |  | 10: Hz; 14: Lux; 17: 0; 18: 0; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-nike-LUX |  |  | Hz
Light-nike |  |  |  | 10: A; 14: Lux; 17: 0; 18: 0; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-nike-LUX |  |  | A
Light-nike |  |  |  | 10: CWF; 14: Lux; 17: 0; 18: 0; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-nike-LUX |  |  | CWF
Light-nike |  |  |  | 10: TL84; 14: Lux; 17: 0; 18: 0; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-nike-LUX |  |  | TL84
Light-nike |  |  |  | 10: TL83; 14: Lux; 17: 0; 18: 0; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-nike-LUX |  |  | TL83
Light-nike |  |  |  | 10: U30; 14: Lux; 17: 0; 18: 0; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-nike-LUX |  |  | U30
Light-nike |  |  |  | 10: U35; 14: Lux; 17: 0; 18: 0; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-nike-LUX |  |  | U35
Light-nike |  |  |  | 10: UV; 14: Lux; 17: 0; 18: 0; 19: LUX; 21: 0.012 |  | 0.01 | 0.012 | Light-nike-LUX |  |  | UV
Light-Above1m2 | ASTM 1729, ISO 3664:2009 | Light box | TAV-QTHAS TAV-QT-LIGHT5 | Light box 120, 150 - đo 3 vị trí | Light box | Tủ đèn so màu | Colour temperature Illuminance ADJUSTMENT RECOMMENDED | Colour temperature (*) | D65 - L - 6500 | ±300 | Colour temperature (*) - K | Nhiệt độ màu | K | 6500 | 300 | 6200 | 6800 | KEL |  | 40 | 1500000 | Service, Verification & Recommendation | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: D65 - M - 6500; 12: ±300; 14: K; 15: 6500; 16: 300; 17: 6200; 18: 6800; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: D65 - R - 6500; 12: ±300; 14: K; 15: 6500; 16: 300; 17: 6200; 18: 6800; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: Hz - L - 2300; 12: ±300; 14: K; 15: 2300; 16: 300; 17: 2000; 18: 2600; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: Hz - M - 2300; 12: ±300; 14: K; 15: 2300; 16: 300; 17: 2000; 18: 2600; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: Hz - R - 2300; 12: ±300; 14: K; 15: 2300; 16: 300; 17: 2000; 18: 2600; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: A - L - 2856; 12: ±300; 14: K; 15: 2856; 16: 300; 17: 2556; 18: 3156; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: A - M - 2856; 12: ±300; 14: K; 15: 2856; 16: 300; 17: 2556; 18: 3156; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: A - R - 2856; 12: ±300; 14: K; 15: 2856; 16: 300; 17: 2556; 18: 3156; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: F - L - 2800; 12: ±300; 14: K; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: F - M - 2800; 12: ±300; 14: K; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: F - R - 2800; 12: ±300; 14: K; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: CWF - L - 4150; 12: ±300; 14: K; 15: 4150; 16: 300; 17: 3850; 18: 4450; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: CWF - M - 4150; 12: ±300; 14: K; 15: 4150; 16: 300; 17: 3850; 18: 4450; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: CWF - R - 4150; 12: ±300; 14: K; 15: 4150; 16: 300; 17: 3950; 18: 4350; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: TL84 - L - 4000; 12: ±300; 14: K; 15: 4000; 16: 300; 17: 3800; 18: 4200; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: TL84 - M - 4000; 12: ±300; 14: K; 15: 4000; 16: 300; 17: 3800; 18: 4200; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: TL84 - R - 4000; 12: ±300; 14: K; 15: 4000; 16: 300; 17: 3800; 18: 4200; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: TL83 - L - 3000; 12: ±300; 14: K; 15: 3000; 16: 300; 17: 2800; 18: 3200; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: TL83 - M - 3000; 12: ±300; 14: K; 15: 3000; 16: 300; 17: 2800; 18: 3200; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL
Light-Above1m2 |  |  |  | 10: TL83 - R - 3000; 12: ±300; 14: K; 15: 3000; 16: 300; 17: 2800; 18: 3200; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL |  |  | TL83 - R - 3000
Light-Above1m2 |  |  |  | 10: U30 - L - 3000; 12: ±300; 14: K; 15: 3000; 16: 300; 17: 2800; 18: 3200; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL |  |  | U30 - L - 3000
Light-Above1m2 |  |  |  | 10: U30 - M - 3000; 12: ±300; 14: K; 15: 3000; 16: 300; 17: 2800; 18: 3200; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL |  |  | U30 - M - 3000
Light-Above1m2 |  |  |  | 10: U30 - R - 3000; 12: ±300; 14: K; 15: 3000; 16: 300; 17: 2800; 18: 3200; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL |  |  | U30 - R - 3000
Light-Above1m2 |  |  |  | 10: U35 - L - 3500; 12: ±300; 14: K; 15: 3500; 16: 300; 17: 3300; 18: 3700; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL |  |  | U35 - L - 3500
Light-Above1m2 |  |  |  | 10: U35 - M - 3500; 12: ±300; 14: K; 15: 3500; 16: 300; 17: 3300; 18: 3700; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL |  |  | U35 - M - 3500
Light-Above1m2 |  |  |  | 10: U35 - R - 3500; 12: ±300; 14: K; 15: 3500; 16: 300; 17: 3300; 18: 3700; 19: KEL; 21: 40 |  | 1 | 40 | Light-Above1m2-KEL |  |  | U35 - R - 3500
Light-Above1m2 |  |  |  | 7: Tủ đèn so màu; 9: Illuminance (*); 10: D65 - L; 12: N/A; 11: Illuminance (*) - Lux; 13: Độ rọi; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX |  |  | D65 - L - Illuminance (*) - Lux
Light-Above1m2 |  |  |  | 10: D65 - M; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX |  |  | D65 - M
Light-Above1m2 |  |  |  | 10: D65 - R; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX |  |  | D65 - R
Light-Above1m2 |  |  |  | 10: Hz - L; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX |  |  | Hz - L
Light-Above1m2 |  |  |  | 10: Hz - M; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX |  |  | Hz - M
Light-Above1m2 |  |  |  | 10: Hz - R; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX |  |  | Hz - R
Light-Above1m2 |  |  |  | 10: A - L; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX |  |  | A - L
Light-Above1m2 |  |  |  | 10: A - M; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX |  |  | A - M
Light-Above1m2 |  |  |  | 10: A - R; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX |  |  | A - R
Light-Above1m2 |  |  |  | 10: F - L; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX |  |  | F - L
Light-Above1m2 |  |  |  | 10: F - M; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX |  |  | F - M
Light-Above1m2 |  |  |  | 10: F - R; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX |  |  | F - R
Light-Above1m2 |  |  |  | 10: CWF - L; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX |  |  | CWF - L
Light-Above1m2 |  |  |  | 10: CWF - M; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX |  |  | CWF - M
Light-Above1m2 |  |  |  | 10: CWF - R; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX
Light-Above1m2 |  |  |  | 10: TL84 - L; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX
Light-Above1m2 |  |  |  | 10: TL84 - M; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX
Light-Above1m2 |  |  |  | 10: TL84 - R; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX
Light-Above1m2 |  |  |  | 10: TL83 - L; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX
Light-Above1m2 |  |  |  | 10: TL83 - M; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX
Light-Above1m2 |  |  |  | 10: TL83 - R; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX
Light-Above1m2 |  |  |  | 10: U30 - L; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX
Light-Above1m2 |  |  |  | 10: U30 - M; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX
Light-Above1m2 |  |  |  | 10: U30 - R; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX
Light-Above1m2 |  |  |  | 10: U35 - L; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX
Light-Above1m2 |  |  |  | 10: U35 - M; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX
Light-Above1m2 |  |  |  | 10: U35 - R; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX
Light-Above1m2 |  |  |  | 10: UV - L; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX
Light-Above1m2 |  |  |  | 10: UV - M; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX
Light-Above1m2 |  |  |  | 10: UV - R; 12: N/A; 14: Lux; 15: N/A; 19: LUX; 21: 0.012 |  | 1 | 0.012 | Light-Above1m2-LUX
Perspi | ISO 105 series, AATCC 15 | Perspiration, Phenolic Yellowing, Colour Migration | TAV-QT-PERS TAV - QTHKL | Perspirometer | Perspirometer | Bộ thử bền màu mồ hôi | Weight | Top plates mass (*) | 454 | N/A | Top plates mass (*) - g | Khối lượng tấm ép mẫu | g |  |  |  |  | MASS |  | 0.1 | 800000 | Service & Verification | 0.01 | 0.1 | Perspi-MASS
Perspi | [8 empty] | 910 | N/A |  | Khối lượng tấm ép mẫu | g | 910 | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 | 800000 | Service & Verification | 0.01 | 0.1 | Perspi-MASS
Perspi | [6 empty] | Weight | Weight (*) | 3630 | N/A | Weight (*) - g | Khối lượng | g | 3630 | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Perspi-MASS
Perspi | [7 empty] | Total compression mass - AATCC (*) | 4540 | ±100 | Total compression mass - AATCC (*) - g | Tổng khối lượng - AATCC | g | 4540 | 100 | 4440 | 4640 | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Perspi-MASS
Perspi | [7 empty] | Total compression mass - ISO 105 | 5000 | ±100 | Total compression mass - ISO 105 (*) - g | Tổng khối lượng - ISO 105 | g | 5000 | 100 | 4900 | 5100 | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Perspi-MASS
Perspi | [6 empty] | Weight | Top plates mass (*) | Top | N/A | Top plates mass (*) - g | Khối lượng tấm ép mẫu | g | N/A | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Perspi-MASS
Perspi |  |  |  | 10: Bottom; 12: N/A; 14: g; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Perspi-MASS |  |  | Bottom
Perspi |  |  |  | 10: Total mass; 12: N/A; 14: g; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Perspi-MASS |  |  | Total mass
Perspi |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Perspi- |  |  | *
Perspi | [27 empty] | * | bar: 11
pH | ĐLVN 31: 2017 | PH meter | TAV-QT.18-Ph |  | PH meter | Máy đo pH | pH value | - pH value | 4.01 | ±0.05 | - pH value - pH | Giá trị pH | pH | 4.01 | 0.05 | 3.96 | 4.06 | pH |  | 0.01 | 600000 | Service & Verification | 0 | 0 | pH-pH |  |  | 4.01 - - pH value - pH
pH |  |  |  | 10: 7.01; 12: ±0.05; 14: pH; 15: 7.01; 16: 0.05; 17: 6.96; 18: 7.06; 19: pH; 21: 0.01 |  | 0 | 0 | pH-pH |  |  | 7.01
pH |  |  |  | 10: 10.01; 12: ±0.05; 14: pH; 15: 10.01; 16: 0.05; 17: 9.96; 18: 10.06; 19: pH; 21: 0.01 |  | 0 | 0 | pH-pH
pH |  |  |  | 19: pH; 21: 0.01 |  | 0 | 0 | pH-pH
pH |  |  |  | 19: pH; 21: 0.01 |  | 0 | 0 | pH-pH
pH |  |  |  | 19: pH; 21: 0.01 |  | 0 | 0 | pH-pH
pH | [7 empty] | Temperature - ⁰C | 4.01 | N/A | Temperature (*) - ⁰C | Nhiệt độ | ℃ |  | N/A | #VALUE! | #VALUE! | DEG |  | 0.2 |  |  | 0.1 | 0.2 | pH-DEG
pH |  |  |  | 10: 7.01; 12: N/A; 14: ℃; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | pH-DEG
pH |  |  |  | 10: 10.01; 12: N/A; 14: ℃; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | pH-DEG
pH |  |  |  | 14: ℃; 19: DEG; 20: 0.01; 21: 0.2 |  | 0.1 | 0.2 | pH-DEG
pH |  |  |  | 14: ℃; 19: DEG; 20: 0.01; 21: 0.2 |  | 0.1 | 0.2 | pH-DEG
pH |  |  |  | 14: ℃; 19: DEG; 21: 0
pH |  |  |  | 9: mV; 10: 1.68; 12: +255 to +375; 11: mV; 14: mV; 15: 315; 16: 60; 17: 255; 18: 375; 19: pH
pH |  |  |  | 10: 4.01; 12: +117 to +237; 14: mV; 15: 177; 16: 60; 17: 117; 18: 237; 19: pH
pH |  |  |  | 10: 6.86; 12: -52 to +68; 14: mV; 15: 8; 16: 60; 17: -52; 18: 68; 19: pH
pH |  |  |  | 10: 7.01; 12: -60 to +60; 14: mV; 15: 0; 16: 60; 17: -60; 18: 60; 19: pH
pH |  |  |  | 10: 9.18; 12: -189 to -69; 14: mV; 15: -158; 16: 31; 17: -189; 18: -69; 19: pH
pH |  |  |  | 10: 10.01 -; 12: -237 to -117; 14: mV; 15: -177; 16: 60; 17: -237; 18: -117; 19: pH
pH |  |  |  | 10: 12.46; 12: -383 to -263; 14: mV; 15: -323; 16: 60; 17: -383; 18: -263; 19: pH
pH
pH |  |  |  | 9: Slope - %; 10: 100; 12: 85-110; 11: Slope - %; 14: %; 15: 97.5; 16: 12.5; 17: 85; 18: 110; 19: pH
pH |  |  |  | 10: 100; 12: 85-110; 14: %; 15: 97.5; 16: 12.5; 17: 85; 18: 110; 19: pH
Pill-snagging | ISO 12945‐1, M&S P18A, P18B, P21A | ICI Pilling & Snagging | TAV-QTHTD TAV-QT-PILL1 | Orbitor, Rotapill | ICI Pilling & Snagging Tester | Thiết bị thử độ xù lông và xước móc | Speed Reversing Action Counter Box/Drum Dimension Pin height | Speed (*) - rpm | 30 | ±2 | Speed (*) - rpm | Tốc độ | rpm | 30 | 2 | 28 | 32 | RPM |  | 0.1 | 2000000 | Service & Calibration | 0.1 | 0.1 | Pill-snagging-RPM |  |  | 30 - Speed (*) - rpm
Pill-snagging |  |  |  | 10: 30; 12: ±2; 14: rpm; 15: 30; 16: 2; 17: 28; 18: 32; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Pill-snagging-RPM |  |  | 30
Pill-snagging |  |  |  | 10: 60; 12: ±2; 14: rpm; 15: 60; 16: 2; 17: 58; 18: 62; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Pill-snagging-RPM |  |  | 60
Pill-snagging |  |  |  | 10: 60; 12: ±2; 14: rpm; 15: 60; 16: 2; 17: 58; 18: 62; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Pill-snagging-RPM |  |  | 60
Pill-snagging | [7 empty] | Reversing Action after 50 cycles - Cycle | 50 | ±1 | Reversing Action after 50 cycles - Cycle | Đảo chiều sau 50 vòng | Cycle | 50 | 1 | 49 | 51 | SPM |  | 0.02 |  |  | 0.001 | 0.02 | Pill-snagging-SPM |  |  | 50 - Reversing Action after 50 cycles - Cycle
Pill-snagging | [7 empty] | Counter - Cycle | 100 | ±1 | Counter - Cycle | Bộ đếm | Cycle | 100 | 1 | 99 | 101 | TIM1 |  | 0.02 |  |  | 0.001 | 0.02 | Pill-snagging-TIM1 |  |  | 100 - Counter - Cycle
Pill-snagging |  |  |  | 9: Box/Drum Dimensions (before lining) - mm; 10: Box/Drum 1 - W - 229; 12: N/A; 11: Box/Drum Dimensions (before lining) - mm; 13: Kích thước hộp; 14: mm; 15: 229; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 1 - W - 229 - Box/Drum Dimensions (before lining) - mm
Pill-snagging |  |  |  | 10: Box/Drum 1 - H - 229; 12: N/A; 14: mm; 15: 229; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 1 - H - 229
Pill-snagging |  |  |  | 10: Box/Drum 1 - D - 229; 12: N/A; 14: mm; 15: 229; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 1 - D - 229
Pill-snagging |  |  |  | 10: Box/Drum 2 - W - 229; 12: N/A; 14: mm; 15: 229; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 2 - W - 229
Pill-snagging |  |  |  | 10: Box/Drum 2 - H - 229; 12: N/A; 14: mm; 15: 229; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 2 - H - 229
Pill-snagging |  |  |  | 10: Box/Drum 2 - D - 229; 12: N/A; 14: mm; 15: 229; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 2 - D - 229
Pill-snagging |  |  |  | 10: Box/Drum 3 - W - 229; 12: N/A; 14: mm; 15: 229; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 3 - W - 229
Pill-snagging |  |  |  | 10: Box/Drum 3 - H - 229; 12: N/A; 14: mm; 15: 229; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 3 - H - 229
Pill-snagging |  |  |  | 10: Box/Drum 3 - D - 229; 12: N/A; 14: mm; 15: 229; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 3 - D - 229
Pill-snagging |  |  |  | 10: Box/Drum 4 - W - 229; 12: N/A; 14: mm; 15: 229; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 4 - W - 229
Pill-snagging |  |  |  | 10: Box/Drum 4 - H - 229; 12: N/A; 14: mm; 15: 229; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 4 - H - 229
Pill-snagging |  |  |  | 10: Box/Drum 4 - D - 229; 12: N/A; 14: mm; 15: 229; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 4 - D - 229
Pill-snagging | [7 empty] | Pin height - mm | Box/Drum 1 - 10 | ±0.5 | Pin height - mm | Chiều cao đinh | mm | 10 | 0.5 | 9.5 | 10.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 1 - 10 - Pin height - mm
Pill-snagging |  |  |  | 10: Box/Drum 1 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 1 - 10
Pill-snagging |  |  |  | 10: Box/Drum 1 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 1 - 10
Pill-snagging |  |  |  | 10: Box/Drum 1 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 1 - 10
Pill-snagging |  |  |  | 10: Box/Drum 1 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 1 - 10
Pill-snagging |  |  |  | 10: Box/Drum 1 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 1 - 10
Pill-snagging |  |  |  | 10: Box/Drum 2 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM |  |  | Box/Drum 2 - 10
Pill-snagging |  |  |  | 10: Box/Drum 2 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 2 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 2 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 2 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 2 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 3 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 3 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 3 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 3 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 3 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 3 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 3 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 4 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 4 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 4 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 4 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 4 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 4 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagging |  |  |  | 10: Box/Drum 4 - 10; 12: ±0.5; 14: mm; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagging-DIM
Pill-snagpod | BS 8479 | ICI Pilling & Snagging | TAV-QTHTD TAV-QT-PILL2 | Snagpod | ICI Pilling & Snagging Tester | Thiết bị thử độ xù lông và xước móc | Speed Reversing Action Counter Dimension between opposite parallel faces Depth Number of pin Pin Distance Pin angle Pin height | Speed (*) - rpm | 30 | ±2 | Speed (*) - rpm | Tốc độ | rpm | 30 | 2 | 28 | 32 | RPM |  | 0.1 | 2000000 | Service & Calibration | 0.1 | 0.1 | Pill-snagpod-RPM
Pill-snagpod |  |  |  | 10: 60; 12: ±2; 14: rpm; 15: 60; 16: 2; 17: 58; 18: 62; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Pill-snagpod-RPM
Pill-snagpod | [7 empty] | Reversing Action after 50 cycles - Cycle | 50 cycles | ± 1 | Reversing Action after 50 cycles - Cycle | Đảo chiều sau 50 vòng | Cycle | 50 | 1 | 49 | 51 | SPM |  | #N/A |  |  | #N/A | #N/A | Pill-snagpod-SPM
Pill-snagpod | [7 empty] | Counter - Cycle | 100 cycles | ±1 | Counter - Cycle | Bộ đếm | Cycle | 100 | 1 | 99 | 101 | TIM1 |  | #N/A |  |  | #N/A | #N/A | Pill-snagpod-TIM1
Pill-snagpod | [7 empty] | Dimension between opposite parallel faces - mm | Snagpod 1 - 224.5 | ±0.3 | Dimension between opposite parallel faces - mm | Khoảng cách 2 mặt đối diện | mm | 224.5 | 0.3 | 224.2 | 224.8 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 1 - 224.5; 12: ±0.3; 14: mm; 15: 224.5; 16: 0.3; 17: 224.2; 18: 224.8; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 1 - 224.5; 12: ±0.3; 14: mm; 15: 224.5; 16: 0.3; 17: 224.2; 18: 224.8; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 1 - 224.5; 12: ±0.3; 14: mm; 15: 224.5; 16: 0.3; 17: 224.2; 18: 224.8; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 2 - 224.5; 12: ±0.3; 14: mm; 15: 224.5; 16: 0.3; 17: 224.2; 18: 224.8; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 2 - 224.5; 12: ±0.3; 14: mm; 15: 224.5; 16: 0.3; 17: 224.2; 18: 224.8; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 2 - 224.5; 12: ±0.3; 14: mm; 15: 224.5; 16: 0.3; 17: 224.2; 18: 224.8; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 2 - 224.5; 12: ±0.3; 14: mm; 15: 224.5; 16: 0.3; 17: 224.2; 18: 224.8; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 3 - 224.5; 12: ±0.3; 14: mm; 15: 224.5; 16: 0.3; 17: 224.2; 18: 224.8; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 3 - 224.5; 12: ±0.3; 14: mm; 15: 224.5; 16: 0.3; 17: 224.2; 18: 224.8; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 3 - 224.5; 12: ±0.3; 14: mm; 15: 224.5; 16: 0.3; 17: 224.2; 18: 224.8; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 3 - 224.5; 12: ±0.3; 14: mm; 15: 224.5; 16: 0.3; 17: 224.2; 18: 224.8; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 4 - 224.5; 12: ±0.3; 14: mm; 15: 224.5; 16: 0.3; 17: 224.2; 18: 224.8; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 4 - 224.5; 12: ±0.3; 14: mm; 15: 224.5; 16: 0.3; 17: 224.2; 18: 224.8; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 4 - 224.5; 12: ±0.3; 14: mm; 15: 224.5; 16: 0.3; 17: 224.2; 18: 224.8; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 4 - 224.5; 12: ±0.3; 14: mm; 15: 224.5; 16: 0.3; 17: 224.2; 18: 224.8; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod | [7 empty] | Depth - mm | Snagpod 1 - 228.0 | ±0.1 | Depth - mm | Chiều sâu | mm | 228 | 0.1 | 227.9 | 228.1 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 2 - 228.0; 12: ±0.1; 14: mm; 15: 228; 16: 0.1; 17: 227.9; 18: 228.1; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 3 - 228.0; 12: ±0.1; 14: mm; 15: 228; 16: 0.1; 17: 227.9; 18: 228.1; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 4 - 228.0; 12: ±0.1; 14: mm; 15: 228; 16: 0.1; 17: 227.9; 18: 228.1; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod | [7 empty] | Pin angle - ⁰ | Snagpod 1 - 60 | ±5 | Pin angle - ⁰ | Góc nghiêng của đinh | ⁰ | 60 | 5 | 55 | 65 | ANG |  | 0.03 |  |  | 0.05 | 0.03 | Pill-snagpod-ANG
Pill-snagpod |  |  |  | 10: Snagpod 1 - 60; 12: ±5; 14: ⁰; 15: 60; 16: 5; 17: 55; 18: 65; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Pill-snagpod-ANG
Pill-snagpod |  |  |  | 10: Snagpod 1 - 60; 12: ±5; 14: ⁰; 15: 60; 16: 5; 17: 55; 18: 65; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Pill-snagpod-ANG |  |  | Snagpod 1 - 60
Pill-snagpod |  |  |  | 10: Snagpod 1 - 60; 12: ±5; 14: ⁰; 15: 60; 16: 5; 17: 55; 18: 65; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Pill-snagpod-ANG |  |  | Snagpod 1 - 60
Pill-snagpod |  |  |  | 10: Snagpod 2 - 60; 12: ±5; 14: ⁰; 15: 60; 16: 5; 17: 55; 18: 65; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Pill-snagpod-ANG |  |  | Snagpod 2 - 60
Pill-snagpod |  |  |  | 10: Snagpod 2 - 60; 12: ±5; 14: ⁰; 15: 60; 16: 5; 17: 55; 18: 65; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Pill-snagpod-ANG |  |  | Snagpod 2 - 60
Pill-snagpod |  |  |  | 10: Snagpod 2 - 60; 12: ±5; 14: ⁰; 15: 60; 16: 5; 17: 55; 18: 65; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Pill-snagpod-ANG |  |  | Snagpod 2 - 60
Pill-snagpod |  |  |  | 10: Snagpod 2 - 60; 12: ±5; 14: ⁰; 15: 60; 16: 5; 17: 55; 18: 65; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Pill-snagpod-ANG |  |  | Snagpod 2 - 60
Pill-snagpod |  |  |  | 10: Snagpod 3 - 60; 12: ±5; 14: ⁰; 15: 60; 16: 5; 17: 55; 18: 65; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Pill-snagpod-ANG |  |  | Snagpod 3 - 60
Pill-snagpod |  |  |  | 10: Snagpod 3 - 60; 12: ±5; 14: ⁰; 15: 60; 16: 5; 17: 55; 18: 65; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Pill-snagpod-ANG |  |  | Snagpod 3 - 60
Pill-snagpod |  |  |  | 10: Snagpod 3 - 60; 12: ±5; 14: ⁰; 15: 60; 16: 5; 17: 55; 18: 65; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Pill-snagpod-ANG |  |  | Snagpod 3 - 60
Pill-snagpod |  |  |  | 10: Snagpod 3 - 60; 12: ±5; 14: ⁰; 15: 60; 16: 5; 17: 55; 18: 65; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Pill-snagpod-ANG |  |  | Snagpod 3 - 60
Pill-snagpod |  |  |  | 10: Snagpod 4 - 60; 12: ±5; 14: ⁰; 15: 60; 16: 5; 17: 55; 18: 65; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Pill-snagpod-ANG |  |  | Snagpod 4 - 60
Pill-snagpod |  |  |  | 10: Snagpod 4 - 60; 12: ±5; 14: ⁰; 15: 60; 16: 5; 17: 55; 18: 65; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Pill-snagpod-ANG |  |  | Snagpod 4 - 60
Pill-snagpod |  |  |  | 10: Snagpod 4 - 60; 12: ±5; 14: ⁰; 15: 60; 16: 5; 17: 55; 18: 65; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Pill-snagpod-ANG |  |  | Snagpod 4 - 60
Pill-snagpod |  |  |  | 10: Snagpod 4 - 60; 12: ±5; 14: ⁰; 15: 60; 16: 5; 17: 55; 18: 65; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Pill-snagpod-ANG |  |  | Snagpod 4 - 60
Pill-snagpod | [7 empty] | Pin length - mm | Snagpod 1 - 5 | ±0.3 | Pin length - mm | Chiều dài đinh | mm | 5 | 0.3 | 4.7 | 5.3 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Pill-snagpod-DIM |  |  | Snagpod 1 - 5 - Pin length - mm
Pill-snagpod |  |  |  | 10: Snagpod 1 - 5; 12: ±0.3; 14: mm; 15: 5; 16: 0.3; 17: 4.7; 18: 5.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM |  |  | Snagpod 1 - 5
Pill-snagpod |  |  |  | 10: Snagpod 1 - 5; 12: ±0.3; 14: mm; 15: 5; 16: 0.3; 17: 4.7; 18: 5.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM |  |  | Snagpod 1 - 5
Pill-snagpod |  |  |  | 10: Snagpod 1 - 5; 12: ±0.3; 14: mm; 15: 5; 16: 0.3; 17: 4.7; 18: 5.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM |  |  | Snagpod 1 - 5
Pill-snagpod |  |  |  | 10: Snagpod 2 - 5; 12: ±0.3; 14: mm; 15: 5; 16: 0.3; 17: 4.7; 18: 5.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM |  |  | Snagpod 2 - 5
Pill-snagpod |  |  |  | 10: Snagpod 2 - 5; 12: ±0.3; 14: mm; 15: 5; 16: 0.3; 17: 4.7; 18: 5.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM |  |  | Snagpod 2 - 5
Pill-snagpod |  |  |  | 10: Snagpod 2 - 5; 12: ±0.3; 14: mm; 15: 5; 16: 0.3; 17: 4.7; 18: 5.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM |  |  | Snagpod 2 - 5
Pill-snagpod |  |  |  | 10: Snagpod 2 - 5; 12: ±0.3; 14: mm; 15: 5; 16: 0.3; 17: 4.7; 18: 5.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM |  |  | Snagpod 2 - 5
Pill-snagpod |  |  |  | 10: Snagpod 3 - 5; 12: ±0.3; 14: mm; 15: 5; 16: 0.3; 17: 4.7; 18: 5.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 3 - 5; 12: ±0.3; 14: mm; 15: 5; 16: 0.3; 17: 4.7; 18: 5.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 3 - 5; 12: ±0.3; 14: mm; 15: 5; 16: 0.3; 17: 4.7; 18: 5.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 3 - 5; 12: ±0.3; 14: mm; 15: 5; 16: 0.3; 17: 4.7; 18: 5.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 4 - 5; 12: ±0.3; 14: mm; 15: 5; 16: 0.3; 17: 4.7; 18: 5.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 4 - 5; 12: ±0.3; 14: mm; 15: 5; 16: 0.3; 17: 4.7; 18: 5.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 4 - 5; 12: ±0.3; 14: mm; 15: 5; 16: 0.3; 17: 4.7; 18: 5.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-snagpod |  |  |  | 10: Snagpod 4 - 5; 12: ±0.3; 14: mm; 15: 5; 16: 0.3; 17: 4.7; 18: 5.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Pill-snagpod-DIM
Pill-random | ASTM D3512, ISO 12945-3, JIS L1076, GB/T 4802.4 |  | TAV-QTHTD TAV-QTHAP | Impulse | Random Tumble Pilling Tester | Thiết bị thử độ xù lông | Speed Impeller Position Time Air pressure | Speed (*) | 1200 | ±20 | Speed (*) - rpm | Tốc độ | rpm | 1200 | 20 | 1180 | 1220 | RPM |  | 0.1 | 1500000 | Service & Calibration | 0.1 | 0.1 | Pill-random-RPM
Pill-random |  |  |  | 10: 1200; 12: ±20; 14: rpm; 15: 1200; 16: 20; 17: 1180; 18: 1220; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Pill-random-RPM
Pill-random |  |  |  | 10: 1200; 12: ±20; 14: rpm; 15: 1200; 16: 20; 17: 1180; 18: 1220; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Pill-random-RPM
Pill-random |  |  |  | 10: 1200; 12: ±20; 14: rpm; 15: 1200; 16: 20; 17: 1180; 18: 1220; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Pill-random-RPM
Pill-random |  |  |  | 9: Impeller Position; 10: Visual check; 11: Impeller Position; 13: Vị trí Impeller; 14: pos; 17: 0; 18: 0; 19: POS; 21: #N/A |  | #N/A | #N/A | Pill-random-POS
Pill-random | [7 empty] | Time | 1800 | N/A | Time - sec | Thời gian | sec | 1800 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Pill-random-TIME
Pill-random |  |  |  | 10: 3600; 12: N/A; 14: sec; 15: 3600; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Pill-random-TIME
Pill-random | [7 empty] | Air pressure (*) | 1 | N/A | Air pressure - psi | Áp suất | psi | 1 | N/A | #VALUE! | #VALUE! | PRESS |  | 0.1 |  |  | 0.05 | 0.1 | Pill-random-PRESS
Pill-random |  |  |  | 10: 2; 12: N/A; 14: psi; 15: 2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Pill-random-PRESS
Pill-random |  |  |  | 10: 3; 12: N/A; 14: psi; 15: 3; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Pill-random-PRESS
Pill-random |  |  |  | 10: 4; 12: N/A; 14: psi; 15: 4; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Pill-random-PRESS
Pill-random |  |  |  | 10: 5; 12: N/A; 14: psi; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: PRESS; 21: 0.1 |  | 0.05 | 0.1 | Pill-random-PRESS
Ruler10 |  | Ruler | TAV-QT-DIM | 5: Ruler >1m; 6: Ruler; 7: Thước thẳng; 9: Dimension of Devided Fractions - mm; 10: …....... mm; 11: Dimension of Devided Fractions - mm; 13: Kích thước các khoảng chia; 21: #N/A; 22: 400000 | Service & Verification | #N/A | #N/A | Ruler10-
Ruler10 |  |  |  | 10: 1 -; 12: N/A; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 2 -; 12: N/A; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 3 -; 12: N/A; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 4 -; 12: N/A; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 5 -; 12: N/A; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 6 -; 12: N/A; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 7 -; 12: N/A; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 8 -; 12: N/A; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 9 -; 12: N/A; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 10 -; 12: N/A; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 9: Dimension of Devided Fractions - (inch/feet/….........) Kích thước các khoảng chia; 10: …...... (inch/feet/…....); 11: *; 21: #N/A |  | #N/A | #N/A | Ruler10-
Ruler10 |  |  |  | 10: 1 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 2 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 3 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 4 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 5 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 6 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 7 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 8 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 9 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler10 |  |  |  | 10: 10 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler10-DIM
Ruler5 |  | Ruler | TAV-QT-DIM | 6: Ruler; 7: Thước thẳng; 9: Dimension of Devided Fractions - mm; 10: …....... mm; 11: Dimension of Devided Fractions - mm; 13: Kích thước các khoảng chia; 14: mm; 21: #N/A; 22: 400000 | Service & Verification | #N/A | #N/A | Ruler5-
Ruler5 |  |  |  | 10: 1 -; 12: N/A; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler5-DIM
Ruler5 |  |  |  | 10: 2 -; 12: N/A; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler5-DIM
Ruler5 |  |  |  | 10: 3 -; 12: N/A; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler5-DIM
Ruler5 |  |  |  | 10: 4 -; 12: N/A; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler5-DIM
Ruler5 |  |  |  | 10: 5 -; 12: N/A; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler5-DIM
Ruler5 |  |  |  | 9: Dimension of Devided Fractions - (inch/feet/….........) Kích thước các khoảng chia; 10: …...... (inch/feet/…....); 11: *; 21: #N/A |  | #N/A | #N/A | Ruler5-
Ruler5 |  |  |  | 10: 1 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler5-DIM
Ruler5 |  |  |  | 10: 2 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler5-DIM
Ruler5 |  |  |  | 10: 3 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler5-DIM
Ruler5 |  |  |  | 10: 4 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler5-DIM
Ruler5 |  |  |  | 10: 5 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler5-DIM
Ruler-Eke |  |  | TAV-QT-DIM | 5: Eke Ruler; 6: Eke Ruler; 7: Thước góc; 9: Dimension of Devided Fractions - Edge 1 - (inch/feet/….........) Kích thước các khoảng chia - Cạnh 1; 10: …...... (inch/feet/…....); 11: *; 13: Kích thước các khoảng chia - Cạnh 1; 21: #N/A; 22: 600000 | Service & Verification | #N/A | #N/A | Ruler-Eke-
Ruler-Eke |  |  |  | 10: 1 -; 12: N/A; 19: DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 2 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 3 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 4 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 5 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 6 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 7 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 8 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 9 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 10 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 9: Dimension of Devided Fractions - Edge 2 - (inch/feet/….........) Kích thước các khoảng chia - Cạnh 2; 10: …...... (inch/feet/…....); 12: N/A; 11: *; 13: Kích thước các khoảng chia - Cạnh 2; 21: #N/A |  | #N/A | #N/A | Ruler-Eke-
Ruler-Eke |  |  |  | 10: 1 -; 12: N/A; 19: DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 2 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 3 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 4 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 5 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 6 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 7 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 8 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 9 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 10 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 9: Dimension of Devided Fractions - Edge 3 - (inch/feet/….........) Kích thước các khoảng chia - Cạnh 3; 10: …...... (inch/feet/…....); 12: N/A; 11: *; 13: Kích thước các khoảng chia - Cạnh 3; 21: #N/A |  | #N/A | #N/A | Ruler-Eke-
Ruler-Eke |  |  |  | 10: 1 -; 12: N/A; 19: DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 2 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 3 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 4 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 5 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 6 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 7 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 8 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 9 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 10 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 9: Dimension of Devided Fractions - Edge 4 - (inch/feet/….........) Kích thước các khoảng chia - Cạnh 4; 10: …...... (inch/feet/…....); 12: N/A; 11: *; 13: Kích thước các khoảng chia - Cạnh 4; 21: #N/A |  | #N/A | #N/A | Ruler-Eke-
Ruler-Eke |  |  |  | 10: 1 -; 12: N/A; 19: DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 2 -; 12: N/A; 14: mm; 15: 100; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 3 -; 12: N/A; 14: mm; 15: 100; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 4 -; 12: N/A; 14: mm; 15: 100; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 5 -; 12: N/A; 14: mm; 15: 100; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 6 -; 12: N/A; 14: mm; 15: 100; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 7 -; 12: N/A; 14: mm; 15: 100; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 8 -; 12: N/A; 14: mm; 15: 100; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 9 -; 12: N/A; 14: mm; 15: 100; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 10: 10 -; 12: N/A; 14: mm; 15: 100; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-Eke-DIM
Ruler-Eke |  |  |  | 6: Eke Ruler; 9: Angle - ⁰; 10: 90⁰; 12: N/A; 11: Angle - ⁰; 13: Góc; 14: ⁰; 15: 90; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Ruler-Eke-ANG
Template-shrinkage | AATCC/ISO Shrinkage, Appreance Change | Template | TAV-QT-DIM | Shrinkage Template, Perspex | Shrinkage Template | Thước đo độ co rút của vải |  | Dimension 1 | 10 inch | N/A | Dimension 1 - Inch | Kích thước các khoảng chia - Thước 1 | Inch | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 350 mm; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 15 inch; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 18 inch; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 500 mm; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  |  | Shrinkage Ruler | Thước đo độ co rút của vải |  | Dimension 2 | 10 inch | N/A | Dimension 2 - Inch | Kích thước các khoảng chia - Thước 2 | Inch | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 350 mm; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 15 inch; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 18 inch; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 500 mm; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage | [7 empty] | Dimension 3 | 10 inch | N/A | Dimension 3 - Inch | Kích thước các khoảng chia - Thước 3 | Inch | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 350 mm; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 15 inch; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 18 inch; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 500 mm; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage | [7 empty] | Dimension 4 | 10 inch | N/A | Dimension 4 - Inch | Kích thước các khoảng chia - Thước 4 | Inch | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 350 mm; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 15 inch; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 18 inch; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 500 mm; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage | [7 empty] | 10 inch % Shrinkage | 0-5 | N/A | 10 inch % Shrinkage - inch |  | inch | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 5-10; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 10-15; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 15-20; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage | [7 empty] | 10 inch % Stretch | 0-5 | N/A | 10 inch % Stretch - inch |  | inch | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 5-10; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 10-15; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 15-20; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | Universal Shrinkage Percentage | Universal Shrinkage Percentage (%) Ruler | Thước đo độ co rút của vải |  | Dimension of Devided Fractions | 0-10 inch | N/A | Dimension of Devided Fractions - inch | Kích thước các khoảng chia | inch | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 0-350 mm; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 0-18 inch; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 0-500 mm; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage | [7 empty] | 10 inch % Shrinkage | 0-5 | N/A | 10 inch % Shrinkage - inch |  | inch | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 5-10; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 10-15; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 15-20; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM |  |  | 15-20
Template-shrinkage | [7 empty] | 10 inch % Stretch | 0-5 | N/A | 10 inch % Stretch - inch |  | inch | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-shrinkage-DIM |  |  | 0-5 - 10 inch % Stretch - inch
Template-shrinkage |  |  |  | 10: 5-10; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM |  |  | 5-10
Template-shrinkage |  |  |  | 10: 10-15; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM |  |  | 10-15
Template-shrinkage |  |  |  | 10: 15-20; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM |  |  | 15-20
Template-shrinkage | [7 empty] | 350mm % Shrinkage | 0-5 | N/A | 350mm % Shrinkage - mm |  | mm | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-shrinkage-DIM |  |  | 0-5 - 350mm % Shrinkage - mm
Template-shrinkage |  |  |  | 10: 5-10; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM |  |  | 5-10
Template-shrinkage |  |  |  | 10: 10-15; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM |  |  | 10-15
Template-shrinkage |  |  |  | 10: 15-20; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM |  |  | 15-20
Template-shrinkage | [7 empty] | 350mm % Stretch | 0-5 | N/A | 350mm % Stretch - mm |  | mm | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-shrinkage-DIM |  |  | 0-5 - 350mm % Stretch - mm
Template-shrinkage |  |  |  | 10: 5-10; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 10-15; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 15-20; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage | [7 empty] | 18 inch % Shrinkage | 0-5 | N/A | 18 inch % Shrinkage - inch |  | inch | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 5-10; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 10-15; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 15-20; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage | [7 empty] | 18 inch % Stretch | 0-5 | N/A | 18 inch % Stretch - inch |  | inch | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 5-10; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 10-15; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 15-20; 12: N/A; 14: inch; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage | [7 empty] | 500mm % Shrinkage | 0-5 | N/A | 500mm % Shrinkage - mm |  | mm | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 5-10; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 10-15; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 15-20; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage | [7 empty] | 500mm % Stretch | 0-5 | N/A | 500mm % Stretch - mm |  | mm | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 5-10; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM |  |  | 5-10
Template-shrinkage |  |  |  | 10: 10-15; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template-shrinkage |  |  |  | 10: 15-20; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-shrinkage-DIM
Template | ASTM D2261, ASTM D1424, ASTM D5034, ASTM D1683, ISO 13937-1, ISO 13935-2 | Template | TAV-QT-DIM | Tearing Test Template for ASTM D2261 | Tearing Test Templete for ASTM D2261 |  |  | Template 1 | Template 1 | N/A | Template 1 - mm | Kích thước các khoảng chia | mm | 200 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-DIM
Template |  |  |  | 12: N/A; 14: mm; 15: 75; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-DIM
Template |  |  |  | 12: N/A; 14: mm; 15: 75; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-DIM
Template |  |  |  | 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-DIM
Template |  |  |  | 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-DIM
Template | [7 empty] | Template 2 | Template 2 | N/A | Template 2 - mm | Kích thước các khoảng chia | mm | 200 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-DIM
Template |  |  |  | 12: N/A; 14: mm; 15: 75; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-DIM |  |  | 0
Template |  |  |  | 12: N/A; 14: mm; 15: 75; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-DIM |  |  | 0
Template |  |  |  | 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-DIM |  |  | 0
Template |  |  |  | 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-DIM |  |  | 0
Template | [7 empty] | Template 3 | Template 3 | N/A | Template 3 - mm | Kích thước các khoảng chia | mm | 102 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Template-DIM |  |  | Template 3 - Template 3 - mm
Template |  |  |  | 12: N/A; 14: mm; 15: 63; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-DIM |  |  | 0
Template |  |  |  | 12: N/A; 14: mm; 15: 75; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-DIM |  |  | 0
Template |  |  |  | 12: N/A; 14: mm; 15: 63; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-DIM |  |  | 0
Template |  |  |  | 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-DIM |  |  | 0
Template |  |  |  | 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-DIM |  |  | 0
Template |  |  |  | 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-DIM |  |  | 0
Template |  |  |  | 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-DIM
Template |  |  |  | 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Template-DIM |  |  | 0
Ruler-T |  | T Ruler | TAV-QT-DIM |  | T Ruler | Thước chữ T |  | Dimension of Devided Fractions | Templete 1 | N/A | Dimension of Devided Fractions - cm | Kích thước các khoảng chia | cm | 10 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 | 400000 | Service & Verification | 0.01 | 0.02 | Ruler-T-DIM |  |  | Templete 1 - Dimension of Devided Fractions - cm
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 20; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 30; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 40; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 50; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 60; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 10: Templete 2; 12: N/A; 14: cm; 15: 1; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | Templete 2
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 3; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 4; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 6; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 7; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 8; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 9; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 10; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 10: Templete 3; 12: N/A; 14: cm; 15: 2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | Templete 3
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 4; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 6; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 8; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 10; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 12; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM |  |  | 0
Ruler-T |  |  |  | 12: N/A; 14: cm; 15: 14; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ruler-T-DIM
Flexiframe |  | FlexiFrame | TAV-QT-DIM | 5: FlexiFrame, Stretch Recovery Tester; 6: FlexiFrame; 7: Thiết bị kiểm tra độ giãn dài; 9: Weight 1 - (kg/lb) Tải trọng; 10: ..........; 11: *; 13: Tải trọng; 21: #N/A |  | #N/A | #N/A | Flexiframe- |  |  | .......... - *
Flexiframe |  |  |  | 10: 1 -; 12: N/A; 17: 0; 18: 0; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Flexiframe-MASS |  |  | 1 -
Flexiframe |  |  |  | 10: 2 -; 12: N/A; 17: 0; 18: 0; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Flexiframe-MASS |  |  | 2 -
Flexiframe |  |  |  | 10: 3 -; 12: N/A; 17: 0; 18: 0; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Flexiframe-MASS |  |  | 3 -
Flexiframe |  |  |  | 9: Weight 2 - (kg/lb) Tải trọng; 10: ..........; 11: *; 21: #N/A |  | #N/A | #N/A | Flexiframe- |  |  | .......... - *
Flexiframe |  |  |  | 10: 1 -; 12: N/A; 17: 0; 18: 0; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Flexiframe-MASS |  |  | 1 -
Flexiframe |  |  |  | 10: 2 -; 12: N/A; 17: 0; 18: 0; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Flexiframe-MASS |  |  | 2 -
Flexiframe |  |  |  | 10: 3 -; 12: N/A; 17: 0; 18: 0; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Flexiframe-MASS |  |  | 3 -
Flexiframe |  |  |  | 9: Weight 3 - (kg/lb) Tải trọng; 10: ..........; 11: *; 21: #N/A |  | #N/A | #N/A | Flexiframe- |  |  | .......... - *
Flexiframe |  |  |  | 10: 1 -; 12: N/A; 17: 0; 18: 0; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Flexiframe-MASS |  |  | 1 -
Flexiframe |  |  |  | 10: 2 -; 12: N/A; 17: 0; 18: 0; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Flexiframe-MASS |  |  | 2 -
Flexiframe |  |  |  | 10: 3 -; 12: N/A; 17: 0; 18: 0; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Flexiframe-MASS |  |  | 3 -
Flexiframe |  |  |  | 9: Ruler1 - (mm/inch) Thước 1; 10: 1 -; 12: N/A; 11: *; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM |  |  | 1 - - *
Flexiframe |  |  |  | 10: 2 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 3 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 4 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 5 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 6 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 7 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 8 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 9 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 10 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 11 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 12 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 9: Ruler2 - (mm/inch) Thước 2; 10: 1 -; 12: N/A; 11: *; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 2 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 3 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 4 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 5 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 6 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 7 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 8 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 9 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 10 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 11 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 12 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 9: Ruler3 - (mm/inch) Thước 3; 10: 1 -; 12: N/A; 11: *; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 2 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 3 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 4 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 5 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 6 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 7 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 8 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 9 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 10 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 11 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe |  |  |  | 10: 12 -; 12: N/A; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Flexiframe-DIM
Flexiframe | [7 empty] | Times1 | 30 | N/A | Times1 - sec | Thời gian | sec | 30 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | #N/A | 0.02 | Flexiframe-TIME
Flexiframe |  |  |  | 10: 30; 12: N/A; 14: sec; 15: 30; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | #N/A | 0.02 | Flexiframe-TIME
Flexiframe |  |  |  | 10: 60; 12: N/A; 14: sec; 15: 60; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | #N/A | 0.02 | Flexiframe-TIME
Flexiframe |  |  |  | 10: 60; 12: N/A; 14: sec; 15: 60; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | #N/A | 0.02 | Flexiframe-TIME
Flexiframe | [7 empty] | Times2 | 30 | N/A | Times2 - sec | Thời gian | sec | 30 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | #N/A | 0.02 | Flexiframe-TIME
Flexiframe |  |  |  | 10: 30; 12: N/A; 14: sec; 15: 30; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | #N/A | 0.02 | Flexiframe-TIME
Flexiframe |  |  |  | 10: 60; 12: N/A; 14: sec; 15: 60; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | #N/A | 0.02 | Flexiframe-TIME
Flexiframe |  |  |  | 10: 60; 12: N/A; 14: sec; 15: 60; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | #N/A | 0.02 | Flexiframe-TIME
Flexiframe | [7 empty] | Times3 | 30 | N/A | Times3 - sec | Thời gian | sec | 30 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | #N/A | 0.02 | Flexiframe-TIME
Flexiframe |  |  |  | 10: 30; 12: N/A; 14: sec; 15: 30; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | #N/A | 0.02 | Flexiframe-TIME
Flexiframe |  |  |  | 10: 60; 12: N/A; 14: sec; 15: 60; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | #N/A | 0.02 | Flexiframe-TIME |  |  | 60
Flexiframe |  |  |  | 10: 60; 12: N/A; 14: sec; 15: 60; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | #N/A | 0.02 | Flexiframe-TIME |  |  | 60
Crimp |  | Yarn Crimp Tester | TAV-QT-DIM TAV - QTHKL | Yarn Crimp Tester | Yarn Crimp Tester |  |  | Dimension of Devided Fractions - mm | 1x100 | N/A | Dimension of Devided Fractions - mm - mm | Kích thước các khoảng chia | mm | 100 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 | 400000 | Service & Verification | 0.01 | 0.02 | Crimp-DIM |  |  | 1x100 - Dimension of Devided Fractions - mm - mm
Crimp |  |  |  | 10: 2x100; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM |  |  | 2x100
Crimp |  |  |  | 10: 3x100; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM |  |  | 3x100
Crimp |  |  |  | 10: 4x100; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM |  |  | 4x100
Crimp |  |  |  | 10: 5x100; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 6x100; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 7x100; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 8x100; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 9x100; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 10x100; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 11x100; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 12x100; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 13x100; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 14x100; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 15x100; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp | [7 empty] | Dimension of Devided Fractions - inch | 1x5 | N/A | Dimension of Devided Fractions - inch - inch | Kích thước các khoảng chia | inch | 5 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 2x5; 12: N/A; 14: inch; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 3x5; 12: N/A; 14: inch; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 4x5; 12: N/A; 14: inch; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 5x5; 12: N/A; 14: inch; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 6x5; 12: N/A; 14: inch; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 7x5; 12: N/A; 14: inch; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 8x5; 12: N/A; 14: inch; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 9x5; 12: N/A; 14: inch; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 10x5; 12: N/A; 14: inch; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 11x5; 12: N/A; 14: inch; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 12x5; 12: N/A; 14: inch; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 13x5; 12: N/A; 14: inch; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 14x5; 12: N/A; 14: inch; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp |  |  |  | 10: 15x5; 12: N/A; 14: inch; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Crimp-DIM
Crimp | [7 empty] | Weight | 0-35 | N/A | Weight - g | Tải trọng | g | 35 | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Crimp-MASS
Crimp |  |  |  | 10: 0-175; 12: N/A; 14: g; 15: 175; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Crimp-MASS
Cut-sample | EN12127 | Sample Cutter | TAV-QT-DIM | Sample Cutter | Sample Cutter | Dao cắt mẫu | Diameter Area Calculation | Diameter - mm (Area - 10 cm²) | 36 | N/A | Diameter - mm (Area - 10 cm²) - mm | Đường kính | mm | 36 |  | 36 | 36 | DIM |  | 0.02 | 600000 | Service & Verification | 0.01 | 0.02 | Cut-sample-DIM
Cut-sample |  |  |  | 9: Area Calculation; 12: N/A; 11: Area Calculation - cm²; 13: Diện tích; 14: cm²; 15: 10; 17: 10; 18: 10; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-sample-DIM |  |  | - Area Calculation - cm²
Cut-sample | [6 empty] | Diameter Area Calculation | Diameter - mm (Area - 100 cm²) | 112.84 | ±0.56 | Diameter - mm (Area - 100 cm²) - mm | Đường kính | mm | 112.84 | 0.56 | 112.28 | 113.4 | DIM |  | 0.02 | 600000 | Service & Verification | 0.01 | 0.02 | Cut-sample-DIM
Cut-sample | [7 empty] | Area Calculation | 100 | ±1% | Area Calculation - cm² | Diện tích | cm² | 100 | 1 | 99 | 101 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-sample-DIM
Cut-sample | [6 empty] | Diameter Area Calculation | Diameter - mm (Area - 12 cm²) | 38 | N/A | Diameter - mm (Area - 12 cm²) - mm | Đường kính | mm | 38 |  | 38 | 38 | DIM |  | 0.02 | 600000 | Service & Verification | 0.01 | 0.02 | Cut-sample-DIM
Cut-sample |  |  |  | 9: Area Calculation; 12: N/A; 11: Area Calculation - cm²; 13: Diện tích; 14: cm²; 15: 12; 17: 12; 18: 12; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-sample-DIM
Cut-sample | [6 empty] | Diameter Area Calculation | Diameter - mm (Area - 154 cm²) | 140 | N/A | Diameter - mm (Area - 154 cm²) - mm | Đường kính | mm | 140 |  | 140 | 140 | DIM |  | 0.02 | 600000 | Service & Verification | 0.01 | 0.02 | Cut-sample-DIM
Cut-sample |  |  |  | 9: Area Calculation; 12: N/A; 11: Area Calculation - cm²; 13: Diện tích; 14: cm²; 15: 154; 17: 154; 18: 154; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-sample-DIM |  |  | Area Calculation - cm²
Cut-sample | [6 empty] | Diameter Area Calculation | Diameter - mm (Area - 50 cm²) | 80 | N/A | Diameter - mm (Area - 50 cm²) - mm | Đường kính | mm | 80 |  | 80 | 80 | DIM |  | 0.02 | 600000 | Service & Verification | 0.01 | 0.02 | Cut-sample-DIM |  |  | 80 - Diameter - mm (Area - 50 cm²) - mm
Cut-sample |  |  |  | 9: Area Calculation; 12: N/A; 11: Area Calculation - cm²; 13: Diện tích; 14: cm²; 15: 50; 17: 50; 18: 50; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-sample-DIM |  |  | Area Calculation - cm²
Cut-sample |  |  |  | 9: Diameter - mm (Area - ….... cm²); 12: N/A; 11: Diameter - mm (Area - ….... cm²) - mm; 13: Đường kính; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-sample-DIM |  |  | Diameter - mm (Area - ….... cm²) - mm
Cut-sample |  |  |  | 9: Area Calculation; 12: N/A; 11: Area Calculation - cm²; 13: Diện tích; 14: cm²; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-sample-DIM
Cut-sample |  |  |  | 9: Diameter - mm (Area - ….... cm²); 12: N/A; 11: Diameter - mm (Area - ….... cm²) - mm; 13: Đường kính; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-sample-DIM
Cut-sample |  |  |  | 9: Area Calculation; 12: N/A; 11: Area Calculation - cm²; 13: Diện tích; 14: cm²; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-sample-DIM |  |  | Area Calculation - cm²
Spray | AATCC TM 22, ISO 4920 | Spray Rating | LAB-SPRAY-01 | Spray Rating - Manual | Spray Rating | Thiết bị thử độ thấm nước | Distance to specimen Specimen holder diameter Diameter of Funnel Funnel height (include nozzle) - AATCC TM 22 Funnel height (include nozzle) - ISO 4920 Nozzle Dimension - AATCC TM 22 Nozzle Dimension - ISO 4920 Duration flow Angle Water Volume sprayed from machine Cycles per minutes Time | Distance to specimen - mm | 150 | ±2 | Distance to specimen - mm | Khoảng cách tới mẫu | mm | 150 | 2 | 148 | 152 | DIM |  | 0.02 | 1200000 | Service & Verification | 0.01 | 0.02 | Spray-DIM |  | mm | 150 - Distance to specimen - mm
Spray | [7 empty] | Specimen holder diameter - mm | 155 | ±5 | Specimen holder diameter - mm | Đường kính vòng giữ mẫu | mm | 155 | 5 | 150 | 160 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Spray-DIM |  |  | 155 - Specimen holder diameter - mm
Spray | [7 empty] | Diameter of Funnel - mm | 150 | ±5 | Diameter of Funnel - mm | Đường kính phễu | mm | 150 | 5 | 145 | 155 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Spray-DIM |  |  | 150 - Diameter of Funnel - mm
Spray | [7 empty] | Funnel height (include nozzle) - AATCC TM 22 - mm | 200 | ±10 | Funnel height (include nozzle) - AATCC TM 22 - mm | Chiều cao phễu đã lắp vòi phun - AATCC TM 22 | mm | 200 | 10 | 190 | 210 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Spray-DIM |  |  | 200 - Funnel height (include nozzle) - AATCC TM 22 - mm
Spray | [7 empty] | Funnel height (include nozzle) - ISO 4920 - mm | 190 | ±10 | Funnel height (include nozzle) - ISO 4920 - mm | Chiều cao phễu đã lắp vòi phun - ISO 4920 | mm | 190 | 10 | 180 | 200 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Spray-DIM |  |  | 190 - Funnel height (include nozzle) - ISO 4920 - mm
Spray | [7 empty] | Nozzle Dimension | 8.8 | ±0.25 | Nozzle Dimension | Kích thước vòi phun - AATCC TM 22 | mm | 8.8 | 0.25 | 8.55 | 9.05 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Spray-DIM |  |  | 8.8 - Nozzle Dimension
Spray | [7 empty] | Height of Nozzle - AATCC TM 22 -mm | 49 | ±0.38 | Height of Nozzle - AATCC TM 22 -mm | Chiều cao đầu phun | mm | 49 | 0.38 | 48.62 | 49.38 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Spray-DIM |  |  | 49 - Height of Nozzle - AATCC TM 22 -mm
Spray | [7 empty] | Diameter of head Nozzle - AATCC TM 22 -mm | 32 | ±0.5 | Diameter of head Nozzle - AATCC TM 22 -mm | Đường kính đầu vòi phun | mm | 32 | 0.5 | 31.5 | 32.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Spray-DIM |  |  | 32 - Diameter of head Nozzle - AATCC TM 22 -mm
Spray | [7 empty] | Height of Nozzle - ISO 4920 - mm | 49 | ±1 | Height of Nozzle - ISO 4920 - mm | Chiều cao đầu phun | mm | 49 | 1 | 48 | 50 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Spray-DIM |  |  | 49 - Height of Nozzle - ISO 4920 - mm
Spray | [7 empty] | Diameter of spray hole | 0.86 | ±0.05 | Diameter of spray hole | Đường kính lỗ phun | mm | 0.86 | 0.05 | 0.81 | 0.91 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Spray-DIM |  |  | 0.86 - Diameter of spray hole
Spray | [7 empty] | Diameter a circle of 6 hole | 10 | ±0.5 | Diameter a circle of 6 hole | Đường kính vòng tròn 6 lỗ | mm | 10 | 0.5 | 9.5 | 10.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Spray-DIM |  |  | 10 - Diameter a circle of 6 hole
Spray | [7 empty] | Diameter a circle of 12 hole | 21 | ±0.5 | Diameter a circle of 12 hole | Đường kính vòng tròn 12 lỗ | mm | 21 | 0.5 | 20.5 | 21.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Spray-DIM |  |  | 21 - Diameter a circle of 12 hole
Spray | [7 empty] | Duration flow with water 250ml - sec | 25 - 30 | 25 - 30 | Duration flow with water 250ml - sec | Thời gian phun với 250ml nước | sec | 27.5 | 2.5 | 25 | 30 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Spray-TIME |  |  | 25 - 30 - Duration flow with water 250ml - sec
Spray |  |  |  | 9: Angle - ⁰; 10: 45; 12: N/A; 11: Angle - ⁰; 13: Góc; 14: ⁰; 15: 45; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Spray-ANG |  |  | 45 - Angle - ⁰
Spray | [7 empty] | Water Volume sprayed from machine - ml/min | 250 | ± 25 | Water Volume sprayed from machine - ml/min | Thể tích nước được phun | g | 250 | 25 | 225 | 275 | VOL/TIME |  | 0.1 |  |  | 0.01 | 0.1 | Spray-VOL/TIME |  |  | 250 - Water Volume sprayed from machine - ml/min
Spray | [7 empty] | Outer circle - Փ - mm | 21 | ±0.5 | Outer circle - Փ - mm | Đường kính vòng ngoài |  | 21 | 0.5 | 20.5 | 21.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Spray-DIM |  |  | 21 - Outer circle - Փ - mm
Spray |  |  |  | 9: Number of holes; 10: 12; 12: N/A; 11: Number of holes; 13: Số lượng lỗ; 14: num; 15: 12; 16: N/A; 19: NUM | [6 empty] | 12 - Number of holes
Spray | [7 empty] | Inner circle - Փ - mm | 10 | ±0.5 | Inner circle - Փ - mm | Đường kính vòng trong |  | 10 | 0.5 | 9.5 | 10.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Spray-DIM |  |  | 10 - Inner circle - Փ - mm
Spray |  |  |  | 9: Number of holes; 10: 6; 12: N/A; 11: Number of holes; 14: M; 15: 6; 16: N/A; 19: NUM | [6 empty] | 60rpm - 60rpm
Spray |  |  |  | 9: Hole diameter - mm; 10: 0.86; 12: ±0.05; 11: Hole diameter - mm; 13: Đường kính lỗ; 14: mm; 15: 0.86; 16: 0.05; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Spray-DIM
Spray |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Spray-
Spray |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Spray-
Spray |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Spray-
Spray |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Spray- |  |  | *
Spray |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Spray- |  |  | *
Spray-adidas | Adidas Spray Test | Spray Rating | TAV-QT-SPR | Spray Rating - FGT-64 | Spray Rating | Thiết bị thử độ thấm nước | Water volume sprayed Cycles per minutes Timer | Water volume sprayed | 175ml/min | ±25 | Water volume sprayed - mm | Lưu lượng | mm | 150 | 2 | 148 | 152 | DIM |  | 0.02 | 800000 | Service & Verification | 0.01 | 0.02 | Spray-adidas-DIM |  |  | 175ml/min - Water volume sprayed - mm
Spray-adidas |  |  |  | 9: Cycles per minutes; 10: 12 - 15; 12: N/A; 11: Cycles per minutes; 21: #N/A |  | #N/A | #N/A | Spray-adidas- |  |  | 12 - 15 - Cycles per minutes
Spray-adidas |  |  |  | 9: Timer; 10: 60; 12: ±1; 11: Timer; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Spray-adidas-TIME |  |  | 60 - Timer
Spray-adidas | [6 empty] | Nozzle distance Flow duration | Distance from top funnel to bottom nozzle | 190 | ±25 | Distance from top funnel to bottom nozzle - mm | Khoảng cách từ đỉnh phễu đến vòi phun | mm | 150 | 2 | 148 | 152 | DIM |  | 0.02 | 800000 | Service & Verification | 0.01 | 0.02 | Spray-adidas-DIM |  |  | 60rpm - 60rpm
Spray-adidas | [6 empty] | Nozzle distance Flow duration | Distance from nozzle to specimen position | 150 | ±25 | Distance from nozzle to specimen position - mm | Khoảng cách từ vòi phun đến vị trí đặt mẫu | mm | 150 | 2 | 148 | 152 | DIM |  | 0.02 | 800000 | Service & Verification | 0.01 | 0.02 | Spray-adidas-DIM
Spray-adidas |  |  |  | 9: Flow duration; 10: 25 - 30; 11: Flow duration - sec; 13: Thời gian; 14: sec; 21: #N/A |  | #N/A | #N/A | Spray-adidas-
Spray-salt |  | Salt Spray Test Chamber (NaCl) | TAV-QTHND | 5: Salt Spray Test Chamber (NaCl); 6: Salt Spray Test Chamber (NaCl); 17: 0; 18: 0; 22: 1200000 | Service & Verification | #N/A | #N/A | Spray-salt-
Spray-salt |  |  |  | Precision Salt Sprayting Tester | Precision Salt Sprayting Tester |  |  | Chamber Temp (*) | 50 | N/A | Chamber Temp (*) - ℃ | Nhiệt độ buồng | ℃ | 50 | N/A | #VALUE! | #VALUE! | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Spray-salt-DEG
Spray-salt | [7 empty] | Tank Temp (*) | 63 | N/A | Tank Temp (*) - ℃ |  | ℃ | 63 | N/A | #VALUE! | #VALUE! | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Spray-salt-DEG
Spray-salt | [7 empty] | Time | 60 | N/A | Time - s | Thời gian | s | 60 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Spray-salt-TIME
Spray-salt |  |  |  | 10: 10; 12: N/A; 14: min; 15: 10; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Spray-salt-TIME
Spray-salt | [7 empty] | Air pressure setting | 1 | N/A | Air pressure setting - kg/cm2 |  | kg/cm2 | 1 | N/A | #VALUE! | #VALUE! | Press |  | 0.05 |  |  | 0.01 | 0.05 | Spray-salt-Press |  |  | 1 - Air pressure setting - kg/cm2
Spray-salt |  |  |  | 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Spray-salt- |  |  | 0
Spray-salt |  |  |  | 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Spray-salt- |  |  | 0
Tensile | ISO 7500 class 1 | Tensile Strength | TAV-QTHKN TAV-QT-TENSILE-STD | Titan, Instrons, Tinius Olsen, Zwick | Tensile Strength | Thiết bị thử cường lực | Jaw Parallelism Speed Extension Force | Jaw Parallelism Only for Titan | 100 | ±1 | Jaw Parallelism Only for Titan - mm | Khoảng cách giữa các ngàm | mm | 100 | 1 | 99 | 101 | DIM |  | 0.02 | 2000000 | Service & Calibration - 5 points | 0.01 | 0.02 | Tensile-DIM |  |  | 100 - Jaw Parallelism Only for Titan - mm
Tensile | [7 empty] | Speed | 50 | ±1% | Speed - mm/min | Tốc độ | mm/min | 50 | 0.5 | 49.5 | 50.5 | DIM/TIME |  | 0.02 |  |  | 0.001 | 0.02 | Tensile-DIM/TIME |  |  | 50 - Speed - mm/min
Tensile |  |  |  | 10: 100; 12: ±1%; 14: mm/min; 15: 100; 16: 1; 17: 99; 18: 101; 19: DIM/TIME; 21: 0.02 |  | 0.001 | 0.02 | Tensile-DIM/TIME |  |  | 100
Tensile |  |  |  | 10: 200; 12: ±1%; 14: mm/min; 15: 200; 16: 2; 17: 198; 18: 202; 19: DIM/TIME; 21: 0.02 |  | 0.001 | 0.02 | Tensile-DIM/TIME |  |  | 200
Tensile |  |  |  | 10: 250; 12: ±1%; 14: mm/min; 15: 250; 16: 2.5; 17: 247.5; 18: 252.5; 19: DIM/TIME; 21: 0.02 |  | 0.001 | 0.02 | Tensile-DIM/TIME |  |  | 250
Tensile |  |  |  | 10: 300; 12: ±1%; 14: mm/min; 15: 300; 16: 3; 17: 297; 18: 303; 19: DIM/TIME; 21: 0.02 |  | 0.001 | 0.02 | Tensile-DIM/TIME |  |  | 300
Tensile |  |  |  | 10: 500; 12: ±1%; 14: mm/min; 15: 500; 16: 5; 17: 495; 18: 505; 19: DIM/TIME; 21: 0.02 |  | 0.001 | 0.02 | Tensile-DIM/TIME |  |  | 500
Tensile | [7 empty] | Extension | 50 | ±1% | Extension - mm | Hành trình | mm | 50 | 0.5 | 49.5 | 50.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Tensile-DIM
Tensile |  |  |  | 10: 100; 12: ±1%; 14: mm; 15: 100; 16: 1; 17: 99; 18: 101; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tensile-DIM
Tensile |  |  |  | 10: 200; 12: ±1%; 14: mm; 15: 200; 16: 2; 17: 198; 18: 202; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tensile-DIM
Tensile |  |  |  | 10: 250; 12: ±1%; 14: mm; 15: 250; 16: 2.5; 17: 247.5; 18: 252.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tensile-DIM
Tensile |  |  |  | 10: 300; 12: ±1%; 14: mm; 15: 300; 16: 3; 17: 297; 18: 303; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tensile-DIM
Tensile |  |  |  | 10: 500; 12: ±1%; 14: mm; 15: 500; 16: 5; 17: 495; 18: 505; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Tensile-DIM
Tensile |  |  |  | 9: Force (*) - 100/200/500/1000/….. N Lực; 12: ±1%; 11: *; 13: Lực; 21: #N/A |  | #N/A | #N/A | Tensile-
Tensile |  |  |  | 10: 1 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 2 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 3 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 4 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 5 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 6 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 7 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 8 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 9 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 9: Force (*) - 100/200/500/1000/….. N Lực; 12: ±1%; 11: *; 13: Lực; 21: #N/A |  | #N/A | #N/A | Tensile-
Tensile |  |  |  | 10: 1 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 2 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 3 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 4 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 5 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 6 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 7 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 8 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 9 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 9: Force (*) - 100/200/500/1000/….. N Lực; 12: ±1%; 11: *; 13: Lực; 21: #N/A |  | #N/A | #N/A | Tensile-
Tensile |  |  |  | 10: 1 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 2 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 3 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 4 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 5 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 6 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 7 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 8 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 9 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Tensile-FORCE
Tensile |  |  |  | 10: 10 -; 12: ±1%; 14: N; 17: 0; 18: 0; 19: FORCE
Cut-TDM |  | TDM 100 Machine | TAV-QTHKL | 5: TDM 100 Machine; 6: TDM 100 Machine; 7: Thiết bị; 8: Force Position Blade Length Speed; 21: #N/A |  | #N/A | #N/A | Cut-TDM-
Cut-TDM | [7 empty] | Position | 100 - Left | N/A | Position - mm |  | mm | 100 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-TDM-DIM
Cut-TDM |  |  |  | 10: 100 - Right; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-TDM-DIM
Cut-TDM | [7 empty] | Blade Length | N/A | N/A | Blade Length - mm |  | mm | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-TDM-DIM |  |  | N/A - Blade Length - mm
Cut-TDM |  |  |  | 9: Force Cut - g; 10: 5; 12: N/A; 11: Force Cut - g; 14: g; 15: 5; 17: 5; 18: 5; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Cut-TDM-FORCES
Cut-TDM |  |  |  | 10: 10; 12: N/A; 14: g; 15: 10; 17: 10; 18: 10; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Cut-TDM-FORCES
Cut-TDM |  |  |  | 10: 20; 12: N/A; 14: g; 15: 20; 17: 20; 18: 20; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Cut-TDM-FORCES
Cut-TDM |  |  |  | 10: 50; 12: N/A; 14: g; 15: 50; 17: 50; 18: 50; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Cut-TDM-FORCES
Cut-TDM |  |  |  | 10: 100; 12: N/A; 14: g; 15: 100; 17: 100; 18: 100; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Cut-TDM-FORCES
Cut-TDM |  |  |  | 10: 200; 12: N/A; 14: g; 15: 200; 17: 200; 18: 200; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Cut-TDM-FORCES
Cut-TDM |  |  |  | 10: 500; 12: N/A; 14: g; 15: 500; 17: 500; 18: 500; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Cut-TDM-FORCES
Cut-TDM |  |  |  | 10: 1000; 12: N/A; 14: g; 15: 1000; 17: 1000; 18: 1000; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Cut-TDM-FORCES
Cut-TDM |  |  |  | 10: 2000; 12: N/A; 14: g; 15: 2000; 17: 2000; 18: 2000; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Cut-TDM-FORCES
Cut-TDM | [7 empty] | Length of Mounting area | Length | N/A | Length of Mounting area - mm |  | mm | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-TDM-DIM |  |  | Length - Length of Mounting area - mm
Cut-TDM |  |  |  | 10: Width; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-TDM-DIM
Cut-TDM | [7 empty] | Specimen holder curved radius | 38 | ±0.5 | Specimen holder curved radius - mm |  | mm | 38 | 0.5 | 37.5 | 38.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-TDM-DIM
Cut-TDM | [7 empty] | Multiple cut on the beneath of the Specimen holder | 10 | N/A | Multiple cut on the beneath of the Specimen holder - mm |  | mm | 10 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-TDM-DIM
Cut-TDM | [7 empty] | Blade with Exposed on the blade Holder | 12 | ±0.5 | Blade with Exposed on the blade Holder - mm |  | mm | 12 | 0.5 | 11.5 | 12.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-TDM-DIM
Cut-TDM | [7 empty] | Verification of the cutting motion system | 20 | ±2.5 | Verification of the cutting motion system - mm/sec |  | mm/sec | 20 | 2.5 | 17.5 | 22.5 | DIM/TIME |  | #N/A |  |  | #N/A | #N/A | Cut-TDM-DIM/TIME
Cut-TDM |  |  |  | 10: 40; 12: ±2.5; 14: mm/sec; 15: 40; 16: 2.5; 17: 37.5; 18: 42.5; 19: DIM/TIME; 21: #N/A |  | #N/A | #N/A | Cut-TDM-DIM/TIME
Cut-TDM |  |  |  | 10: 60; 12: ±2.5; 14: mm/sec; 15: 60; 16: 2.5; 17: 57.5; 18: 62.5; 19: DIM/TIME; 21: #N/A |  | #N/A | #N/A | Cut-TDM-DIM/TIME |  |  | 60
Cut-TDM |  | Tomodynano Meter | TAV-QTHKL |  |  |  |  | Mass (*) | 10 | N/A | Mass (*) - g |  | g | 10 | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Cut-TDM-MASS |  |  | 10 - Mass (*) - g
Cut-TDM |  |  |  | 10: 20; 12: N/A; 14: g; 15: 20; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Cut-TDM-MASS |  |  | 20
Cut-TDM |  |  |  | 10: 50; 12: N/A; 14: g; 15: 50; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Cut-TDM-MASS |  |  | 50
Cut-TDM |  |  |  | 10: 100; 12: N/A; 14: g; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Cut-TDM-MASS
Cut-TDM |  |  |  | 10: 200; 12: N/A; 14: g; 15: 200; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Cut-TDM-MASS
Cut-TDM |  |  |  | 10: 500; 12: N/A; 14: g; 15: 500; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Cut-TDM-MASS
Cut-TDM |  |  |  | 10: 1000; 12: N/A; 14: g; 15: 1000; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Cut-TDM-MASS
Cut-TDM |  |  |  | 10: 2000; 12: N/A; 14: g; 15: 2000; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Cut-TDM-MASS
Cut-other |  |  | TAV - QTHTD TAV - QTHKL | 5: Cutting machine - all kinds; 6: Cutting machine - all kinds; 7: Cutting machine - all kinds; 21: #N/A |  | #N/A | #N/A | Cut-other-
Cut-other | [7 empty] | Template 1 | …..................... | N/A | Template 1 - mm | Kích thước | mm |  | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-other-DIM
Cut-other |  |  |  | 10: ….....................; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-other-DIM
Cut-other | [7 empty] | Template 2 | …..................... | N/A | Template 2 - mm | Kích thước | mm |  | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-other-DIM |  |  | …..................... - Template 2 - mm
Cut-other |  |  |  | 10: ….....................; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-other-DIM |  |  | ….....................
Cut-other | [7 empty] | Template 3 | …..................... | N/A | Template 3 - mm | Kích thước | mm |  | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-other-DIM
Cut-other |  |  |  | 10: ….....................; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-other-DIM
Cut-other | [7 empty] | Template 4 | …..................... | N/A | Template 4 - mm | Kích thước | mm |  | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-other-DIM |  |  | …..................... - Template 4 - mm
Cut-other |  |  |  | 10: ….....................; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-other-DIM
Cut-other | [7 empty] | Template 5 | …..................... | N/A | Template 5 - mm | Kích thước | mm |  | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-other-DIM
Cut-other |  |  |  | 10: ….....................; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-other-DIM
Cut-other | [7 empty] | Template 6 | …..................... | N/A | Template 6 - mm | Kích thước | mm |  | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-other-DIM
Cut-other |  |  |  | 10: ….....................; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-other-DIM
Cut-other | [7 empty] | Template 7 | …..................... | N/A | Template 7 - mm | Kích thước | mm |  | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-other-DIM
Cut-other |  |  |  | 10: ….....................; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-other-DIM
Cut-other | [7 empty] | Template 8 | …..................... | N/A | Template 8 - mm | Kích thước | mm |  | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-other-DIM
Cut-other |  |  |  | 10: ….....................; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-other-DIM
Cut-other | [7 empty] | Template 9 | …..................... | N/A | Template 9 - mm | Kích thước | mm |  | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-other-DIM
Cut-other |  |  |  | 10: ….....................; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-other-DIM
Cut-other | [7 empty] | Template 10 | …..................... | N/A | Template 10 - mm | Kích thước | mm |  | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cut-other-DIM
Cut-other |  |  |  | 10: ….....................; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Cut-other-DIM
LabDye |  | Lab Dyeing Machine | TAV-QT-LABDYE TAV-QTHTD | Sandolab | Lab Dyeing Machine | Máy nhuộm thí nghiệm | Speed Temp Time | Speed (*) | 40 | ±2 | Speed (*) - rpm | Tốc độ | rpm | 40 | 2 | 38 | 42 | RPM |  | 0.1 | 2000000 | Service & Calibration | 0.1 | 0.1 | LabDye-RPM
LabDye | [7 empty] | Temp (*) | 60 | ±3 | Temp (*) - ͦC | Nhiệt độ | ͦC | 60 | 3 | 57 | 63 | DEG |  | 0.2 |  |  | 0.1 | 0.2 | LabDye-DEG
LabDye |  |  |  | 10: 100; 12: ±3; 14: ͦC; 15: 100; 16: 3; 17: 97; 18: 103; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | LabDye-DEG
LabDye |  |  |  | 10: 130; 12: ±3; 14: ͦC; 15: 130; 16: 3; 17: 127; 18: 133; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | LabDye-DEG
LabDye |  |  |  | 11: *; 14: G; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | LabDye-DEG
LabDye |  |  |  | 11: *; 14: G; 19: DEG; 21: 0.2; 22: 2000000 | Service & Calibration | 0.1 | 0.2 | LabDye-DEG
LabDye |  |  |  | 9: Speed Motor (*); 11: Speed Motor (*) - rpm; 13: Tốc độ Motor Băng tải; 14: rpm; 17: 0; 18: 0; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | LabDye-RPM
LabDye |  |  |  | 14: rpm; 17: 0; 18: 0; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | LabDye-RPM
LabDye |  |  |  | 14: rpm; 17: 0; 18: 0; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | LabDye-RPM
LabDye |  |  |  | 14: rpm; 17: 0; 18: 0; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | LabDye-RPM
LabDye |  |  |  | 14: rpm; 17: 0; 18: 0; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | LabDye-RPM
LabDye | [7 empty] | Time | 600 | N/A | Time - sec | Thời gian | sec | 600 | N/A | N/A | N/A | TIME |  | 0.02 |  |  | 0.001 | 0.02 | LabDye-TIME
LabDye |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | LabDye-
LabDye |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | LabDye-
LabDye |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | LabDye-
LabDye |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | LabDye-
Hydrostatic | AATCC 127,ISO 1420, ISO 811, JIS L1092 | Hydrostatic Head | TAV-QT.41-HHT | Hydrostatic Head, HydroPro | Hydrostatic Head | Thiết bị thử độ thẩm thấu nước của vật liệu | Pressure Pressure Gradient Test area | Pressure - mbar | 50 | ±1 | Pressure - mbar | Áp suất | mbar | 50 | 1 | 49 | 51 | Press |  | 0.1 | 2500000 | Service & Calibration | 0.05 | 0.1 | Hydrostatic-Press
Hydrostatic |  |  |  | 10: 100; 12: ±1; 14: mbar; 15: 100; 16: 1; 17: 99; 18: 101; 19: Press; 21: 0.1 |  | 0.05 | 0.1 | Hydrostatic-Press
Hydrostatic |  |  |  | 10: 200; 12: ±1; 14: mbar; 15: 200; 16: 1; 17: 199; 18: 201; 19: Press; 21: 0.1 |  | 0.05 | 0.1 | Hydrostatic-Press
Hydrostatic |  |  |  | 10: 500; 12: ±1; 14: mbar; 15: 500; 16: 1; 17: 499; 18: 501; 19: Press; 21: 0.1 |  | 0.05 | 0.1 | Hydrostatic-Press
Hydrostatic |  |  |  | 10: 1000; 12: ±1; 14: mbar; 15: 1000; 16: 1; 17: 999; 18: 1001; 19: Press; 21: 0.1 |  | 0.05 | 0.1 | Hydrostatic-Press
Hydrostatic |  |  |  | 10: 1500; 12: ±1; 14: mbar; 15: 1500; 16: 1; 17: 1499; 18: 1501; 19: Press; 21: 0.1 |  | 0.05 | 0.1 | Hydrostatic-Press
Hydrostatic |  |  |  | 10: 2000; 12: ±1; 14: mbar; 15: 2000; 16: 1; 17: 1999; 18: 2001; 19: Press; 21: 0.1 |  | 0.05 | 0.1 | Hydrostatic-Press
Hydrostatic |  |  |  |  | Hydrostatic Head | Thiết bị thử độ thẩm thấu nước của vật liệu |  | Pressure Gradient - mbar/min | 3 | ±0.5 | Pressure Gradient - mbar/min | Tốc độ tăng áp suất |  | 3 | 0.5 | 2.5 | 3.5 | Press |  | 0.1 |  |  | 0.05 | 0.1 | Hydrostatic-Press
Hydrostatic |  |  |  | 10: 5; 12: ±0.5; 14: mbar; 15: 5; 16: 0.5; 17: 4.5; 18: 5.5; 19: Press; 21: 0.1 |  | 0.05 | 0.1 | Hydrostatic-Press
Hydrostatic |  |  |  | 10: 10; 12: ±0.5; 14: mbar; 15: 10; 16: 0.5; 17: 9.5; 18: 10.5; 19: Press; 21: 0.1 |  | 0.05 | 0.1 | Hydrostatic-Press
Hydrostatic |  |  |  | 10: 60; 12: ±3; 14: mbar; 15: 60; 16: 3; 17: 57; 18: 63; 19: Press; 21: 0.1 |  | 0.05 | 0.1 | Hydrostatic-Press
Hydrostatic |  |  |  | 10: 200; 12: ±3; 14: mbar; 15: 200; 16: 3; 17: 197; 18: 203; 19: Press; 21: 0.1 |  | 0.05 | 0.1 | Hydrostatic-Press
Hydrostatic |  |  |  | 10: 500; 12: ±3; 14: mbar; 15: 500; 16: 3; 17: 497; 18: 503; 19: Press; 21: 0.1 |  | 0.05 | 0.1 | Hydrostatic-Press
Hydrostatic |  |  |  | 11: *; 19: Press; 21: 0.1 |  | 0.05 | 0.1 | Hydrostatic-Press
Hydrostatic |  |  |  | 9: Timer; 11: Timer - sec; 13: Thời gian; 14: sec; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Hydrostatic-TIME |  |  | - Timer - sec
Hydrostatic |  |  |  | 11: *; 14: ME; 19: TIME | [6 empty] | *
Hydrostatic |  |  |  | 11: *; 14: ME; 19: TIME | [6 empty] | *
Hydrostatic |  |  |  | 9: Test area; 11: Test area; 21: #N/A |  | #N/A | #N/A | Hydrostatic- |  |  | - Test area
Hydrostatic |  |  |  | 9: Diameter - mm (Area - ….... cm²) Đường kính; 12: N/A; 11: *; 13: Đường kính; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Hydrostatic-DIM |  |  | *
Hydrostatic |  |  |  | 9: Area Calculation; 12: N/A; 11: Area Calculation - cm²; 13: Diện tích; 14: cm²; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Hydrostatic-DIM
Hydrostatic |  |  |  | 9: Diameter - mm (Area - ….... cm²) Đường kính; 12: N/A; 11: *; 13: Đường kính; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Hydrostatic-DIM |  |  | *
Hydrostatic |  |  |  | 9: Area Calculation; 12: N/A; 11: Area Calculation - cm²; 13: Diện tích; 14: cm²; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Hydrostatic-DIM
Hydrostatic |  |  |  | 9: Diameter - mm (Area - ….... cm²) Đường kính; 12: N/A; 11: *; 13: Đường kính; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Hydrostatic-DIM |  |  | *
Hydrostatic |  |  |  | 9: Area Calculation; 12: N/A; 11: Area Calculation - cm²; 13: Diện tích; 14: cm²; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Hydrostatic-DIM
Pull-universal | ASTM F963, EN 71-1 | Universal Mechanical Safety | TAV-QTHKN | SafeGuard, Snap Tester | Universal Mechanical Safety | Thiết bị thử độ bứt nút |  | Force (*) - kgf | 5 | ±0.3 | Force (*) - kg | Lực | kg | 5 | 0.3 | 4.7 | 5.3 | FORCES |  | 0.14 |  |  | 0.0001 | 0.14 | Pull-universal-FORCES |  |  | 5 - Force (*) - kg
Pull-universal |  |  |  | 10: 10; 12: ±0.3; 14: kg; 15: 10; 16: 0.3; 17: 9.7; 18: 10.3; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Pull-universal-FORCES |  |  | 10
Pull-universal |  |  |  | 10: 15; 12: ±0.3; 14: kg; 15: 15; 16: 0.3; 17: 14.7; 18: 15.3; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Pull-universal-FORCE |  |  | 15
Pull-universal |  |  |  | 10: 20; 12: ±0.3; 14: kg; 15: 20; 16: 0.3; 17: 19.7; 18: 20.3; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Pull-universal-FORCE
Pull-universal |  |  |  | 10: 25; 12: ±0.3; 14: kg; 15: 25; 16: 0.3; 17: 24.7; 18: 25.3; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Pull-universal-FORCE |  |  | 25
Pull-universal |  |  |  | 10: 30; 12: ±0.3; 14: kg; 15: 30; 16: 0.3; 17: 29.7; 18: 30.3; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Pull-universal-FORCE |  |  | 30
Pull-force |  |  |  | 5: Pull force; 6: Pull Force Tester; 7: Thiết bị thử lực kéo cổ; 9: Force (*) …......... N/lb Lực; 10: 1 -; 11: *; 13: Lực; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Pull-force-FORCE |  |  | 1 - - *
Pull-force |  |  |  | 10: 2 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Pull-force-FORCE
Pull-force |  |  |  | 10: 3 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Pull-force-FORCE |  |  | 3 -
Pull-force |  |  |  | 10: 4 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Pull-force-FORCE |  |  | 4 -
Pull-force |  |  |  | 10: 5 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Pull-force-FORCE |  |  | 5 -
Pull-force |  |  |  | 10: 6 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Pull-force-FORCE
Pull-force |  |  |  | 10: 7 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Pull-force-FORCE
Pull-force |  |  |  | 10: 8 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Pull-force-FORCE
Pull-force |  |  |  | 10: 9 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Pull-force-FORCE
Pull-force |  |  |  | 10: 10 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Pull-force-FORCE
Compression-foam |  | IFD Machine |  | IFD Machine | IFD Machine | Thiết bị thử nén foam | Force | Force (*) …......... N/lb Lực | 1 - | ±1% | * | Lực | SS |  |  |  |  | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Compression-foam-MASS
Compression-foam |  |  |  | 10: 2 -; 12: ±1%; 14: SS; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Compression-foam-MASS
Compression-foam |  |  |  | 10: 3 -; 12: ±1%; 14: SS; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Compression-foam-MASS
Compression-foam |  |  |  | 10: 4 -; 12: ±1%; 14: SS; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Compression-foam-MASS
Compression-foam |  |  |  | 10: 5 -; 12: ±1%; 14: SS; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Compression-foam-MASS
Compression-foam |  |  |  | 10: 6 -; 12: ±1%; 14: SS; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Compression-foam-MASS
Compression-foam |  |  |  | 10: 7 -; 12: ±1%; 14: SS; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Compression-foam-MASS
Compression-foam |  |  |  | 10: 8 -; 12: ±1%; 14: SS; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Compression-foam-MASS
Compression-foam |  |  |  | 10: 9 -; 12: ±1%; 14: SS; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Compression-foam-MASS
Compression-foam |  |  |  | 10: 10 -; 12: ±1%; 14: SS; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Compression-foam-MASS
Compression-foam | [7 empty] | Ruler | 0-100 | N/A | Ruler - mm | Thước | mm | 100 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Compression-foam-DIM
Compression-foam |  |  |  | 10: 100-200; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Compression-foam-DIM
Compression-foam |  |  |  | 10: 200-300; 12: N/A; 14: mm; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Compression-foam-DIM
Tension-yarn |  |  |  | Tension Meter | Tension Meter |  |  | Down Ward Force - N |  | N/A | Down Ward Force - N | Lực hướng xuống | N |  |  | 0 | 0 | FORCES |  | 0.14 |  |  | 0.0001 | 0.14 | Tension-yarn-FORCES
Tension-yarn |  |  |  | 9: Force Tension - N/cm; 12: ±5%; 11: Force Tension - N/cm; 13: Lực căng; 14: N; 17: 0; 18: 0; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Tension-yarn-FORCES
Tension-yarn |  |  |  | 12: ±5%; 14: N; 17: 0; 18: 0; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Tension-yarn-FORCES
Tension-yarn |  |  |  | 12: ±5%; 14: N; 17: 0; 18: 0; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Tension-yarn-FORCES
Tension-yarn |  |  |  | 12: ±5%; 14: N; 17: 0; 18: 0; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Tension-yarn-FORCES
Tension-yarn |  |  |  | 9: Point calibration; 10: Check; 11: Point calibration; 17: 0; 18: 0; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Tension-yarn-FORCES
Tension-yarn |  |  |  | 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Tension-yarn-
Tension-yarn |  |  |  | 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Tension-yarn-
Tension-yarn |  |  |  | 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Tension-yarn-
Tension-yarn |  |  |  | 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Tension-yarn-
Tension-yarn |  |  |  | 11: *; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Tension-yarn-
View-pilling | EN ISO 12945-1, EN ISO 12945-2, ASTM D 3512 | Pilling Assessment Viewer | TAV-QT-VIEW | Pilling Assessment Viewer | Pilling Assessment Viewer | Hộp đánh giá mẫu | Lightsource Temp Lightsource illuminance Tilt angle | Lightsource Temp (*) | D65 |  | Lightsource Temp (*) - K | Nhiệt độ màu | K |  |  |  |  | KEL |  | 40 | 800000 | Service & Verification | 1 | 40 | View-pilling-KEL
View-pilling |  |  |  | 10: CWF; 14: K; 19: KEL; 21: 40 |  | 1 | 40 | View-pilling-KEL
View-pilling |  |  |  | 9: Lightsource illuminance (*); 10: D65; 11: Lightsource illuminance (*) - Lux; 13: Độ rọi; 14: Lux; 19: LUX; 21: 0.012 |  | 1 | 0.012 | View-pilling-LUX |  |  | D65 - Lightsource illuminance (*) - Lux
View-pilling |  |  |  | 10: CWF; 14: Lux; 19: LUX; 21: 0.012 |  | 1 | 0.012 | View-pilling-LUX |  |  | CWF
View-pilling |  |  |  | 9: Tilt angle; 11: Tilt angle - ⁰; 13: Góc nghiêng; 14: ⁰; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | View-pilling-ANG |  |  | - Tilt angle - ⁰
View-pilling |  |  |  | 9: Light Angle; 10: 45; 11: Light Angle - ⁰; 13: Góc rọi; 14: ⁰; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | View-pilling-ANG |  |  | 45 - Light Angle - ⁰
View-board | AATCC Viewing Board | Viewing Board | TAV-QT-VIEW | 5: Viewing Board; 6: AATCC Viewing Board; 7: Bảng đánh giá mẫu AATCC; 8: Lightsource Temp Lightsource illuminance Tilt angle; 9: Lightsource Temp (*); 10: D65; 11: Lightsource Temp (*) - K; 13: Nhiệt độ màu; 21: #N/A; 22: 800000 | Service & Verification | #N/A | #N/A | View-board- |  |  | D65 - Lightsource Temp (*) - K
View-board |  |  |  | 10: Left; 14: K; 17: 0; 18: 0; 19: KEL; 21: 40 |  | 1 | 40 | View-board-KEL |  |  | Left
View-board |  |  |  | 10: Center; 14: K; 17: 0; 18: 0; 19: KEL; 21: 40 |  | 1 | 40 | View-board-KEL |  |  | Center
View-board |  |  |  | 10: Right; 14: K; 17: 0; 18: 0; 19: KEL; 21: 40 |  | 1 | 40 | View-board-KEL |  |  | Right
View-board |  |  |  | 10: CWF; 21: #N/A |  | #N/A | #N/A | View-board-
View-board |  |  |  | 10: Left; 14: K; 17: 0; 18: 0; 19: KEL; 21: 40 |  | 1 | 40 | View-board-KEL
View-board |  |  |  | 10: Center; 14: K; 17: 0; 18: 0; 19: KEL; 21: 40 |  | 1 | 40 | View-board-KEL
View-board |  |  |  | 10: Right; 14: K; 17: 0; 18: 0; 19: KEL; 21: 40 |  | 1 | 40 | View-board-KEL
View-board |  |  |  |  | Pilling Assessment Viewer | Hộp đánh giá mẫu |  | Lightsource illuminance (*) | D65 | N/A | Lightsource illuminance (*) - Lux | Độ rọi | Lux |  | N/A | #VALUE! | #VALUE! | LUX |  | 0.012 |  |  | 1 | 0.012 | View-board-LUX
View-board |  |  |  | 10: CWF; 12: N/A; 14: Lux; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: LUX; 21: 0.012 |  | 1 | 0.012 | View-board-LUX
View-board |  |  |  |  | Pilling Assessment Viewer | Hộp đánh giá mẫu |  | Tilt angle | 5 | ±0.1 | Tilt angle - ⁰ | Góc nghiêng | ⁰ | 5 | 0.1 | 4.9 | 5.1 | ANG |  | 0.03 |  |  | 0.05 | 0.03 | View-board-ANG
View-board |  |  |  | 10: 5; 12: ±0.1; 14: ⁰; 15: 5; 16: 0.1; 17: 4.9; 18: 5.1; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | View-board-ANG
View-board |  |  |  | 10: 5; 12: ±0.1; 14: ⁰; 15: 5; 16: 0.1; 17: 4.9; 18: 5.1; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | View-board-ANG |  |  | 5
View-aatcc |  | Viewing Board | TAV-QT-VIEW | Viewing Board | Viewing Board | Hộp đánh giá mẫu | Lightsource Temp Lightsource illuminance Tilt angle | Lightsource Temp | D65 | N/A | Lightsource Temp (*) - K | Nhiệt độ màu | K |  |  |  |  | KEL |  | 40 |  |  | 1 | 40 | View-aatcc-KEL |  |  | D65 - Lightsource Temp (*) - K
View-aatcc |  |  |  | 10: CWF; 12: N/A; 14: K; 19: KEL; 21: 40 |  | 1 | 40 | View-aatcc-KEL |  |  | CWF
View-aatcc |  |  |  | 10: 500W; 12: N/A; 14: K; 19: KEL; 21: 40 |  | 1 | 40 | View-aatcc-KEL |  |  | 500W
View-aatcc |  |  |  | 6: Pilling Assessment Viewer; 7: Hộp đánh giá mẫu; 9: Lightsource illuminance; 10: D65; 12: N/A; 11: Lightsource illuminance - Lux; 13: Độ rọi; 14: Lux; 19: LUX; 21: 0.012 |  | 1 | 0.012 | View-aatcc-LUX |  |  | D65 - Lightsource illuminance - Lux
View-aatcc |  |  |  | 10: CWF; 12: N/A; 14: Lux; 19: LUX; 21: 0.012 |  | 1 | 0.012 | View-aatcc-LUX |  |  | CWF
View-aatcc |  |  |  | 10: 500W; 12: N/A; 14: Lux; 19: LUX; 21: 0.012 |  | 1 | 0.012 | View-aatcc-LUX |  |  | 500W
View-aatcc |  |  |  |  | Pilling Assessment Viewer | Hộp đánh giá mẫu |  | Angle Board - ⁰ | 5 | ±0.1 | Angle Board - ⁰ | Góc nghiêng | ⁰ | 5 | 0.1 | 4.9 | 5.1 | ANG |  | 0.03 |  |  | 0.05 | 0.03 | View-aatcc-ANG |  |  | 5 - Angle Board - ⁰
View-aatcc | [27 empty] | 0
View-aatcc | [7 empty] | Floor distance to lamp - Inch | 93 | ±1 | Floor distance to lamp - Inch | Khoảng cách từ sàn tới đèn | Inch | 93 | 1 | 92 | 94 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | View-aatcc-DIM |  |  | 93 - Floor distance to lamp - Inch
View-aatcc | [27 empty] | 0
View-aatcc | [7 empty] | Table size - Inch | 72 | ±1 | Table size - Inch | Kích thước bảng | Inch | 72 | 1 | 71 | 73 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | View-aatcc-DIM |  |  | 72 - Table size - Inch
View-aatcc | [27 empty] | 0
View-aatcc | [7 empty] | Height Footer to board - Inch | 25 | ±0.5 | Height Footer to board - Inch | Chiếu cao từ chân đến bảng | Inch | 25 | 0.5 | 24.5 | 25.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | View-aatcc-DIM |  |  | 25 - Height Footer to board - Inch
View-aatcc | [27 empty] | 0
View-aatcc | [7 empty] | Distance lamp to board - Inch | 24.4 | ±0.3 | Distance lamp to board - Inch | Khoảng cách từ đèn đến bảng | Inch | 24.4 | 0.3 | 24.1 | 24.7 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | View-aatcc-DIM
View-aatcc
View-aatcc |  |  |  | 9: Vị trí người đứng đén bảng - Inch; 10: 48; 12: ±1; 14: Inch; 15: 48; 16: 1; 17: 47; 18: 49; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | View-aatcc-DIM
View-aatcc | bar: 12
View-aatcc |  |  |  | 9: 500W Reflector Flood lamp - Inch; 10: 60; 12: ±1; 14: Inch; 15: 60; 16: 1; 17: 59; 18: 61; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | View-aatcc-DIM
View-aatcc
View-aatcc |  |  |  | 9: Góc đèn 500W và người - ⁰; 10: 45; 12: ±1; 14: ⁰; 15: 45; 16: 1; 17: 44; 18: 46; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | View-aatcc-ANG |  |  | 45
View-aatcc | [27 empty] | 48
View-aatcc |  |  |  | 9: Table distance to the operator - Inch; 10: 48; 12: N/A; 11: Table distance to the operator - Inch; 13: Khoảng cách từ bảng đến người điều khiển; 14: Inch; 15: 48; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | View-aatcc-DIM
View-aatcc
View-aatcc |  |  |  | 9: Table distance to the lamp 500W - Inch; 10: 72; 12: N/A; 11: Table distance to the lamp 500W - Inch; 13: Khoảng cách từ bảng đến đèn 500W; 14: Inch; 15: 72; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | View-aatcc-DIM
View-aatcc
View-aatcc |  |  |  | 9: Chiều dài góc nghiêng - Inch; 10: 6.4; 12: N/A; 14: Inch; 15: 6.4; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | View-aatcc-DIM
Sharp-point | BS EN 71, ASTM F963 | Sharp Point | TAV-QT-SHARP | Sharp Point | Sharp Point | Thiết bị thử độ nhọn | Dimension Depth of Slot Movement per 5 Divisions Test force | Dimension | BS EN 71 - 1.02x1.15mm | ±0.02 | Dimension | Kích thước | mm |  |  |  |  | DIM |  | 0.02 | 800000 | Service & Verification | 0.01 | 0.02 | Sharp-point-DIM
Sharp-point |  |  |  | 8: Dimension Depth of Slot Movement per 5 Divisions Test force; 12: ±0.02; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Sharp-point-DIM
Sharp-point |  |  |  | 10: ASTM F963 - 0.040x0.045"; 12: N/A; 14: inch; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Sharp-point-DIM
Sharp-point |  |  |  | 12: N/A; 14: inch; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Sharp-point-DIM
Sharp-point |  |  |  | 9: Depth of Slot; 10: BS EN 71 - 0.38mm; 12: ±0.02; 11: Depth of Slot; 13: Chiều sâu; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Sharp-point-DIM
Sharp-point |  |  |  | 10: ASTM F963 - 0.015"; 12: N/A; 14: inch; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Sharp-point-DIM
Sharp-point |  |  |  | 9: Movement per 5 Divisions; 10: 0.12mm; 12: ±0.02; 11: Movement per 5 Divisions; 13: Movement per 5 Divisions; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Sharp-point-DIM
Sharp-point |  |  |  | 10: 0.005"; 12: N/A; 14: inch; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Sharp-point-DIM
Sharp-point |  |  |  | 9: Test force; 10: BS EN 71 - 4.5N; 12: N/A; 11: Test force; 13: Lực; 14: N; 15: 4.5; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Sharp-point-FORCES
Sharp-point |  |  |  | 10: ASTM F963 - 1lbs (4.45N); 12: N/A; 14: lb; 15: 1; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Sharp-point-FORCES
Sharp-edge | BS EN 71, ASTM F963 | Sharp Edge | TAV-QT-SHARP | Sharp Edge | Sharp Edge | Thiết bị thử cạnh sắc | Force Diameter of mandrel Rotate speed | Force | BS EN 71 - 6N | ±0.5 | Force | Lực | N | 6 | 0.5 | 3 | 9 | FORCES |  | 0.1 | 800000 | Service & Verification | 0.01 | 0.1 | Sharp-edge-FORCES
Sharp-edge |  |  |  | 10: ASTM F963 - 1.35 lbs; 12: N/A; 14: lb; 15: 1.35; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Sharp-edge-FORCES
Sharp-edge | [7 empty] | Diameter of mandrel | BS EN 71 - 9.53mm | ±0.12 | Diameter of mandrel | Đường kính của trục | mm | 9.53 | 0.12 | 9.41 | 9.65 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Sharp-edge-DIM
Sharp-edge |  |  |  | 10: ASTM F963 - 0.375"; 12: ±0.005; 14: inch; 15: 0.375; 16: 0.005; 17: 0.37; 18: 0.38; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Sharp-edge-DIM
Sharp-edge | [7 empty] | Rotate speed | BS EN 71 - 23mm/s | ±4 | Rotate speed | Tốc độ xoay | mm/sec | 23 | 4 | 19 | 27 | DIM/TIME |  | 0.02 |  |  | 0.01 | 0.02 | Sharp-edge-DIM/TIME
Sharp-edge |  |  |  | 10: ASTM F963 - 1 inch/s; 12: ±0.08; 14: inch/sec; 15: 1; 16: 0.08; 17: 0.92; 18: 1.08; 19: DIM/TIME; 21: 0.02 |  | 0.01 | 0.02 | Sharp-edge-DIM/TIME
Volume |  | 250 ml | TAV-QT-VOL TAV-QTHKL | 5: Cylinder/Burette/Pippet/Flask; 6: 'Cylinder/Burette/Pippet/Flask; 7: Dụng cụ thí nghiệm; 8: Volume; 9: Cylinder/Burette/Pippet/Flask; 10: …....... (ml/µl); 11: *; 22: 400000 | Service & Verification | #N/A | #N/A | Volume-
Volume | [7 empty] | Volume - (ml/µl) Thể tích | 1 - | N/A | * | Thể tích | ml |  | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Volume-MASS
Volume |  |  |  | 10: 2 -; 12: N/A; 14: ml; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Volume-MASS
Volume |  |  |  | 10: 3 -; 12: N/A; 14: ml; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Volume-MASS |  |  | 3 -
Volume |  |  |  | 10: 4 -; 12: N/A; 14: ml; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Volume-MASS |  |  | 4 -
Volume |  |  |  | 10: 5 -; 12: N/A; 14: ml; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Volume-MASS |  |  | 5 -
Volume |  |  |  | 10: 6 -; 12: N/A; 14: ml; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Volume-MASS
Volume |  |  |  | 10: 7 -; 12: N/A; 14: ml; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Volume-MASS
Volume |  |  |  | 10: 8 -; 12: N/A; 14: ml; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Volume-MASS
Volume |  |  |  | 10: 9 -; 12: N/A; 14: ml; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Volume-MASS
Volume |  |  |  | 10: 10 -; 12: N/A; 14: ml; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Volume-MASS
Weight-set |  | Set of Standard Weight | TAV-QTHKL | Set of Standard Weight | Set of Standard Weight | Bộ quả cân chuẩn | Weight | Mass (*) - g | 0.001 | N/A | Mass (*) - g | Khối lượng | g | 0.001 | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 | 500000 | Service & Verification | 0.01 | 0.1 | Weight-set-MASS
Weight-set |  |  |  | 10: 0.002; 12: N/A; 14: SS; 15: 0.002; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS |  |  | 0.002
Weight-set |  |  |  | 10: 0.005; 12: N/A; 14: SS; 15: 0.005; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS |  |  | 0.005
Weight-set |  |  |  | 10: 0.01; 12: N/A; 14: SS; 15: 0.01; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS
Weight-set |  |  |  | 10: 0.02; 12: N/A; 14: SS; 15: 0.02; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS
Weight-set |  |  |  | 10: 0.05; 12: N/A; 14: SS; 15: 0.05; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS
Weight-set |  |  |  | 10: 0.1; 12: N/A; 14: SS; 15: 0.1; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS
Weight-set |  |  |  | 10: 0.2; 12: N/A; 14: SS; 15: 0.2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS
Weight-set |  |  |  | 10: 0.5; 12: N/A; 14: SS; 15: 0.5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS
Weight-set |  |  |  | 10: 1; 12: N/A; 14: SS; 15: 1; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS
Weight-set |  |  |  | 10: 2; 12: N/A; 14: SS; 15: 2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS
Weight-set |  |  |  | 10: 5; 12: N/A; 14: SS; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS
Weight-set |  |  |  | 10: 10; 12: N/A; 14: SS; 15: 10; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS
Weight-set |  |  |  | 10: 20; 12: N/A; 14: SS; 15: 20; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS
Weight-set |  |  |  | 10: 50; 12: N/A; 14: SS; 15: 50; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS
Weight-set |  |  |  | 10: 100; 12: N/A; 14: SS; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS
Weight-set |  |  |  | 10: 200; 12: N/A; 14: SS; 15: 200; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS
Weight-set |  |  |  | 10: 500; 12: N/A; 14: SS; 15: 500; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-set-MASS
Weight-set |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Weight-set-
Weight-set |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Weight-set-
Weight |  | Weight | TAV-QTHKL | 5: 1kg, 2kg, 5kg,10kg; 6: Standard Weight; 7: Tạ chuẩn; 8: Weight; 9: Weight - (g/kg/….....) Khối lượng; 10: ….... (g/kg/….....); 12: N/A; 11: *; 21: #N/A |  | #N/A | #N/A | Weight-
Weight |  |  |  | 12: N/A; 14: SS; 17: 0; 18: 0; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-MASS
Weight |  |  |  | 12: N/A; 14: SS; 17: 0; 18: 0; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-MASS |  |  | Point 1 - 37°C | Point 1 - 37°C -
Weight |  |  |  | 12: N/A; 14: SS; 17: 0; 18: 0; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-MASS |  |  | Point 2 - 37°C | Point 2 - 37°C -
Weight |  |  |  | 12: N/A; 14: SS; 17: 0; 18: 0; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Weight-MASS |  |  | Point 3 - 37°C | Point 3 - 37°C -
Balance |  | Balance | TAV-QTHCA | 5: Balance; 6: Balance; 7: Cân điện tử; 8: Display Value; 9: Load (*) - …............ g; 11: Load (*) - g; 13: Mức cân; 22: 800000 | Service & Verification | #N/A | #N/A | Balance- |  |  | Point 4 - 37°C | Point 4 - 37°C -
Balance |  |  |  | 10: 0.001; 12: N/A; 14: g; 15: 0.001; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.06; 21: 0 |  | 0.06 | 1.6 | Balance-MASS |  |  | Point 5 - 37°C | Point 5 - 37°C -
Balance |  |  |  | 10: 0.002; 12: N/A; 14: g; 15: 0.002; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.06; 21: 6e-06 |  | 0.06 | 1.6 | Balance-MASS |  |  | Point 6 - 37°C | Point 6 - 37°C -
Balance |  |  |  | 10: 0.005; 12: N/A; 14: g; 15: 0.005; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.06; 21: 6e-06 |  | 0.06 | 1.6 | Balance-MASS
Balance |  |  |  | 10: 0.01; 12: N/A; 14: g; 15: 0.01; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.06; 21: 8e-06 |  | 0.06 | 1.6 | Balance-MASS |  |  | Point 7 - 37°C | Point 7 - 37°C -
Balance |  |  |  | 10: 0.02; 12: N/A; 14: g; 15: 0.02; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.06; 21: 1e-05 |  | 0.06 | 1.6 | Balance-MASS |  |  | Point 8 - 37°C | Point 8 - 37°C -
Balance |  |  |  | 10: 0.05; 12: N/A; 14: g; 15: 0.05; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.06; 21: 1.2e-05 |  | 0.06 | 1.6 | Balance-MASS |  |  | Point 9 - 37°C | Point 9 - 37°C -
Balance |  |  |  | 10: 0.1; 12: N/A; 14: g; 15: 0.1; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.06; 21: 1.6e-05 |  | 0.06 | 1.6 | Balance-MASS
Balance |  |  |  | 10: 0.2; 12: N/A; 14: g; 15: 0.2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.06; 21: 2e-05 |  | 0.06 | 1.6 | Balance-MASS
Balance |  |  |  | 10: 0.5; 12: N/A; 14: g; 15: 0.5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.06; 21: 2.5e-05 |  | 0.06 | 1.6 | Balance-MASS
Balance | [8 empty] | 1 | N/A |  |  | g | 1 | N/A | #VALUE! | #VALUE! | MASS | 0.06 | 3e-05 | 800000 | Service & Verification | 0.06 | 1.6 | Balance-MASS
Balance |  |  |  | 10: 2; 12: N/A; 14: g; 15: 2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.06; 21: 4e-05 |  | 0.06 | 1.6 | Balance-MASS
Balance |  |  |  | 10: 5; 12: N/A; 14: g; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.06; 21: 5e-05 |  | 0.06 | 1.6 | Balance-MASS
Balance |  |  |  | 10: 10; 12: N/A; 14: g; 15: 10; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0; 21: 6e-05 |  | 0 | 18 | Balance-MASS
Balance |  |  |  | 10: 20; 12: N/A; 14: g; 15: 20; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0; 21: 8e-05 |  | 0 | 18 | Balance-MASS
Balance |  |  |  | 10: 50; 12: N/A; 14: g; 15: 50; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0; 21: 0.0001 |  | 0 | 18 | Balance-MASS
Balance |  |  |  | 10: 100; 12: N/A; 14: g; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0; 21: 0.00016 |  | 0 | 0.012 | Balance-MASS
Balance |  |  |  | 10: 200; 12: N/A; 14: g; 15: 200; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0; 21: 0.0003 |  | 0 | 0.012 | Balance-MASS |  |  | Point 1 - 37°C | Point 1 - 37°C -
Balance |  |  |  | 10: 500; 12: N/A; 14: g; 15: 500; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0; 21: 0.0008 |  | 0 | 0.012 | Balance-MASS |  |  | Point 2 - 37°C | Point 2 - 37°C -
Balance |  |  |  | 10: 1000; 12: N/A; 14: g; 15: 1000; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0; 21: 0.0016 |  | 0 | 0.012 | Balance-MASS |  |  | Point 3 - 37°C | Point 3 - 37°C -
Balance |  |  |  | 10: 2000; 12: N/A; 14: g; 15: 2000; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0; 21: 0.08 |  | 0 | 0.012 | Balance-MASS |  |  | Point 4 - 37°C | Point 4 - 37°C -
Balance |  |  |  | 10: 3000; 12: N/A; 14: g; 15: 3000; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.08 |  | 0 | 0.012 | Balance-MASS |  |  | Point 5 - 37°C | Point 5 - 37°C -
Balance |  |  |  | 10: 4000; 12: N/A; 14: g; 15: 4000; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.08 |  | 0 | 0.012 | Balance-MASS |  |  | Point 6 - 37°C | Point 6 - 37°C -
Balance |  |  |  | 10: 5000; 12: N/A; 14: g; 15: 5000; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0; 21: 0.08 |  | 0 | 0.012 | Balance-MASS |  |  | Point 7 - 37°C | Point 7 - 37°C -
Balance |  |  |  | 10: 10000; 12: N/A; 14: g; 15: 10000; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.06; 21: 0.16 |  | 0.06 | 1.6 | Balance-MASS |  |  | Point 8 - 37°C | Point 8 - 37°C -
Balance |  |  |  | 10: 15000; 12: N/A; 14: g; 15: 15000; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.06; 21: 0.3 |  | 0.06 | 1.6 | Balance-MASS |  |  | Point 9 - 37°C | Point 9 - 37°C -
Balance |  |  |  | 10: 20000; 12: N/A; 14: g; 15: 20000; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.06; 21: 0.3 |  | 0.06 | 1.6 | Balance-MASS
Balance |  |  |  | 10: 25000; 12: N/A; 14: g; 15: 25000; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.06; 21: 23 |  | 0.06 | 1.6 | Balance-MASS
Balance |  |  |  | 10: 30000; 12: N/A; 14: g; 15: 30000; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.06; 21: 23 |  | 0.06 | 1.6 | Balance-MASS
Balance |  |  |  | 10: 40000; 12: N/A; 14: g; 15: 40000; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 23 |  |  |  | Balance-MASS
Balance |  |  |  | 10: 50000; 12: N/A; 14: g; 15: 50000; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 23 |  |  |  | Balance-MASS
Balance |  |  |  | 9: Other - …........ g; 12: N/A; 11: *; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS |  |  |  | Balance-MASS
Balance |  |  |  | 12: N/A; 11: *; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS |  |  |  | Balance-MASS
Balance |  |  |  | 12: N/A; 11: *; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS |  |  |  | Balance-MASS
Thickness |  | Thickness Gauge | TAV-QT.06-DIM | Thickness Gauge | Thickness Gauge | Thiết bị đo độ dày | Distance | Thickness 1 - …........ mm | …........ mm |  | Thickness 1 - mm | Độ dày 1 | [6 empty] | #N/A | #N/A | 800000 | Service & Verification | #N/A | #N/A | Thickness -
Thickness |  |  |  | 10: 1 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM
Thickness |  |  |  | 10: 2 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM
Thickness |  |  |  | 10: 3 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM
Thickness |  |  |  | 10: 4 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM |  |  | Point 1 - 37°C | Point 1 - 37°C -
Thickness |  |  |  | 10: 5 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM |  |  | Point 2 - 37°C | Point 2 - 37°C -
Thickness |  |  |  | 10: 6 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM |  |  | Point 3 - 37°C | Point 3 - 37°C -
Thickness |  |  |  | 10: 7 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM |  |  | Point 4 - 37°C | Point 4 - 37°C -
Thickness |  |  |  | 10: 8 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM |  |  | Point 5 - 37°C | Point 5 - 37°C -
Thickness |  |  |  | 10: 9 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM |  |  | Point 6 - 37°C | Point 6 - 37°C -
Thickness |  |  |  | 10: 10 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM |  |  | Point 7 - 37°C | Point 7 - 37°C -
Thickness |  |  |  | 9: Thickness 2 - …........ mm; 10: …........ mm; 11: Thickness 2 - mm; 13: Độ dày 1; 20: #N/A; 21: #N/A |  | #N/A | #N/A | Thickness - |  |  | Point 8 - 37°C | Point 8 - 37°C -
Thickness |  |  |  | 10: 1 -; 12: N/A; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM |  |  | Point 9 - 37°C | Point 9 - 37°C -
Thickness |  |  |  | 10: 2 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM
Thickness |  |  |  | 10: 3 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM
Thickness |  |  |  | 10: 4 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM
Thickness |  |  |  | 10: 5 -; 12: N/A; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM
Thickness |  |  |  | 10: 6 -; 12: N/A; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM
Thickness |  |  |  | 10: 7 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM
Thickness |  |  |  | 10: 8 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM
Thickness |  |  |  | 10: 9 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM
Thickness |  |  |  | 10: 10 -; 12: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Thickness -DIM
Thickness1 |  | Thickness Gauge |  | Thickness Gauge | Thickness Gauge | Thiết bị đo độ dày | Accuracy of indication Diameter of presser foot Diameter of anvil | Accuracy of indication | 1 | N/A | Accuracy of indication - mm | 1 | mm | 1 | N/A | #VALUE! | #VALUE! | DIM | 0.01 | 0.02 | 800000 | Service & Verification | 0.01 | 0.02 | Thickness1-DIM
Thickness1 | [8 empty] | 1.5 | N/A |  | 1.5 | mm | 1.5 | N/A | #VALUE! | #VALUE! | DIM | 0.01 | 0.02 |  |  | 0.01 | 0.02 | Thickness1-DIM
Thickness1 | [8 empty] | 2 | N/A |  | 2 | mm | 2 | N/A | #VALUE! | #VALUE! | DIM | 0.01 | 0.02 |  |  | 0.01 | 0.02 | Thickness1-DIM
Thickness1 | [8 empty] | 3 | N/A |  | 3 | mm | 3 | N/A | #VALUE! | #VALUE! | DIM | 0.01 | 0.02 |  |  | 0.01 | 0.02 | Thickness1-DIM
Thickness1 | [8 empty] | 5 | N/A |  | 5 | mm | 5 | N/A | #VALUE! | #VALUE! | DIM | 0.01 | 0.02 |  |  | 0.01 | 0.02 | Thickness1-DIM
Thickness1 | [8 empty] | 10 | N/A |  | 10 | mm | 10 | N/A | #VALUE! | #VALUE! | DIM | 0.01 | 0.02 |  |  | 0.01 | 0.02 | Thickness1-DIM
Thickness1 | [7 empty] | Diameter of presser foot | 10 | N/A | Diameter of presser foot - mm | 10 | mm | 10 | N/A | #VALUE! | #VALUE! | DIM | 0.01 | 0.02 |  |  | 0.01 | 0.02 | Thickness1-DIM
Thickness1 | [7 empty] | Diameter of anvil | 10 | N/A | Diameter of anvil - mm | 10 | mm | 10 | N/A | #VALUE! | #VALUE! | DIM | 0.01 | 0.02 |  |  | 0.01 | 0.02 | Thickness1-DIM
Thickness2 | SATRA TM1 | Thickness Gauge |  | Thickness Gauge | Thickness Gauge | Thiết bị đo độ dày | Accuracy of dial gauge Diameter of presser foot Diameter of anvil Downward pressure exerted on test specimen | Accuracy of dial gauge | 0.5-10 | N/A | Accuracy of dial gauge - mm | 0.5-10 | mm | 0.5-10 | N/A | #VALUE! | #VALUE! | DIM | 0.01 | 0.02 | 800000 | Service & Verification | 0.01 | 0.02 | Thickness2-DIM
Thickness2 | [7 empty] | Diameter of presser foot | 10 | ±1 | Diameter of presser foot - mm | 10 | mm | 10 | 1 | 9 | 11 | DIM | 0.01 | 0.02 |  |  | 0.01 | 0.02 | Thickness2-DIM
Thickness2 | [7 empty] | Diameter of anvil | >= diameter of foot | >= diameter of foot | Diameter of anvil - mm | >= diameter of foot | mm | >= diameter of foot | >= diameter of foot | #VALUE! | #VALUE! | DIM | 0.01 | 0.02 |  |  | 0.01 | 0.02 | Thickness2-DIM
Thickness2 | [7 empty] | Downward pressure exerted on test specimen | 49 | ±5 | Downward pressure exerted on test specimen - kpa | 49 | kpa | 49 | 5 | 44 | 54 | Press | #N/A | #N/A |  |  | #N/A | #N/A | Thickness2-Press
Thickness3 |  | Thickness Gauge |  | Thickness Gauge | Thickness Gauge | Thiết bị đo độ dày | Accuracy of indication Ring Check Rubber Standard | Accuracy of indication | 1 | N/A | Accuracy of indication - mm |  | mm | 1 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Thickness3-DIM
Thickness3 |  |  |  | 10: 2; 12: N/A; 14: mm; 15: 2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Thickness3-DIM
Thickness3 |  |  |  | 10: 3; 12: N/A; 14: mm; 15: 3; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Thickness3-DIM
Thickness3 |  |  |  | 10: 4; 12: N/A; 14: mm; 15: 4; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Thickness3-DIM
Thickness3 |  |  |  | 10: 5; 12: N/A; 14: mm; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Thickness3-DIM
Thickness3 |  |  |  | 10: 6; 12: N/A; 14: mm; 15: 6; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Thickness3-DIM
Thickness3 |  |  |  | 10: 7; 12: N/A; 14: mm; 15: 7; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Thickness3-DIM
Thickness3 |  |  |  | 10: 8; 12: N/A; 14: mm; 15: 8; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Thickness3-DIM
Thickness3 |  |  |  | 10: 9; 12: N/A; 14: mm; 15: 9; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Thickness3-DIM
Thickness3 |  |  |  | 10: 10; 12: N/A; 14: mm; 15: 10; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Thickness3-DIM
Thickness3 | [7 empty] | Check Rubber Standard | 1.6 | N/A | Check Rubber Standard - mm |  | mm | 1.6 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Thickness3-DIM
Thickness3 |  |  |  | 10: 1.9; 12: N/A; 14: mm; 15: 1.9; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Thickness3-DIM
Thickness3 | [7 empty] | Ring | Ø20 | N/A | Ring - mm |  | mm | 20 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Thickness3-DIM
Thickness3 |  |  |  | 10: Ø25; 12: N/A; 14: mm; 15: 25; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Thickness3-DIM
Caliper |  | Caliper | TAV-QT.06-DIM | 5: Caliper; 6: Thickness Gauge; 7: Thước kẹp điện tử; 8: Distance; 9: Dimension 1 - …........ mm Kích thước 1; 11: *; 20: #N/A; 21: #N/A; 22: 1200000 | Service & Verification | #N/A | #N/A | Caliper-
Caliper |  |  |  | 10: 1 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 2 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 3 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 4 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 5 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 6 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 7 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 8 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 9 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 10 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 9: Dimension 2 - …........ mm Kích thước 2; 11: *; 20: #N/A; 21: #N/A |  | #N/A | #N/A | Caliper-
Caliper |  |  |  | 10: 1 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 2 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 3 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 4 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 5 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 6 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 7 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 8 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 9 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Caliper |  |  |  | 10: 10 -; 12: N/A; 14: mm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.14 |  | 0.01 | 0.14 | Caliper-DIM
Pick |  |  | TAV-QTH-DIM | Pick Glass | Pick Glass | Kính đếm số lượng sợi |  | Dimension of Devided Fractions - Ruler 1 | 0-5 | N/A | Dimension of Devided Fractions - Ruler 1 - mm | Thước 1 | mm | 0-5 | N/A | #VALUE! | #VALUE! | DIM | 0.01 | 0.02 |  |  | 0.01 | 0.02 | Pick -DIM
Pick |  |  |  | 10: 5-10; 12: N/A; 14: mm; 15: 5-10; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Pick -DIM
Pick |  |  |  | 10: 10-15; 12: N/A; 14: mm; 15: 10-15; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Pick -DIM
Pick |  |  |  | 10: 15-20; 12: N/A; 14: mm; 15: 15-20; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Pick -DIM
Pick |  |  |  | 10: 20-25; 12: N/A; 14: mm; 15: 20-25; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Pick -DIM
Pick | [7 empty] | Dimension of Devided Fractions - Ruler 2 | 0-5 | N/A | Dimension of Devided Fractions - Ruler 2 - mm | Thước 2 | mm | 0-5 | N/A | #VALUE! | #VALUE! | DIM | 0.01 | 0.02 |  |  | 0.01 | 0.02 | Pick -DIM
Pick |  |  |  | 10: 5-10; 12: N/A; 14: mm; 15: 5-10; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Pick -DIM
Pick |  |  |  | 10: 10-15; 12: N/A; 14: mm; 15: 10-15; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Pick -DIM
Pick |  |  |  | 10: 15-20; 12: N/A; 14: mm; 15: 15-20; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Pick -DIM
Pick |  |  |  | 10: 20-25; 12: N/A; 14: mm; 15: 20-25; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Pick -DIM
Pick | [7 empty] | Dimension of Devided Fractions - Ruler 3 | 0-5 | N/A | Dimension of Devided Fractions - Ruler 3 - mm | Thước 3 | mm | 0-5 | N/A | #VALUE! | #VALUE! | DIM | 0.01 | 0.02 |  |  | 0.01 | 0.02 | Pick -DIM
Pick |  |  |  | 10: 5-10; 12: N/A; 14: mm; 15: 5-10; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Pick -DIM
Pick |  |  |  | 10: 10-15; 12: N/A; 14: mm; 15: 10-15; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Pick -DIM
Pick |  |  |  | 10: 15-20; 12: N/A; 14: mm; 15: 15-20; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Pick -DIM
Pick |  |  |  | 10: 20-25; 12: N/A; 14: mm; 15: 20-25; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Pick -DIM
Pick | [7 empty] | Dimension of Devided Fractions - Ruler 4 | 0-5 | N/A | Dimension of Devided Fractions - Ruler 4 - mm | Thước 4 | mm | 0-5 | N/A | #VALUE! | #VALUE! | DIM | 0.01 | 0.02 |  |  | 0.01 | 0.02 | Pick -DIM
Pick |  |  |  | 10: 5-10; 12: N/A; 14: mm; 15: 5-10; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Pick -DIM
Pick |  |  |  | 10: 10-15; 12: N/A; 14: mm; 15: 10-15; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Pick -DIM
Pick |  |  |  | 10: 15-20; 12: N/A; 14: mm; 15: 15-20; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Pick -DIM
Pick |  |  |  | 10: 20-25; 12: N/A; 14: mm; 15: 20-25; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Pick -DIM
Tacho |  | Digital Tacho Metter | TAV-QTHTĐ | Digital Tacho Metter | Digital Tacho Metter | Thiết bị đo tốc độ | Speed | Speed (*) |  |  | Speed (*) - rpm |  | rpm |  |  | 0 | 0 | RPM | 0.1 | 0.1 | 800000 | Service & Verification | 0.1 | 0.1 | Tacho-RPM
Tacho |  |  |  | 12: N/A; 11: *; 14: rpm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RPM; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Tacho-RPM
Tacho |  |  |  | 12: N/A; 11: *; 14: rpm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RPM; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Tacho-RPM
Tacho |  |  |  | 12: N/A; 11: *; 14: rpm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RPM; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Tacho-RPM
Tacho |  |  |  | 12: N/A; 11: *; 14: rpm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RPM; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Tacho-RPM
Tacho |  |  |  | 12: N/A; 11: *; 14: rpm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RPM; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Tacho-RPM
Tacho |  |  |  | 12: N/A; 11: *; 14: rpm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RPM; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Tacho-RPM
Tacho |  |  |  | 12: N/A; 11: *; 14: rpm; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RPM; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Tacho-RPM
Timer |  | Stop Watch, Timer | TAV-QT-TIME | Stop Watch, Timer | Stop Watch, Timer | Đồng hồ bấm giây |  | Display Value | 30 | N/A | Display Value - sec | Giá trị hiển thị | sec | 30 | N/A | #VALUE! | #VALUE! | TIME | 0.001 | 0.02 | 500000 | Service & Verification | 0.001 | 0.02 | Timer-TIME
Timer |  |  |  | 10: 60; 12: N/A; 14: sec; 15: 60; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 20: 0.001; 21: 0.02 |  | 0.001 | 0.02 | Timer-TIME
Timer |  |  |  | 10: 600; 12: N/A; 14: sec; 15: 600; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 20: 0.001; 21: 0.02 |  | 0.001 | 0.02 | Timer-TIME
Timer |  |  |  | 10: 1800; 12: N/A; 14: sec; 15: 1800; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 20: 0.001; 21: 0.02 |  | 0.001 | 0.02 | Timer-TIME
Timer | [7 empty] | Customer's request |  | N/A | Customer's request - sec | Yêu cầu của khách hàng | sec |  | N/A | #VALUE! | #VALUE! | TIME | 0.001 | 0.02 |  |  | 0.001 | 0.02 | Timer-TIME
Timer |  |  |  | 12: N/A; 14: sec; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 20: 0.001; 21: 0.02 |  | 0.001 | 0.02 | Timer-TIME
Timer |  |  |  | 12: N/A; 14: sec; 19: TIME; 20: 0.001; 21: 0.02 |  | 0.001 | 0.02 | Timer-TIME
Shaker |  | Shaker, Shake Water Bath | TAV-QT-TIME TAV-QTHTD TAV-QTHND | Shaker, Shake Water Bath | Shaker, Shake Water Bath | Máy lắc mẫu |  | Display Value - sec | 30 | N/A | Display Value - sec | Giá trị hiển thị | sec | 1800 | N/A | #VALUE! | #VALUE! | TIME | 0.001 | 0.02 | 800000 | Service & Verification | 0.001 | 0.02 | Shaker-TIME
Shaker |  |  |  | 10: 45; 12: N/A; 14: sec; 15: 3600; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 20: 0.001; 21: 0.02 |  | 0.001 | 0.02 | Shaker-TIME
Shaker |  |  |  | 14: sec; 15: 7200; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 20: 0.001; 21: 0.02 |  | 0.001 | 0.02 | Shaker-TIME
Shaker | [7 empty] | Speed (*) - rpm | 30 | ±2 | Speed (*) - rpm | Tốc độ | rpm | 30 | 20 | 28 | 32 | RPM | 0.1 | 0.1 |  |  | 0.1 | 0.1 | Shaker-RPM
Shaker |  |  |  | 10: 60; 12: ±2; 14: rpm; 15: 60; 16: 20; 17: 58; 18: 62; 19: RPM; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Shaker-RPM
Shaker |  |  |  | 10: 90; 12: ±2; 14: rpm; 15: 90; 16: 20; 17: 88; 18: 92; 19: RPM; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Shaker-RPM
Shaker |  |  |  | 10: 120; 12: ±2; 14: rpm; 15: 120; 16: 20; 17: 118; 18: 122; 19: RPM; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Shaker-RPM
Shaker |  |  |  | 10: 150; 12: ±2; 14: rpm; 15: 150; 16: 20; 17: 148; 18: 152; 19: RPM; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Shaker-RPM
Shaker |  |  |  | 10: 200; 12: ±2; 14: rpm; 15: 200; 16: 20; 17: 198; 18: 202; 19: RPM; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Shaker-RPM
Shaker | [5 empty] | Máy kiểm tra độ nhăn |  | Speed Motor (*) - rpm |  | N/A | ư | Tốc độ | rpm |  | N/A | #VALUE! | #VALUE! | RPM | 0.1 | 0.1 |  |  | 0.1 | 0.1 | Shaker-RPM
Shaker | [7 empty] | Temperature (*) - ℃ | 30 | N/A | Temperature (*) - ℃ | Nhiệt độ | ℃ | 30 | N/A | #VALUE! | #VALUE! | DEG | 0.1 | 0.2 |  |  | 0.1 | 0.2 | Shaker-DEG
Shaker |  |  |  | 10: 40; 12: N/A; 14: ℃; 15: 40; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Shaker-DEG
Shaker |  |  |  | 10: 50; 12: N/A; 14: ℃; 15: 50; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Shaker-DEG
Shaker |  |  |  | 10: 60; 12: N/A; 14: ℃; 15: 60; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Shaker-DEG
Shaker | [7 empty] | Times - sec | 60 | N/A | Times - sec | Thời gian | sec | 60 | N/A | #VALUE! | #VALUE! | TIME | 0.001 | 0.02 |  |  | 0.001 | 0.02 | Shaker-TIME
Hygro-thermo |  | Hygro-Thermometer | TAV-QTHND | Hygro-Thermometer | Hygro-Thermometer | Thiết bị nhiệt độ |  | Temperature - current point | 20 | ±1 | Temperature - current point - ⁰C | Nhiệt độ | ℃ | 20 | 1 | 19 | 21 | DEG |  | 0.2 | 800000 | Service & Verification | 0.1 | 0.2 | Hygro-thermo-DEG
Hygro-thermo |  |  |  | 10: 21; 12: ±2; 14: G; 15: 21; 16: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Hygro-thermo-DEG
Hygro-thermo |  |  |  | 10: 22; 12: ±2; 14: G; 15: 22; 16: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Hygro-thermo-DEG
Hygro-thermo |  |  |  | 9: Temperature - ….....; 12: ±2; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Hygro-thermo-DEG
Hygro-thermo |  |  |  | 12: ±2; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Hygro-thermo-DEG
Hygro-thermo |  |  |  | 12: ±2; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Hygro-thermo-DEG |  |  | 0
Hygro-thermo | [7 empty] | Hụmidity - current point | 55 | ±4 | Hụmidity - current point - % RH | Độ ẩm | % RH | 55 | 4 | 51 | 59 | RH |  | 0.3 |  |  | 0.01 | 0.3 | Hygro-thermo-RH
Hygro-thermo |  |  |  | 10: 55; 12: ±2; 14: % RH; 15: 55; 16: 2; 17: 53; 18: 57; 19: RH; 21: 0.3 |  | 0.01 | 0.3 | Hygro-thermo-RH
Hygro-thermo |  |  |  | 10: 65; 12: ±4; 14: % RH; 15: 65; 16: 4; 17: 61; 18: 69; 19: RH; 21: 0.3 |  | 0.01 | 0.3 | Hygro-thermo-RH
Hygro-thermo |  |  |  | 10: 65; 12: ±2; 14: % RH; 15: 65; 16: 2; 17: 63; 18: 67; 19: RH; 21: 0.3 |  | 0.01 | 0.3 | Hygro-thermo-RH
Moisture |  |  | TAV-QTHH | Moisture Humidity Meter - Aquaboy | Moisture Humidity Meter | Thiết bị đo độ ẩm |  | Moisture | 0 | N/A | Moisture - % |  | % | 0 | N/A | #VALUE! | #VALUE! | RH |  | 0.2 |  |  | 0.1 | 0.2 | Moisture-RH |  |  | 0 - Moisture - %
Moisture |  |  |  | 10: 70; 12: N/A; 14: %; 15: 70; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RH; 21: 0.2 |  | 0.1 | 0.2 | Moisture-RH |  |  | 70
Moisture |  |  |  | 10: 100; 12: N/A; 14: %; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RH; 21: 0.2 |  | 0.1 | 0.2 | Moisture-RH |  |  | 100
Moisture |  |  |  | 12: N/A; 11: *; 19: RH | [6 empty] | *
Moisture |  |  |  | 12: N/A; 11: *; 19: RH | [6 empty] | *
Moisture |  |  |  | 12: N/A; 11: *; 19: RH | [6 empty] | *
Thermometer |  | Thermometer | TAV-QTHND | Thermometer | Thermometer | Thiết bị nhiệt độ |  | Temperature (*) - 20⁰C/21⁰C/22⁰C/50⁰C/70⁰C/90⁰C/…....⁰C |  | N/A | Temperature (*) - ℃ | Nhiệt độ | ℃ |  | N/A | #VALUE! | #VALUE! | DEG |  | 0.2 |  |  | 0.1 | 0.2 | Thermometer-DEG |  |  | Temperature (*) - ℃
Thermometer |  |  |  | 12: N/A; 14: ℃; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Thermometer-DEG |  |  | 0
Thermometer |  |  |  | 12: N/A; 14: ℃; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Thermometer-DEG |  |  | 0
Thermometer |  |  |  | 12: N/A; 14: ℃; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Thermometer-DEG |  |  | 0
Thermometer |  |  |  | 12: N/A; 14: ℃; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Thermometer-DEG |  |  | 0
Thermometer |  |  |  | 12: N/A; 14: ℃; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Thermometer-DEG |  |  | 0
Thermometer |  |  |  | 12: N/A; 14: ℃; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Thermometer-DEG |  |  | 0
Thermometer | [7 empty] | Humidity - Indication Value | 70 | N/A | Humidity - Indication Value - % RH | Độ ẩm - Giá trị chỉ thị | % RH | 70 | N/A | #VALUE! | #VALUE! | RH |  | 0.4 |  |  | 0.1 | 0.4 | Thermometer-RH |  |  | 70 - Humidity - Indication Value - % RH
Thermometer |  |  |  | 12: N/A; 11: *; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RH; 21: 0.4 |  | 0.1 | 0.4 | Thermometer-RH |  |  | *
Thermometer |  |  |  | 12: N/A; 11: *; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RH; 21: 0.4 |  | 0.1 | 0.4 | Thermometer-RH |  |  | *
Thermometer | [7 empty] | Hụmidity - current point | 48 | N/A | Hụmidity - current point - % RH | Độ ẩm - Giá trị hiện tại | % RH | 48 | N/A | #VALUE! | #VALUE! | RH |  | 0.4 |  |  | 0.1 | 0.4 | Thermometer-RH |  |  | 48 - Hụmidity - current point - % RH
Thermometer |  |  |  | 12: N/A; 11: *; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RH; 21: 0.4 |  | 0.1 | 0.4 | Thermometer-RH |  |  | *
Thermometer |  |  |  | 12: N/A; 11: *; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: RH; 21: 0.4 |  | 0.1 | 0.4 | Thermometer-RH |  |  | *
Inspect |  | Fabric Inspection Rolling Machine | TAV-QTHTĐ | Fabric Inspection Rolling Machine | Fabric Inspection Rolling Machine | Máy kiểm vải |  | Rotation Speed (*) - rpm | 40 | ±2 | Rotation Speed (*) - rpm | Tốc độ xoay | rpm | 40 | 2 | 38 | 42 | RPM | 0.1 | 0.1 | 1200000 | Service & Verification | 0.1 | 0.1 | Inspect-RPM |  |  | 40 - Rotation Speed (*) - rpm
Inspect |  |  |  | 10: 50; 12: ±2; 14: rpm; 15: 50; 16: 2; 17: 48; 18: 52; 19: RPM; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Inspect-RPM |  |  | 50
Inspect |  |  |  | 9: Counter - Yard; 10: 20; 12: ±2; 11: Counter - Yard; 13: Bộ đếm; 14: Yard; 15: 20; 19: YARD; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Inspect-YARD |  |  | 20 - Counter - Yard
Inspect |  |  |  | 10: 30; 12: ±2; 14: Yard; 15: 30; 19: YARD; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Inspect-YARD |  |  | 30
Inspect |  |  |  | 10: 40; 12: ±2; 14: Yard; 15: 40; 19: YARD; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Inspect-YARD |  |  | 40
Inspect | [7 empty] | Fabric Rolling Speed - Yards/minute | 20 | ±2 | Fabric Rolling Speed - Yards/minute | Tốc độ cuộn vải | Yards/minute | 20 | 2 | 18 | 22 | RPM | 0.1 | 0.1 |  |  | 0.1 | 0.1 | Inspect-RPM |  |  | 20 - Fabric Rolling Speed - Yards/minute
Inspect |  |  |  | 10: 25; 12: ±2; 14: Yards/minute; 15: 25; 16: 2; 17: 23; 18: 27; 19: RPM; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Inspect-RPM
Inspect |  |  |  | 10: Requested by client:; 12: ±2; 11: *; 14: Yards/minute; 16: 2; 17: -2; 18: 2; 19: RPM; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Inspect-RPM
Inspect |  |  |  | 12: ±2; 11: *; 14: Yards/minute; 16: 2; 19: RPM; 20: 0; 21: 0.1 |  | 0.1 | 0.1 | Inspect-RPM
Inspect |  |  |  | 12: ±2; 11: *; 14: Yards/minute; 16: 2; 19: RPM; 20: 0; 21: 18 |  | 0.1 | 0.1 | Inspect-RPM
Inspect |  |  |  | 12: ±2; 11: *; 14: Yards/minute; 16: 2; 19: RPM; 20: 0; 21: 18 |  | 0.1 | 0.1 | Inspect-RPM
Inspect | [7 empty] | Fabric Rolling stroke - Yard |  | ±2 | Fabric Rolling stroke - Yard | Chiều dài vải | Yard |  | 2 | -2 | 2 | Yard | 0 | 0.012 |  |  | 0.1 | 0.1 | Inspect-Yard
Inspect |  |  |  | 12: ±2; 11: *; 14: Yard; 16: 2; 17: -2; 18: 2; 19: Yard; 20: 0; 21: 0.012 |  | 0.1 | 0.1 | Inspect-Yard
Inspect |  |  |  | 12: ±2; 11: *; 14: Yard; 16: 2; 17: -2; 18: 2; 19: Yard; 20: 0; 21: 0.012 |  | 0.1 | 0.1 | Inspect-Yard
Inspect |  |  |  | 12: ±2; 11: *; 14: Yard; 16: 2; 17: -2; 18: 2; 19: Yard; 20: 0.1; 21: 0.1; 22: 1200000 | Service & Verification | 0.1 | 0.1 | Inspect-Yard
Inspect |  |  |  | 12: ±2; 11: *; 14: Yard; 16: 2; 17: -2; 18: 2; 19: Yard; 20: 0.1; 21: 0.1 |  | 0.1 | 0.1 | Inspect-Yard
Inspect |  |  |  | 9: Illuminance (*) - Lux; 10: Left; 12: N/A; 11: Illuminance (*) - Lux; 13: Độ rọi; 14: Lux; 19: LUX |  |  |  | Inspect-LUX
Inspect |  |  |  | 10: Middle; 12: N/A; 14: Lux; 19: LUX |  |  |  | Inspect-LUX
Inspect |  |  |  | 10: Right; 12: N/A; 14: Lux; 19: LUX |  |  |  | Inspect-LUX
Velcro | Adidas ST-07 SATRA TM123 ISO 22776 | Different strength between layers of textile touch and close fasteners | TAV-QTHTD | Velcro machine | Different strength between layers of textile touch and close fasteners | Máy thử độ mỏi của đế giày | Rotating speed Time | Rotating speed (*) - rpm | 60 | ±5 | Rotating speed (*) - rpm | Tốc độ | rpm | 60 | 5 | 55 | 65 | RPM | 0.1 | 0.1 | 800000 | Service & Verification | 0.1 | 0.1 | Velcro-RPM |  | Y
Velcro | [7 empty] | Counter - Lần/min |  | N/A | Counter - Lần/min | Bộ đếm | Lần/min | N/A | N/A | #VALUE! | #VALUE! | TIM1 |  | 0.02 |  |  | 0.001 | 0.02 | Velcro-TIM1
Velcro | [7 empty] | Weight - g |  | N/A | Weight - g | Tải trọng | g | N/A | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Velcro-MASS
Velcro |  |  |  | 12: N/A; 14: g; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Velcro-MASS
Velcro |  |  |  | 12: N/A; 14: g; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Velcro-MASS
Slide |  | Slide fastness reciprocating tester |  | 5: Slide fastness reciprocating tester; 6: Slide fastness reciprocating tester; 17: 0; 18: 0; 20: #N/A; 21: #N/A; 22: 1200000 | Service & Verification | #N/A | #N/A | Slide- |  | Y
Duro-shore | Adidas GE-09 ASTM D2240 | Durometer, hardness durometer | TAV-QTHKL | Durometer, hardness durometer, Asker | Durometer, hardness durometer | Đồng hồ đo độ cứng/mềm | Weight Reciprocating speed | Weight …............ G - kg Tải trọng | 10 | N/A | * | Tải trọng | kg | 10 | N/A | #VALUE! | #VALUE! | MASS | 0.01 | 0.1 | 1000000 | Service & Verification | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 10: 20; 12: N/A; 14: kg; 15: 20; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 10: 30; 12: N/A; 14: kg; 15: 30; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 10: 40; 12: N/A; 14: kg; 15: 40; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 10: 50; 12: N/A; 14: kg; 15: 50; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 10: 60; 12: N/A; 14: kg; 15: 60; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 10: 70; 12: N/A; 14: kg; 15: 70; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 10: 80; 12: N/A; 14: kg; 15: 80; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 10: 90; 12: N/A; 14: kg; 15: 90; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 10: 100; 12: N/A; 14: kg; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore | [7 empty] | Diameter pin | 1.25 | ±0.15 | Diameter pin - mm | Đường kính | mm | 1.25 | 0.15 | 1.1 | 1.4 | DIM | 0.01 | 0.02 |  |  | 0.01 | 0.02 | Duro-shore-DIM |  | Y
Duro-shore | [7 empty] | Height pin | 2.5 | ±0.04 | Height pin - mm | Chiều cao | mm | 2.5 | 0.04 | 2.46 | 2.54 | DIM | 0.01 | 0.02 |  |  | 0.01 | 0.02 | Duro-shore-DIM |  | Y
Duro-shore |  |  |  | 9: Weight …............ (g/kg) Tải trọng; 11: *; 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 19: MASS; 20: 0.01; 21: 0.1 |  | 0.01 | 0.1 | Duro-shore-MASS |  | Y
Duro-shore |  |  |  | 9: Diameter pin; 11: Diameter pin - mm; 13: Đường kính; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Duro-shore-DIM |  | Y
Duro-shore |  |  |  | 9: Height pin; 11: Height pin - mm; 13: Chiều cao; 19: DIM; 20: 0.01; 21: 0.02 |  | 0.01 | 0.02 | Duro-shore-DIM |  | Y
Chamber-hydro |  | Humidity/ Temperature chamber for Hydrolysis, Blooming Test | TAV-QTHND | 5: Humidity/ Temperature Chamber; 6: Humidity/ Temperature chamber for Hydrolysis, Blooming Test; 8: Temperature Humidity; 9: GT hiển thị; 11: *; 20: #N/A; 22: 1200000 | Service & Calibration - 9 points | #N/A | #N/A | Chamber-hydro-
Chamber-hydro | [6 empty] | Temperature Humidity | Temperature (*) - 55°C/70°C/80°C/..…°C Nhiệt độ | 1 - | ±2 | * | Nhiệt độ | ℃ |  | 2 | -2 | 2 | DEG | 0.1 |  | 1200000 | Service & Calibration - 9 points | 0.1 | 0.2 | Chamber-hydro-DEG |  |  | 1 - - *
Chamber-hydro |  |  |  | 10: 2 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG |  |  | 2 - - *
Chamber-hydro |  |  |  | 10: 3 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG |  |  | 3 - - *
Chamber-hydro |  |  |  | 10: 4 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG |  |  | 4 - - *
Chamber-hydro |  |  |  | 10: 5 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG |  |  | 5 - - *
Chamber-hydro |  |  |  | 10: 6 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG |  |  | 6 - - *
Chamber-hydro |  |  |  | 10: 7 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG |  |  | 7 - - *
Chamber-hydro |  |  |  | 10: 8 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG |  |  | 8 - - *
Chamber-hydro |  |  |  | 10: 9 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG
Chamber-hydro |  |  |  | 9: GT hiển thị; 12: ±2; 11: *
Chamber-hydro |  |  |  | 9: Temperature (*) - 55°C/70°C/80°C/..…°C Nhiệt độ; 10: 1 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG
Chamber-hydro |  |  |  | 10: 2 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG
Chamber-hydro |  |  |  | 10: 3 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG
Chamber-hydro |  |  |  | 10: 4 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG
Chamber-hydro |  |  |  | 10: 5 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG |  |  | 5 - - *
Chamber-hydro |  |  |  | 10: 6 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG |  |  | 6 - - *
Chamber-hydro |  |  |  | 10: 7 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG
Chamber-hydro |  |  |  | 10: 8 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG |  |  | 8 - - *
Chamber-hydro |  |  |  | 10: 9 -; 12: ±2; 11: *; 14: ℃; 16: 2; 17: -2; 18: 2; 19: DEG; 20: 0.1 |  | 0.1 | 0.2 | Chamber-hydro-DEG
Chamber-hydro | [6 empty] | Humidity | Humidity - % | 50 | ±5 | Humidity - % | Độ ẩm | % | 50 | 5 | 45 | 55 | RH | 0.1 |  |  |  | 0.1 | 0.4 | Chamber-hydro-RH |  |  | 50 - Humidity - %
Chamber-hydro |  |  |  | 10: 70; 12: ±5; 14: %; 15: 70; 16: 5; 17: 65; 18: 75; 19: RH; 20: 0.1 |  | 0.1 | 0.4 | Chamber-hydro-RH |  |  | 70
Chamber-hydro |  |  |  | 10: 90; 12: ±5; 14: %; 15: 90; 16: 5; 17: 85; 18: 95; 19: RH; 20: 0.1 |  | 0.1 | 0.4 | Chamber-hydro-RH
Chamber-cold |  | Cold Flexing Chamber | TAV-QTHND | 5: Cold Flexing Chamber; 6: Freezer for Ross Flexing and Bally flexometer; 8: Temperature Humidity; 9: GT hiển thị; 11: *; 20: #N/A; 21: #N/A; 22: 800000 | Service & Verification | #N/A | #N/A | Chamber-cold-
Chamber-cold |  |  |  | 9: Temperature (*) - (-5)°C/(-10)°C/…...°C Nhiệt độ Nhiệt độ; 10: 1 -; 12: ±2; 11: *; 13: Nhiệt độ; 14: ℃; 16: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 10: 2 -; 12: ±2; 11: *; 14: ℃; 16: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 10: 3 -; 12: ±2; 11: *; 14: ℃; 16: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 10: 4 -; 12: ±2; 11: *; 14: ℃; 16: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 10: 5 -; 12: ±2; 11: *; 14: ℃; 16: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 10: 6 -; 12: ±2; 11: *; 14: ℃; 16: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 10: 7 -; 12: ±2; 11: *; 14: ℃; 16: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 10: 8 -; 12: ±2; 11: *; 14: ℃; 16: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 10: 9 -; 12: ±2; 11: *; 14: ℃; 16: 2; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 9: GT hiển thị; 12: ±2; 11: *; 21: #N/A |  | #N/A | #N/A | Chamber-cold-
Chamber-cold |  |  |  | 9: Temperature (*) - (-5)°C/(-10)°C/…...°C Nhiệt độ Nhiệt độ; 10: 1 -; 12: ±2; 11: *; 14: ℃; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 10: 2 -; 12: ±2; 11: *; 14: ℃; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 10: 3 -; 12: ±2; 11: *; 14: ℃; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 10: 4 -; 12: ±2; 11: *; 14: ℃; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 10: 5 -; 12: ±2; 11: *; 14: ℃; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 10: 6 -; 12: ±2; 11: *; 14: ℃; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 10: 7 -; 12: ±2; 11: *; 14: ℃; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 10: 8 -; 12: ±2; 11: *; 14: ℃; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Chamber-cold |  |  |  | 10: 9 -; 12: ±2; 11: *; 14: ℃; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Chamber-cold-DEG
Condition-room |  | Conditioning Room | TAV-QTHND | Conditioning Lab Room | Conditioning Lab Room | Phòng lab chuẩn | Temperature | Temperature (*) - 21°C/22°C/…...°C Nhiệt độ | 1 - | ±1 | * | Nhiệt độ | ℃ |  | 1 | -1 | 1 | DEG | 0.1 | 0.4 | 6000000 | Service & Calibration - 9 points | 0.1 | 0.2 | Condition-room-DEG |  |  | 1 - - *
Condition-room |  |  |  | 10: 2 -; 12: ±1; 14: ℃; 16: 1; 17: -1; 18: 1; 19: DEG; 20: 0.1; 21: 0.4 |  | 0.1 | 0.2 | Condition-room-DEG |  |  | 2 -
Condition-room |  |  |  | 10: 3 -; 12: ±1; 14: ℃; 16: 1; 17: -1; 18: 1; 19: DEG; 20: 0.1; 21: 0.4 |  | 0.1 | 0.2 | Condition-room-DEG
Condition-room |  |  |  | 10: 4 -; 12: ±1; 14: ℃; 16: 1; 17: -1; 18: 1; 19: DEG; 20: 0.1; 21: 0.4 |  | 0.1 | 0.2 | Condition-room-DEG |  |  | 4 -
Condition-room |  |  |  | 10: 5 -; 12: ±1; 14: ℃; 16: 1; 17: -1; 18: 1; 19: DEG; 20: 0.1; 21: 0.4 |  | 0.1 | 0.2 | Condition-room-DEG
Condition-room |  |  |  | 10: 6 -; 12: ±1; 14: ℃; 16: 1; 17: -1; 18: 1; 19: DEG; 20: 0.1; 21: 0.4 |  | 0.1 | 0.2 | Condition-room-DEG
Condition-room |  |  |  | 10: 7 -; 12: ±1; 14: ℃; 16: 1; 17: -1; 18: 1; 19: DEG; 20: 0.1; 21: 0.4 |  | 0.1 | 0.2 | Condition-room-DEG
Condition-room |  |  |  | 10: 8 -; 12: ±1; 14: ℃; 16: 1; 17: -1; 18: 1; 19: DEG; 20: 0.1; 21: 0.4 |  | 0.1 | 0.2 | Condition-room-DEG
Condition-room |  |  |  | 10: 9 -; 12: ±1; 14: ℃; 16: 1; 17: -1; 18: 1; 19: DEG; 20: 0.1; 21: 0.4 |  | 0.1 | 0.2 | Condition-room-DEG
Condition-room | [7 empty] | Humidity - 65% | 1 - 65% | ±2/±4 | Humidity - 65% - % | Độ ẩm | % | 65 |  | 65 | 65 | RH | 0.1 | 0.2 |  |  | 0.1 | 0.4 | Condition-room-RH
Condition-room |  |  |  | 10: 2 - 65%; 12: ±2/±4; 14: %; 15: 65; 17: 65; 18: 65; 19: RH; 20: 0.1; 21: 0.2 |  | 0.1 | 0.4 | Condition-room-RH
Condition-room |  |  |  | 10: 3 - 65%; 12: ±2/±4; 14: %; 15: 65; 17: 65; 18: 65; 19: RH; 20: 0.1; 21: 0.2 |  | 0.1 | 0.4 | Condition-room-RH
Condition-room |  |  |  | 10: 4 - 65%; 12: ±2/±4; 14: %; 15: 65; 17: 65; 18: 65; 19: RH; 20: 0.1; 21: 0.2 |  | 0.1 | 0.4 | Condition-room-RH
Condition-room |  |  |  | 10: 5 - 65%; 12: ±2/±4; 14: %; 15: 65; 17: 65; 18: 65; 19: RH; 20: 0.1; 21: 0.2 |  | 0.1 | 0.4 | Condition-room-RH
Condition-room |  |  |  | 10: 6 - 65%; 12: ±2/±4; 14: %; 15: 65; 17: 65; 18: 65; 19: RH; 20: 0.1; 21: 0.2 |  | 0.1 | 0.4 | Condition-room-RH
Condition-room |  |  |  | 10: 7 - 65%; 12: ±2/±4; 14: %; 15: 65; 17: 65; 18: 65; 19: RH; 20: 0.1; 21: 0.2 |  | 0.1 | 0.4 | Condition-room-RH
Condition-room |  |  |  | 10: 8 - 65%; 12: ±2/±4; 14: %; 15: 65; 17: 65; 18: 65; 19: RH; 20: 0.1; 21: 0.2 |  | 0.1 | 0.4 | Condition-room-RH
Condition-room |  |  |  | 10: 9 - 65%; 12: ±2/±4; 14: %; 15: 65; 17: 65; 18: 65; 19: RH; 20: 0.1; 21: 0.2 |  | 0.1 | 0.4 | Condition-room-RH
Condition-unit |  | Conditioning Unit, Precision Air Conditioning | TAV-QTHND | 5: Conditioning Unit, Precision Air Conditioning, STA Branca; 6: Branca Idealair Type "2/S" Climate Units; 7: Hệ thống vi khí hậu Branca; 8: Temperature; 9: GT hiển thị; 11: *; 21: #N/A |  | #N/A | #N/A | Condition-unit- |  |  | *
Condition-unit | [6 empty] | Temperature | Temperature (*) - 21°C/22°C/…...°C Nhiệt độ | 1 - | ±1 | * | Nhiệt độ | ℃ |  | 1 | -1 | 1 | DEG | 0.1 | 0.2 |  |  | 0.1 | 0.2 | Condition-unit-DEG |  |  | 1 - - *
Condition-unit |  |  |  | 10: 2 -; 12: ±1; 11: *; 14: ℃; 16: 1; 17: -1; 18: 1; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Condition-unit-DEG |  |  | 2 - - *
Condition-unit |  |  |  | 10: 3 -; 12: ±1; 11: *; 14: ℃; 16: 1; 17: -1; 18: 1; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Condition-unit-DEG |  |  | 3 - - *
Condition-unit |  |  |  | 10: 4 -; 12: ±1; 11: *; 14: ℃; 16: 1; 17: -1; 18: 1; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Condition-unit-DEG |  |  | 4 - - *
Condition-unit |  |  |  | 10: 5 -; 12: ±1; 11: *; 14: ℃; 16: 1; 17: -1; 18: 1; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Condition-unit-DEG |  |  | 5 - - *
Condition-unit |  |  |  | 10: 6 -; 12: ±1; 11: *; 14: ℃; 16: 1; 17: -1; 18: 1; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Condition-unit-DEG
Condition-unit |  |  |  | 10: 7 -; 12: ±1; 11: *; 14: ℃; 16: 1; 17: -1; 18: 1; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Condition-unit-DEG |  |  | 7 - - *
Condition-unit |  |  |  | 10: 8 -; 12: ±1; 11: *; 14: ℃; 16: 1; 17: -1; 18: 1; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Condition-unit-DEG |  |  | 8 - - *
Condition-unit |  |  |  | 10: 9 -; 12: ±1; 11: *; 14: ℃; 16: 1; 17: -1; 18: 1; 19: DEG; 20: 0.1; 21: 0.2 |  | 0.1 | 0.2 | Condition-unit-DEG |  |  | 9 - - *
Condition-unit | [7 empty] | Humidity - 65%/….....% | 1 - 65% | ±2/±4 | * | Độ ẩm | % | 65 |  | 65 | 65 | RH | #N/A | #N/A |  |  | #N/A | #N/A | Condition-unit-RH |  |  | 1 - 65% - *
Condition-unit |  |  |  | 10: 2 -; 12: ±2/±4; 11: *; 14: %; 17: 0; 18: 0; 19: RH; 20: #N/A; 21: #N/A |  | #N/A | #N/A | Condition-unit-RH |  |  | 2 - - *
Condition-unit |  |  |  | 10: 3 -; 12: ±2/±4; 11: *; 14: %; 17: 0; 18: 0; 19: RH; 20: #N/A; 21: #N/A |  | #N/A | #N/A | Condition-unit-RH
Condition-unit |  |  |  | 10: 4 -; 12: ±2/±4; 11: *; 14: %; 17: 0; 18: 0; 19: RH; 20: #N/A; 21: #N/A |  | #N/A | #N/A | Condition-unit-RH
Condition-unit |  |  |  | 10: 5 -; 12: ±2/±4; 11: *; 14: %; 17: 0; 18: 0; 19: RH; 20: #N/A; 21: #N/A |  | #N/A | #N/A | Condition-unit-RH
Condition-unit |  |  |  | 10: 6 -; 12: ±2/±4; 11: *; 14: %; 17: 0; 18: 0; 19: RH; 20: #N/A; 21: #N/A |  | #N/A | #N/A | Condition-unit-RH
Condition-unit |  |  |  | 10: 7 -; 12: ±2/±4; 11: *; 14: %; 17: 0; 18: 0; 19: RH; 20: #N/A; 21: #N/A |  | #N/A | #N/A | Condition-unit-RH
Condition-unit |  |  |  | 10: 8 -; 12: ±2/±4; 11: *; 14: %; 17: 0; 18: 0; 19: RH; 20: #N/A; 21: #N/A |  | #N/A | #N/A | Condition-unit-RH
Condition-unit |  |  |  | 10: 9 -; 12: ±2/±4; 11: *; 14: %; 17: 0; 18: 0; 19: RH; 20: #N/A; 21: #N/A |  | #N/A | #N/A | Condition-unit-RH
Heel | Adidas GE-11 | Heel Fatigue, Heel Impact Tester |  | Heel Fatigue, Heel Impact Tester | Heel Fatigue, Heel Impact Tester |  | Angle Weight | Angle |  |  | Angle - ⁰ | Góc | ⁰ |  |  | 0 | 0 | ANG |  | #N/A | 1000000 | Service & Verification | #N/A | #N/A | Heel-ANG
Heel | Adidas GE-11 | Heel Fatigue, Heel Impact Tester |  | 5: Heel Fatigue, Heel Impact Tester; 6: Heel Fatigue, Heel Impact Tester; 9: Weight; 11: Weight - g; 13: Tải trọng; 14: g; 17: 0; 18: 0; 19: MASS; 21: #N/A |  | #N/A | #N/A | Heel-MASS
DIN | Adidas GE-27, GE80 SATRA TM174 ISO 4649 ISO 2087 ASTM D5963 | Abrasion of outsole materials | TAV-QTHTĐ | DIN Abrasion Test | DIN Abrasion Test |  | Weight Frequency of drum per minute Time for one run | Weight | 255g (2.5N) | N/A | Weight - g | Tải trọng | g | 255 | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 | 1200000 | Service & Verification | 0.01 | 0.1 | DIN-MASS
DIN | Adidas GE-27, GE80 | Abrasion of outsole materials | TAV-QTHTĐ | DIN Abrasion Test | DIN Abrasion Test |  |  |  | 510g (5N) | N/A |  |  | g | 510 | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 |  |  | 0.01 | 0.1 | DIN-MASS
DIN |  | Abrasion of outsole materials | TAV-QTHTĐ | DIN Abrasion Test | DIN Abrasion Test |  |  | Sample dimension | 16mm | N/A | Sample dimension - mm | Kích thước mẫu | mm | 16 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | DIN-DIM
DIN |  | Abrasion of outsole materials | TAV-QTHTĐ | DIN Abrasion Test | DIN Abrasion Test |  |  | Speed | 4.2m/v | N/A | Speed - m/v | Tốc độ | m/v | 4.2 | N/A | #VALUE! | #VALUE! | RPM |  | 0.1 |  |  | 0.1 | 0.1 | DIN-RPM
DIN |  | Abrasion of outsole materials | TAV-QTHTĐ | DIN Abrasion Test | DIN Abrasion Test |  |  | Wear Wheel diameter | 150mm | N/A | Wear Wheel diameter - mm | Đường kính | mm | 150 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | DIN-DIM
DIN |  | Abrasion of outsole materials | TAV-QTHTĐ | DIN Abrasion Test | DIN Abrasion Test |  |  | Wear Wheel Length | 160mm | N/A | Wear Wheel Length - mm | Chiều dài | mm | 160 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | DIN-DIM
DIN |  | Abrasion of outsole materials | TAV-QTHTĐ | DIN Abrasion Test | DIN Abrasion Test |  |  | Wear Speed (*) | 40rpm | ±1 | Wear Speed (*) - rpm | Tốc độ | rpm | 40 | 1 | 39 | 41 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | DIN-RPM
DIN |  | Abrasion of outsole materials | TAV-QTHTĐ | DIN Abrasion Test | DIN Abrasion Test |  |  | Inclination Angle | 3⁰ | N/A | Inclination Angle - ⁰ | Góc nhọn | ⁰ | 3 | N/A | #VALUE! | #VALUE! | ANG |  | 0.03 |  |  | 0.05 | 0.03 | DIN-ANG
DIN | Adidas GE-28 | Abrasion of outsole materials | TAV-QTHTĐ | DIN Abrasion Test | DIN Abrasion Test |  |  | Frequency of drum per minute | 40 | ±1 | Frequency of drum per minute - drum/min | Tốc độ | drum/min | 40 | 1 | 39 | 41 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | DIN-RPM
DIN | Adidas GE-29 | Abrasion of outsole materials | TAV-QTHTĐ | DIN Abrasion Test | DIN Abrasion Test |  |  | Speed roller (*) | 40 rpm | ±1 | Speed roller (*) - rpm | Tốc độ | rpm | 40 | 1 | 39 | 41 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | DIN-RPM
DIN | Adidas GE-29 | Abrasion of outsole materials | TAV-QTHTĐ | DIN Abrasion Test | DIN Abrasion Test |  |  | Time for one run | 130 | ±5 | Time for one run - sec | Thời gian | sec | 130 | 5 | 125 | 135 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | DIN-TIME
DIN | Adidas GE-30 | Abrasion of outsole materials | TAV-QTHTĐ | DIN Abrasion Test | DIN Abrasion Test |  |  | Counter | 40spm | N/A | Counter - spm | Bộ đếm | spm | 40 | N/A | #VALUE! | #VALUE! | SPM |  | 0.02 |  |  | 0.001 | 0.02 | DIN-SPM
DIN1 | SATRA TM174 | Soling materials abrasion tester |  | Soling materials abrasion tester |  |  | Sample holder head angle Sample holder loading Test piece, initial protrusion bottom of sample holder Diameter of drum Length of drum Drum rotation speed Time for one run Position of test piece relative to axis of drum Sample holder traversing speed Diameter of test specimen holder Abrasion distance | Sample holder head angle | 3⁰ from vertical | 3⁰ from vertical | Sample holder head angle - ⁰ |  | ⁰ | 3⁰ from vertical | 3⁰ from vertical | #VALUE! | #VALUE! | ANG |  | #N/A | 1200000 | Service & Verification | #N/A | #N/A | DIN1-ANG
DIN1 | SATRA TM174 |  |  | Soling materials abrasion tester |  |  |  | Sample holder loading - N | 5 | ±0.2 | Sample holder loading - N |  | N | 5 | 0.2 | 4.8 | 5.2 | FORCES |  | #N/A |  |  | #N/A | #N/A | DIN1-FORCES
DIN1 | SATRA TM174 |  |  | Soling materials abrasion tester |  |  |  | Sample holder loading - N | 10 | ±0.2 | Sample holder loading - N |  | N | 10 | 0.2 | 9.8 | 10.2 | FORCES |  | #N/A |  |  | #N/A | #N/A | DIN1-FORCES
DIN1 | SATRA TM174 |  |  | Soling materials abrasion tester |  |  |  | Test piece, initial protrusion bottom of sample holder | 2 | ±0.2 | Test piece, initial protrusion bottom of sample holder - mm |  | mm | 2 | 0.2 | 1.8 | 2.2 | DIM |  | #N/A |  |  | #N/A | #N/A | DIN1-DIM
DIN1 | SATRA TM174 |  |  | Soling materials abrasion tester |  |  |  | Diameter of drum | 150 | ±0.2 | Diameter of drum - mm |  | mm | 150 | 0.2 | 149.8 | 150.2 | DIM |  | #N/A |  |  | #N/A | #N/A | DIN1-DIM
DIN1 | SATRA TM174 |  |  | Soling materials abrasion tester |  |  |  | Length of drum | minimum 400 | minimum 400 | Length of drum - mm |  | mm | minimum 400 | minimum 400 | #VALUE! | #VALUE! | DIM |  | #N/A |  |  | #N/A | #N/A | DIN1-DIM
DIN1 | SATRA TM174 |  |  | Soling materials abrasion tester |  |  |  | Drum rotation speed | 40 | ±1 | Drum rotation speed - revs:min |  | revs:min | 40 | 1 | 39 | 41 | TIME |  | #N/A |  |  | #N/A | #N/A | DIN1-TIME
DIN1 | SATRA TM174 |  |  | Soling materials abrasion tester |  |  |  | Time for one run | 130 | ±5 | Time for one run - sec |  | sec | 130 | 5 | 125 | 135 | TIME |  | #N/A |  |  | #N/A | #N/A | DIN1-TIME
DIN1 | SATRA TM174 |  |  | Soling materials abrasion tester |  |  |  | Position of test piece relative to axis of drum | less than 1mm | less than 1mm | Position of test piece relative to axis of drum - mm |  | mm | less than 1mm | less than 1mm | #VALUE! | #VALUE! | DIM |  | #N/A |  |  | #N/A | #N/A | DIN1-DIM
DIN1 | SATRA TM174 |  |  | Soling materials abrasion tester |  |  |  | Sample holder traversing speed | 4.2 | ±0.04 | Sample holder traversing speed - mm:rev |  | mm:rev | 4.2 | 0.04 | 4.16 | 4.24 | TIME |  | #N/A |  |  | #N/A | #N/A | DIN1-TIME
DIN1 | SATRA TM174 |  |  | Soling materials abrasion tester |  |  |  | Diameter of test specimen holder | maximum >=16.3 | maximum >=16.3 | Diameter of test specimen holder - mm |  | mm | maximum >=16.3 | maximum >=16.3 | #VALUE! | #VALUE! | DIM |  | #N/A |  |  | #N/A | #N/A | DIN1-DIM
DIN1 | SATRA TM174 |  |  | 5: Soling materials abrasion tester; 10: minimum<=15.5; 12: minimum<=15.5; 14: mm; 15: minimum<=15.5; 16: minimum<=15.5; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: #N/A |  | #N/A | #N/A | DIN1-DIM
DIN1 | SATRA TM174 |  |  | Soling materials abrasion tester |  |  |  | Abrasion distance | 40 | ±0.2 | Abrasion distance - m |  | m | 40 | 0.2 | 39.8 | 40.2 | DIM |  | #N/A |  |  | #N/A | #N/A | DIN1-DIM
Lace | SATRA TM154 ISO 22774 | Lace Abrasion Tester | TAV-QTHKL TAV-QTHTD | Shoe Lace Abrasion Tester | Shoe Lace Abrasion Tester | Máy thử độ bền dây giày |  | Speed - cpm | 60 | ±3 | Speed - cpm | Tốc độ | cpm | 60 | 3 | 57 | 63 | RPM |  | 0.1 | 1000000 | Service & Verification | 0.1 | 0.1 | Lace-RPM
Lace | bar: 5
Lace | [7 empty] | Counter - cpm | 60 | N/A | Counter - cpm | Bộ đếm | cycle | 60 |  | 60 | 60 | TIM1 |  | 0.02 |  |  | 0.001 | 0.02 | Lace-TIM1
Lace | bar: 5
Lace | [7 empty] | Reciprocating stroke - mm | 35 | ±2 | Reciprocating stroke - mm | Hành trình đảo | mm | 35 | 2 | 33 | 37 | DIM
Lace | bar: 5
Lace | [7 empty] | Weight - g | 1 - 250 | ±3 | Weight - g | Tải trọng | g | 250 | 3 | 247 | 253 | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Lace-MASS
Lace |  |  |  | 10: 2 - 250; 12: ±3; 14: g; 15: 250; 16: 3; 17: 247; 18: 253; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Lace-MASS
Lace |  |  |  | 10: 3 - 250; 12: ±3; 14: g; 15: 250; 16: 3; 17: 247; 18: 253; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Lace-MASS
Lace |  |  |  | 10: 4 - 250; 12: ±3; 14: g; 15: 250; 16: 3; 17: 247; 18: 253; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Lace-MASS
Lace |  |  |  | 10: 5 - 250; 12: ±3; 14: g; 15: 250; 16: 3; 17: 247; 18: 253; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Lace-MASS
Lace |  |  |  | 10: 6 - 250; 12: ±3; 14: g; 15: 250; 16: 3; 17: 247; 18: 253; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Lace-MASS
Lace1 | SATRA TM154 | Lace Abrasion Tester |  | Lace abration machine |  |  | Speed of reciprocation Moving distance Loading force Width of spacer Minimum seperation between stationary & movable clamps Horizontal distance between pully & stationary clamp Width of metal strip Thickness of metal strip Length of metal strip Angle marked on template | Speed of reciprocation | 60 | ±6 | Speed of reciprocation - cycle/min |  | cycle/min | 60 | 6 | 54 | 66 | TIM1 |  | #N/A | 1200000 | Service & Verification | #N/A | #N/A | Lace1-TIM1
Lace1 | SATRA TM154 | Lace Abrasion Tester |  | Lace abration machine |  |  |  | Moving distance | 35 | ±2 | Moving distance - mm |  | mm | 35 | 2 | 33 | 37 | DIM |  | #N/A |  |  | #N/A | #N/A | Lace1-DIM
Lace1 | SATRA TM154 | Lace Abrasion Tester |  | Lace abration machine |  |  |  | Loading force | 250 | ±3 | Loading force - g |  | g | 250 | 3 | 247 | 253 | MASS |  | #N/A |  |  | #N/A | #N/A | Lace1-MASS
Lace1 | SATRA TM154 | Lace Abrasion Tester |  | Lace abration machine |  |  |  | Width of spacer | 25 | ±5 | Width of spacer - mm |  | mm | 25 | 5 | 20 | 30 | DIM |  | #N/A |  |  | #N/A | #N/A | Lace1-DIM
Lace1 | SATRA TM154 | Lace Abrasion Tester |  | 5: Lace abration machine; 9: Minimum seperation between stationary & movable clamps; 10: 280; 12: ±50; 11: Minimum seperation between stationary & movable clamps - mm; 14: mm; 15: 280; 16: 50; 19: DIM; 21: #N/A |  | #N/A | #N/A | Lace1-DIM
Lace1 | SATRA TM154 | Lace Abrasion Tester |  | Lace abration machine |  |  |  | Horizontal distance between pully & stationary clamp | 35 | ±5 | Horizontal distance between pully & stationary clamp - mm |  | mm | 35 | 5 | 30 | 40 | DIM |  | #N/A |  |  | #N/A | #N/A | Lace1-DIM
Lace1 | SATRA TM154 | Lace Abrasion Tester |  | Lace abration machine |  |  |  | Width of metal strip Approx. 25 | 25 | ±5 | Width of metal strip Approx. 25 - mm |  | mm | 25 | 5 | 20 | 30 | DIM |  | #N/A |  |  | #N/A | #N/A | Lace1-DIM
Lace1 | SATRA TM154 | Lace Abrasion Tester |  | Lace abration machine |  |  |  | Thickness of metal strip | Approximately 3mm | Approximately 3mm | Thickness of metal strip - mm |  | mm | Approximately 3mm | Approximately 3mm | #VALUE! | #VALUE! | DIM |  | #N/A |  |  | #N/A | #N/A | Lace1-DIM
Lace1 | SATRA TM154 | Lace Abrasion Tester |  | Lace abration machine |  |  |  | Length of metal strip | Approximately 75mm | Approximately 75mm | Length of metal strip - mm |  | mm | Approximately 75mm | Approximately 75mm | #VALUE! | #VALUE! | DIM |  | #N/A |  |  | #N/A | #N/A | Lace1-DIM
Lace1 | SATRA TM154 | Lace Abrasion Tester |  | Lace abration machine |  |  |  | Angle marked on template | 52.5 | ±5 | Angle marked on template - ⁰ |  | ⁰ | 52.5 | 5 | 47.5 | 57.5 | ANG |  | #N/A |  |  | #N/A | #N/A | Lace1-ANG
Lace-eyelet | SATRA TM93 | Lace Abrasion Tester | TAV-QTHTD | Shoe Lace to Eyelet abration tester | Shoe Lace to Eyelet abration tester | Máy thử độ bền lỗ xâu | Reciprocating stroke of movable clamp Speed Weight Internal diameter of Standard shoes eyelet | Stroke - mm | 76 | ±2 | Stroke - mm | Hành trình | DIM | 76 | 2 | 74 | 78 | DIM |  | 0.02 | 1000000 | Service & Verification | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet
Lace-eyelet | [7 empty] | Speed (*) - rpm | 100 | ±5 | Speed (*) - rpm | Tốc độ | rpm | 100 | 5 | 95 | 105 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Lace-eyelet-RPM
Lace-eyelet | bar: 5; 14
Lace-eyelet | [7 empty] | Weight (*) - g | 455 | ±5 | Weight - g | Tải trọng | g | 455 | 5 | 450 | 460 | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Lace-eyelet-MASS
Lace-eyelet |  |  |  | 10: 455; 12: ±5; 14: g; 15: 455; 16: 5; 17: 450; 18: 460; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Lace-eyelet-MASS
Lace-eyelet |  |  |  | 10: 455; 12: ±5; 14: g; 15: 455; 16: 5; 17: 450; 18: 460; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Lace-eyelet-MASS
Lace-eyelet |  |  |  | 10: 455; 12: ±5; 14: g; 15: 455; 16: 5; 17: 450; 18: 460; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Lace-eyelet-MASS
Lace-eyelet |  |  |  | 10: 455; 12: ±5; 14: g; 15: 455; 16: 5; 17: 450; 18: 460; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Lace-eyelet-MASS
Lace-eyelet |  |  |  | 10: 455; 12: ±5; 14: g; 15: 455; 16: 5; 17: 450; 18: 460; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Lace-eyelet-MASS
Lace-eyelet |  |  |  | 10: 455; 12: ±5; 14: g; 15: 455; 16: 5; 17: 450; 18: 460; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Lace-eyelet-MASS
Lace-eyelet |  |  |  | 10: 455; 12: ±5; 14: g; 15: 455; 16: 5; 17: 450; 18: 460; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Lace-eyelet-MASS
Lace-eyelet |  |  |  | 9: Brass Inner Diameter - mm; 10: 4.5; 12: N/A; 11: Brass Inner Diameter - mm; 13: Đường kính trong ống đồng; 14: mm; 15: 4.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet |  |  |  | 10: 4.5; 12: N/A; 14: mm; 15: 4.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet |  |  |  | 10: 4.5; 12: N/A; 14: mm; 15: 4.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet |  |  |  | 10: 4.5; 12: N/A; 14: mm; 15: 4.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet |  |  |  | 10: 4.5; 12: N/A; 14: mm; 15: 4.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet |  |  |  | 10: 4.5; 12: N/A; 14: mm; 15: 4.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet |  |  |  | 10: 4.5; 12: N/A; 14: mm; 15: 4.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet |  |  |  | 10: 4.5; 12: N/A; 14: mm; 15: 4.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet | [7 empty] | Fiberboard Inner Diameter - mm | 5 | ±0.2 | Fiberboard Inner Diameter - mm | Đường kính trong dây | mm | 5 | 0.2 | 4.8 | 5.2 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet |  |  |  | 10: 5; 12: ±0.2; 14: mm; 15: 5; 16: 0.2; 17: 4.8; 18: 5.2; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet |  |  |  | 10: 5; 12: ±0.2; 14: mm; 15: 5; 16: 0.2; 17: 4.8; 18: 5.2; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet |  |  |  | 10: 5; 12: ±0.2; 14: mm; 15: 5; 16: 0.2; 17: 4.8; 18: 5.2; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet |  |  |  | 10: 5; 12: ±0.2; 14: mm; 15: 5; 16: 0.2; 17: 4.8; 18: 5.2; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet |  |  |  | 10: 5; 12: ±0.2; 14: mm; 15: 5; 16: 0.2; 17: 4.8; 18: 5.2; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet |  |  |  | 10: 5; 12: ±0.2; 14: mm; 15: 5; 16: 0.2; 17: 4.8; 18: 5.2; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet |  |  |  | 10: 5; 12: ±0.2; 14: mm; 15: 5; 16: 0.2; 17: 4.8; 18: 5.2; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Lace-eyelet-DIM
Lace-eyelet | [7 empty] | Counter | 100 | N/A | Counter - cycle | Bộ đếm | cycle | 100 |  | 100 | 100 | TIM1 |  | 0.02 |  |  | 0.001 | 0.02 | Lace-eyelet-TIM1
Elastic tape |  | Elastic tape fatigue |  | Elastic tape fatigue |  |  |  | Speed (*) | 60 | ±5 | Speed (*) - rpm | Tốc độ | rpm | 60 | 5 | 55 | 65 | RPM |  | 0.1 | 1000000 | Service & Verification | 0.1 | 0.1 | Elastic tape-RPM
Elastic tape |  | Elastic tape fatigue |  | Elastic tape fatigue |  |  |  | Stroke length Station 1 | 50 | N/A | Stroke length Station 1 - mm | Hành trình Trạm 1 | mm | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Elastic tape-DIM
Elastic tape |  | Elastic tape fatigue |  | 5: Elastic tape fatigue; 10: 100; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Elastic tape-DIM
Elastic tape |  | Elastic tape fatigue |  | 5: Elastic tape fatigue; 10: 150; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Elastic tape-DIM
Elastic tape |  | Elastic tape fatigue |  | Elastic tape fatigue |  |  |  | Stroke length Station 2 | 50 | N/A | Stroke length Station 2 - mm | Hành trình Trạm 2 | mm | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Elastic tape-DIM
Elastic tape |  | Elastic tape fatigue |  | 5: Elastic tape fatigue; 10: 100; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Elastic tape-DIM
Elastic tape |  | Elastic tape fatigue |  | 5: Elastic tape fatigue; 10: 150; 12: N/A; 14: mm; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Elastic tape-DIM
Elastic tape |  | Elastic tape fatigue |  | Elastic tape fatigue |  |  |  | Counter | 60 | N/A | Counter - min | Bộ đếm | min | 60 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Elastic tape-TIME
Elastic1 | ISO 10768 | Elastic tape fatigue | TAV-QT-DIM | Elastic tape fatigue |  | Elastic tape fatigue | Clamp minimum width Clamp minimum separation Speed | Clamp minimum width | 140 | N/A | Clamp minimum width - mm |  | mm | 140 | N/A | #VALUE! | #VALUE! | DIM |  | #N/A | 800000 | Service & Verification | #N/A | #N/A | Elastic1-DIM
Elastic1 | ISO 10768 | Elastic tape fatigue |  | Elastic tape fatigue |  |  |  | Clamp minimum separation | 60 | ±10 | Clamp minimum separation - mm |  | mm | 60 | 10 | 50 | 70 | DIM |  | #N/A |  |  | #N/A | #N/A | Elastic1-DIM
Elastic1 | ISO 10768 | Elastic tape fatigue |  | Elastic tape fatigue |  |  |  | Speed | 60 | ±5 | Speed - spm |  | spm | 60 | 5 | 55 | 65 | RPM |  | #N/A |  |  | #N/A | #N/A | Elastic1-RPM
Heel1 |  | Heel Fatigue Tester | TAV-QTHTD | Heel Fatigue Tester | Heel Fatigue Tester | Heel Fatigue Tester |  | Impact frequency | 60 | N/A | Impact frequency - spm |  | spm | 60 | N/A | #VALUE! | #VALUE! | SPM |  | 0.02 | 1000000 | Service & Verification | 0.001 | 0.02 | Heel1-SPM
Heel1 |  | Heel Fatigue Tester |  | Heel Fatigue Tester |  |  |  | Angle Sample | 0 | N/A | Angle Sample - ⁰ | Góc | ⁰ | 0 | N/A | #VALUE! | #VALUE! | ANG |  | 0.03 |  |  | 0.05 | 0.03 | Heel1-ANG
Heel1 |  | Heel Fatigue Tester |  | Heel Fatigue Tester |  |  |  | Pendulum | 57 | ±1 | Pendulum - ⁰ |  | ⁰ | 57 | 1 | 56 | 58 | ANG |  | 0.03 |  |  | 0.05 | 0.03 | Heel1-ANG
Heel1 |  | Heel Fatigue Tester |  | Heel Fatigue Tester |  |  |  | Pendulum shaft | 12.5 | ±1 | Pendulum shaft - ⁰ |  | ⁰ | 12.5 | 1 | 11.5 | 13.5 | ANG |  | 0.03 |  |  | 0.05 | 0.03 | Heel1-ANG
Heel1 |  | Heel Fatigue Tester |  | Heel Fatigue Tester |  |  |  | Thickness | 20 | ±1 | Thickness - mm | Độ dày | mm | 20 | 1 | 19 | 21 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel1-DIM
Heel1 |  | Heel Fatigue Tester |  | Heel Fatigue Tester |  |  |  | Impact angle | N/A | N/A | Impact angle - ⁰ | Góc | ⁰ | N/A | N/A | #VALUE! | #VALUE! | ANG |  | 0.03 |  |  | 0.05 | 0.03 | Heel1-ANG
Heel1 |  | Heel Fatigue Tester |  | Heel Fatigue Tester |  |  |  | Speed (*) | 60 | N/A | Speed (*) - rpm | Tốc độ | rpm | 60 | N/A | #VALUE! | #VALUE! | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Heel1-RPM
Heel2 |  | Heel Impact Tester |  | Heel Impact Tester | Heel Impact Tester | Heel Impact Tester |  | Angle | 45 |  | Angle - ⁰ | Góc | ⁰ | 45 |  | 45 | 45 | ANG |  | 0.03 | 800000 | Service & Verification | 0.05 | 0.03 | Heel2-ANG
Heel2 |  | Heel Impact Tester |  | Heel Impact Tester |  |  |  | Specimen holder | 25 |  | Specimen holder - mm | Vòng giữ mẫu | mm | 25 |  | 25 | 25 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel2-DIM
Heel2 |  | Heel Impact Tester |  | Heel Impact Tester |  |  |  | Force | N/A | N/A | Force - mm |  | mm | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel2-DIM
Heel2 |  | Heel Impact Tester |  | Heel Impact Tester |  |  |  | Pendulum | 108 | ±1 | Pendulum - mm |  | mm | 108 | 1 | 107 | 109 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel2-DIM
Heel2 |  | Heel Impact Tester |  | Heel Impact Tester |  |  |  | Pendulum shaft | 25 | ±1 | Pendulum shaft - mm |  | mm | 25 | 1 | 24 | 26 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel2-DIM
Heel2 |  | Heel Impact Tester |  | 5: Heel Impact Tester; 9: Thickness; 10: 50; 12: ±1; 11: Thickness; 15: 50; 16: 1; 21: #N/A |  | #N/A | #N/A | Heel2-
Heel2 |  | Heel Impact Tester |  | Heel Impact Tester |  |  |  | Joules | 1 | N/A | Joules - mm |  | mm | 1 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel2-DIM
Heel2 |  | Heel Impact Tester |  | 5: Heel Impact Tester; 10: 2; 12: N/A; 14: mm; 15: 2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heel2-DIM
Heel2 |  | Heel Impact Tester |  | 5: Heel Impact Tester; 10: 3; 12: N/A; 14: mm; 15: 3; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heel2-DIM
Heel2 |  | Heel Impact Tester |  | 5: Heel Impact Tester; 10: 4; 12: N/A; 14: mm; 15: 4; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heel2-DIM
Heel2 |  | Heel Impact Tester |  | 5: Heel Impact Tester; 10: 5; 12: N/A; 14: mm; 15: 5; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heel2-DIM
Heel2 |  | Heel Impact Tester |  | 5: Heel Impact Tester; 10: 6; 12: N/A; 14: mm; 15: 6; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heel2-DIM
Heel2 |  | Heel Impact Tester |  | 5: Heel Impact Tester; 10: 7; 12: N/A; 14: mm; 15: 7; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heel2-DIM
Heel2 |  | Heel Impact Tester |  | 5: Heel Impact Tester; 10: 8; 12: N/A; 14: mm; 15: 8; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heel2-DIM
Heel2 |  | Heel Impact Tester |  | 5: Heel Impact Tester; 10: 9; 12: N/A; 14: mm; 15: 9; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heel2-DIM
Heel2 |  | Heel Impact Tester |  | Heel Impact Tester |  |  |  | Weight Pendulium | N/A | N/A | Weight Pendulium - kg |  | kg | N/A | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Heel2-MASS
Heel3 | ISO 19953 | Heel Impact Tester |  | Heel Impact Tester |  |  |  | Pendulum | Diameter 108 | ±1 | Pendulum - mm |  | mm | 108 | 1 | 107 | 109 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel3-DIM
Heel3 |  | Heel Impact Tester |  | 5: Heel Impact Tester; 10: Diameter of pendulum shaft 25; 12: ±0.5; 14: mm; 15: 25; 16: 0.5; 17: 24.5; 18: 25.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heel3-DIM
Heel3 |  | Heel Impact Tester |  | 5: Heel Impact Tester; 10: Thickness of pendulum 49; 12: ±2; 14: mm; 15: 49; 16: 2; 17: 47; 18: 51; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Heel3-DIM
Heel3 |  | Heel Impact Tester |  | Heel Impact Tester |  |  |  | Thickness of striker head | 6 | ±0.5 | Thickness of striker head - mm |  | mm | 6 | 0.5 | 5.5 | 6.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel3-DIM
Heel3 |  | Heel Impact Tester |  | Heel Impact Tester |  |  |  | Width of striker head | 25 | ±0.5 | Width of striker head - mm |  | mm | 25 | 0.5 | 24.5 | 25.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel3-DIM
Heel3 |  | Heel Impact Tester |  | Heel Impact Tester |  |  |  | Length of striker head | 35 | ±2 | Length of striker head - mm |  | mm | 35 | 2 | 33 | 37 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel3-DIM
Heel3 |  | Heel Impact Tester |  | Heel Impact Tester |  |  |  | Radius of the rounded striking edge | 3 | ±0.5 | Radius of the rounded striking edge - mm |  | mm | 3 | 0.5 | 2.5 | 3.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel3-DIM
Vapor | SATRA TM172 | Water Vapour Permeability apparatus | TAV-QTHTĐ | Water Vapour Permeability apparatus | Permeability Tester |  | Internal height of test pots Internal diameter of test pots Internal diameter of clamping ring Distance from centre of test pot station to the centre of the axis of rotation of the turntable Clearance between the fan and the front edge of pots Length of fan blades Width of fan blade Angle of fan blades relative to each other Rotation speed of fan Rotation speed of pot holder | Internal height of test pots | 80 | ±10 | Internal height of test pots - mm |  | mm | 80 | 10 | 70 | 90 | DIM |  | #N/A | 1200000 | Service & Verification | #N/A | #N/A | Vapor-DIM
Vapor |  | Water Vapour Permeability apparatus |  | Water Vapour Permeability apparatus |  |  |  | Internal diameter of test pots | 30 | 10 | Internal diameter of test pots - mm |  | mm | 30 | 10 | 20 | 40 | DIM |  | #N/A |  |  | #N/A | #N/A | Vapor-DIM
Vapor |  | Water Vapour Permeability apparatus |  | Water Vapour Permeability apparatus |  |  |  | Internal diameter of clamping ring | 30 | 10 | Internal diameter of clamping ring - mm |  | mm | 30 | 10 | 20 | 40 | DIM |  | #N/A |  |  | #N/A | #N/A | Vapor-DIM
Vapor |  | Water Vapour Permeability apparatus |  | Water Vapour Permeability apparatus |  |  |  | Distance from centre of test pot station to the centre of the axis of rotation of the turntable | 67 | 2 | Distance from centre of test pot station to the centre of the axis of rotation of the turntable - mm |  | mm | 67 | 2 | 65 | 69 | DIM |  | #N/A |  |  | #N/A | #N/A | Vapor-DIM
Vapor |  | Water Vapour Permeability apparatus |  | Water Vapour Permeability apparatus |  |  |  | Clearance between the fan and the front edge of pots | 10 | 5 | Clearance between the fan and the front edge of pots - mm |  | mm | 10 | 5 | 5 | 15 | DIM |  | #N/A |  |  | #N/A | #N/A | Vapor-DIM
Vapor |  | Water Vapour Permeability apparatus |  | Water Vapour Permeability apparatus |  |  |  | Length of fan blades | 90 | N/A | Length of fan blades - mm |  | mm | 90 | N/A | #VALUE! | #VALUE! | DIM |  | #N/A |  |  | #N/A | #N/A | Vapor-DIM
Vapor |  | Water Vapour Permeability apparatus |  | Water Vapour Permeability apparatus |  |  |  | Width of fan blade | 75 | N/A | Width of fan blade - mm |  | mm | 75 | N/A | #VALUE! | #VALUE! | DIM |  | #N/A |  |  | #N/A | #N/A | Vapor-DIM
Vapor |  | Water Vapour Permeability apparatus |  | Water Vapour Permeability apparatus |  |  |  | Angle of fan blades relative to each other | 120 | N/A | Angle of fan blades relative to each other - ℃ |  | ℃ | 120 | N/A | #VALUE! | #VALUE! | DEG |  | #N/A |  |  | #N/A | #N/A | Vapor-DEG
Vapor |  | Water Vapour Permeability apparatus |  | Water Vapour Permeability apparatus |  |  |  | Rotation speed of fan (*) | 1400 | 100 | Rotation speed of fan (*) - rpm |  | rpm | 1400 | 100 | 1300 | 1500 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Vapor-RPM
Vapor |  | Water Vapour Permeability apparatus |  | Water Vapour Permeability apparatus |  |  |  | Rotation speed of pot holder (*) | 75 | 5 | Rotation speed of pot holder (*) - rpm |  | rpm | 75 | 5 | 70 | 80 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Vapor-RPM
Slip |  | Slip Resistence Tester |  | 5: Slip Resistence Tester; 9: Force; 10: 400; 12: ±20; 11: Force; 17: 0; 18: 0; 21: #N/A; 22: 1000000 | Service & Verification | #N/A | #N/A | Slip-
Slip |  | Slip Resistence Tester |  | 5: Slip Resistence Tester; 9: Force; 10: 500; 12: ±25; 11: Force; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Slip-
Slip |  | Slip Resistence Tester |  | 5: Slip Resistence Tester; 9: Angle - ⁰; 10: 7; 11: Angle; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Slip-
Slip |  | Slip Resistence Tester |  | 5: Slip Resistence Tester; 9: Sliding speed; 10: 0.3; 12: ±0.03; 11: Sliding speed; 17: 0; 18: 0; 21: #N/A |  | #N/A | #N/A | Slip-
Shoe | EN ISO 20344:8.4; EN ISO 17707:2005 |  | TAV-QTHKL | Footwear Rigidity Tester | Footwear Rigidity Tester | Thiết bị đo độ bền | Angle Force | Flexing angle 0 ⁰ | 0 | ±1 | Flexing angle 0 ⁰ - ⁰ | Góc | ⁰ | 0 | 1 | -1 | 1 | ANG |  | 0.03 |  |  | 0.05 | 0.03 | Shoe-ANG
Shoe |  |  |  | 10: 10; 12: ±1; 14: ⁰; 15: 10; 16: 1; 17: 9; 18: 11; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Shoe-ANG
Shoe |  |  |  | 10: 15; 12: ±1; 14: ⁰; 15: 15; 16: 1; 17: 14; 18: 16; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Shoe-ANG
Shoe |  |  |  | 10: 20; 12: ±1; 14: ⁰; 15: 20; 16: 1; 17: 19; 18: 21; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Shoe-ANG
Shoe |  |  |  | 10: 25; 12: ±1; 14: ⁰; 15: 25; 16: 1; 17: 24; 18: 26; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Shoe-ANG
Shoe |  |  |  | 10: 30; 12: ±1; 14: ⁰; 15: 30; 16: 1; 17: 29; 18: 31; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Shoe-ANG
Shoe |  |  |  | 10: 35; 12: ±1; 14: ⁰; 15: 35; 16: 1; 17: 34; 18: 36; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Shoe-ANG
Shoe |  |  |  | 10: 40; 12: ±1; 14: ⁰; 15: 40; 16: 1; 17: 39; 18: 41; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Shoe-ANG
Shoe |  |  |  | 10: 45; 12: ±1; 14: ⁰; 15: 45; 16: 1; 17: 44; 18: 46; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Shoe-ANG
Shoe | [7 empty] | Loadcell Force 20kg - N | 0.5 | N/A | Loadcell Force 20kg - N | Lực | N | 0.5 | N/A | #VALUE! | #VALUE! | FORCES |  | 0.1 |  |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 1; 12: N/A; 14: N; 15: 1; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 2; 12: N/A; 14: N; 15: 2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 3; 12: N/A; 14: N; 15: 3; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 4.01; 12: N/A; 14: N; 15: 4.01; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 5.02; 12: N/A; 14: N; 15: 5.02; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 6.03; 12: N/A; 14: N; 15: 6.03; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 7.01; 12: N/A; 14: N; 15: 7.01; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 8.02; 12: N/A; 14: N; 15: 8.02; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 9.01; 12: N/A; 14: N; 15: 9.01; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 10.02; 12: N/A; 14: N; 15: 10.02; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 11.02; 12: N/A; 14: N; 15: 11.02; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 12.01; 12: N/A; 14: N; 15: 12.01; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 13.02; 12: N/A; 14: N; 15: 13.02; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 14.01; 12: N/A; 14: N; 15: 14.01; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 15.02; 12: N/A; 14: N; 15: 15.02; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 16.01; 12: N/A; 14: N; 15: 16.01; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 17.02; 12: N/A; 14: N; 15: 17.02; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Shoe |  |  |  | 10: 18.01; 12: N/A; 14: N; 15: 18.01; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: FORCES; 21: 0.1 |  | 0.01 | 0.1 | Shoe-FORCES
Crease |  |  | TAV-QTHKL | Crease Recovery Tester | Crease Recovery Tester | Thiết bị đo góc hồi nhàu | Weight Angle | Weight | 1 | N/A | Weight - g | Tải trọng | g | 1 | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Crease-MASS
Crease |  |  |  | 10: 2; 12: N/A; 14: g; 15: 2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Crease-MASS
Crease |  |  |  | 10: 3; 12: N/A; 14: g; 15: 3; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Crease-MASS
Crease | [7 empty] | Chân tạ | N/A | N/A | Chân tạ - g | Chân tạ | g | N/A | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Crease-MASS
Crease | [7 empty] | Angle - ⁰ | Left 30 | N/A | Angle - ⁰ | Góc | ⁰ | N/A | N/A | #VALUE! | #VALUE! | ANG |  | 0.03 |  |  | 0.05 | 0.03 | Crease-ANG
Crease |  |  |  | 10: Left 60; 12: N/A; 14: ⁰; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Crease-ANG
Crease |  |  |  | 10: Left 90; 12: N/A; 14: ⁰; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Crease-ANG
Crease |  |  |  | 10: Right 30; 12: N/A; 14: ⁰; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Crease-ANG
Crease |  |  |  | 10: Right 60; 12: N/A; 14: ⁰; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Crease-ANG
Crease |  |  |  | 10: Right 90; 12: N/A; 14: ⁰; 15: N/A; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Crease-ANG
Wrap |  |  | TAV - QTHTD TAV - QT.06-DIM | Motor Wrap Reel | Motor Wrap Reel | Thiết bị guồng sợi | Speed Counter Diameter | Speed (*) - rpm | 200 | ±2 | Speed (*) - rpm | Tốc độ | rpm | 200 | 2 | 198 | 202 | RPM |  | 0.1 | 1200000 |  | 0.1 | 0.1 | Wrap-RPM
Wrap | [7 empty] | Counter | 1 | N/A | Counter - cycle | Bộ đếm | cycle | 100 | N/A | #VALUE! | #VALUE! | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Wrap-TIME
Wrap |  |  |  | 10: 2; 12: N/A; 14: cycle; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wrap-TIME
Wrap |  |  |  | 10: 3; 12: N/A; 14: cycle; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wrap-TIME
Wrap |  |  |  | 10: 4; 12: N/A; 14: cycle; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wrap-TIME
Wrap |  |  |  | 10: 5; 12: N/A; 14: cycle; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wrap-TIME
Wrap |  |  |  | 10: 6; 12: N/A; 14: cycle; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wrap-TIME
Wrap |  |  |  | 10: 7; 12: N/A; 14: cycle; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wrap-TIME
Wrap |  |  |  | 10: 8; 12: N/A; 14: cycle; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wrap-TIME
Wrap |  |  |  | 10: 9; 12: N/A; 14: cycle; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wrap-TIME
Wrap |  |  |  | 10: 10; 12: N/A; 14: cycle; 15: 100; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Wrap-TIME
Wrap | [7 empty] | Circumference of winch | 1000 | ±2.8 | Circumference of winch - mm | Chu vi tay quay | mm | 1000 | 2.8 | 997.2 | 1002.8 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Wrap-DIM
Cycling |  |  |  | Cycling Test Machine | Cycling Test Machine |  |  | Drop Distance | 150 | ±5 | Drop Distance - mm |  | mm | 150 | 5 | 145 | 155 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cycling-DIM
Cycling | [7 empty] | Frequency | 20 | ±2 | Frequency - cycle/min |  | cycle/min | 20 | 2 | 18 | 22 | TIME |  | 0.02 |  |  | 0.001 | 0.02 | Cycling-TIME
Cycling | [7 empty] | Length of Spring | 165 | N/A | Length of Spring - mm |  | mm | 165 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cycling-DIM
Cycling | [7 empty] | Outer Diameter of Spring | 34 | ±1 | Outer Diameter of Spring - mm |  | mm | 34 | 1 | 33 | 35 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cycling-DIM
Cycling | [7 empty] | Inner Diameter of Spring | 24 | ±1 | Inner Diameter of Spring - mm |  | mm | 24 | 1 | 23 | 25 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cycling-DIM
Cycling | [7 empty] | Coil Newton | 16 | N/A | Coil Newton - mm |  | mm | 16 | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Cycling-DIM
Cycling | [7 empty] | Weight | 0.5 | N/A | Weight - kg | Tải trọng | kg | 0.5 | N/A | #VALUE! | #VALUE! | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Cycling-MASS
Cycling |  |  |  | 10: 1; 12: N/A; 14: kg; 15: 1; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Cycling-MASS
Cycling |  |  |  | 10: 2; 12: N/A; 14: kg; 15: 2; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Cycling-MASS
Cycling |  |  |  | 10: 3; 12: N/A; 14: kg; 15: 3; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Cycling-MASS
Cycling |  |  |  | 10: 4; 12: N/A; 14: kg; 15: 4; 16: N/A; 17: #VALUE!; 18: #VALUE!; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Cycling-MASS
Needle |  |  | TAV-QT.34-NEE | 5: Needle Detector; 6: Needle Detector; 7: Máy dò kim; 8: - Positions for 9-Point Test - Sensitivity of machine; 9: Metal; 14: K; 17: 0; 18: 0; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 9: Ferrer; 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 9: Brass; 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Needle |  |  |  | 14: K; 19: TIK; 21: 0 |  | 0 | 0 | Needle-TIK
Spectrophotometer |  |  |  | 5: Spectrophotometer; 6: Spectrophotometer; 7: Máy quang phổ; 14: K; 19: TIK; 21: #N/A |  | #N/A | #N/A | Spectrophotometer-TIK
Plotter |  | Digital Plotter | TAV-QT-DIM | Digital Plotter | Digital Plotter | Máy in sơ đồ |  | Dimension of Devided Fractions 5x5 | 5 | N/A | Dimension of Devided Fractions 5x5 - cm | Kích thước các khoảng chia | cm | 5 | N/A |  |  | DIM |  | 0.02 | 400000 | Service & Verification | 0.01 | 0.02 | Plotter-DIM
Plotter |  | Ruler | TAV-QT-DIM | 6: Ruler; 7: Thước thẳng; 10: 5; 12: N/A; 14: cm; 15: 5; 16: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Plotter-DIM
Plotter |  | Ruler | TAV-QT-DIM |  | Ruler | Thước thẳng |  | Dimension of Devided Fractions 10x10 | 10 | N/A | Dimension of Devided Fractions 10x10 - cm | Kích thước các khoảng chia | cm | 10 | N/A |  |  | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Plotter-DIM
Plotter |  | Ruler | TAV-QT-DIM | 6: Ruler; 7: Thước thẳng; 10: 10; 12: N/A; 14: cm; 15: 10; 16: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Plotter-DIM
Plotter |  | Ruler | TAV-QT-DIM |  | Ruler | Thước thẳng |  | Dimension of Devided Fractions 20x20 | 20 | N/A | Dimension of Devided Fractions 20x20 - cm | Kích thước các khoảng chia | cm | 20 | N/A |  |  | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Plotter-DIM
Plotter |  | Ruler | TAV-QT-DIM | 6: Ruler; 7: Thước thẳng; 10: 20; 12: N/A; 14: cm; 15: 20; 16: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Plotter-DIM
Plotter |  | Ruler | TAV-QT-DIM |  | Ruler | Thước thẳng |  | Dimension of Devided Fractions 30x30 | 30 | N/A | Dimension of Devided Fractions 30x30 - cm | Kích thước các khoảng chia | cm | 30 | N/A |  |  | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Plotter-DIM
Plotter |  | Ruler | TAV-QT-DIM | 6: Ruler; 7: Thước thẳng; 10: 30; 12: N/A; 14: cm; 15: 30; 16: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Plotter-DIM
Plotter |  | Ruler | TAV-QT-DIM |  | Ruler | Thước thẳng |  | Dimension of Devided Fractions 40x40 | 40 | N/A | Dimension of Devided Fractions 40x40 - cm | Kích thước các khoảng chia | cm | 40 | N/A |  |  | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Plotter-DIM
Plotter |  | Ruler | TAV-QT-DIM | 6: Ruler; 7: Thước thẳng; 10: 40; 12: N/A; 14: cm; 15: 40; 16: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Plotter-DIM
Plotter |  |  |  |  | Ruler |  |  | Dimension of Devided Fractions 50x50 | 50 | N/A | Dimension of Devided Fractions 50x50 - cm | Kích thước các khoảng chia | cm | 50 | N/A |  |  | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Plotter-DIM
Plotter | Ruler - 100mm |  |  | 6: Ruler; 10: 50; 12: N/A; 14: cm; 15: 50; 16: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Plotter-DIM
Wick-manual | AATCC 197-2013, GB / T 21655.1 | Vertical Wicking Tester | TAV-QT-DIM | 5: Vertical Wicking Tester; 6: Vertical Wicking Tester; 7: Dụng cụ đo độ dẫn nước; 9: Ruler 1 - (mm/inch/…....) Thước 1; 10: …...... (mm/inch/…....); 11: *; 21: #N/A |  | #N/A | #N/A | Wick-manual-
Wick-manual |  |  |  | 10: 1 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 10: 2 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 10: 3 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 10: 4 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 10: 5 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 9: Ruler 2 - (mm/inch/…....) Thước 2; 10: …...... (mm/inch/…....); 21: #N/A |  | #N/A | #N/A | Wick-manual-
Wick-manual |  |  |  | 10: 1 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 10: 2 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 10: 3 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 10: 4 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 10: 5 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 9: Ruler 3 - (mm/inch/…....) Thước 3; 10: …...... (mm/inch/…....); 21: #N/A |  | #N/A | #N/A | Wick-manual-
Wick-manual |  |  |  | 10: 1 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 10: 2 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 10: 3 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 10: 4 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 10: 5 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 9: Ruler 4 - (mm/inch/…....) Thước 4; 10: …...... (mm/inch/…....); 21: #N/A |  | #N/A | #N/A | Wick-manual-
Wick-manual |  |  |  | 10: 1 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 10: 2 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 10: 3 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 10: 4 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Wick-manual |  |  |  | 10: 5 -; 12: N/A; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Wick-manual-DIM
Evaporator |  |  |  |  | Rotary Evaporator | Thiết bị quay bay hơi |  | Speed (*) - rpm | 30 | ±1% | Speed (*) - rpm | Tốc độ | rpm | 30 | 0.3 |  |  | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Evaporator-RPM
Evaporator |  |  |  | 10: 60; 12: ±1%; 14: rpm; 15: 60; 16: 0.6; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Evaporator-RPM
Evaporator |  |  |  | 10: 90; 12: ±1%; 14: rpm; 15: 90; 16: 0.9; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Evaporator-RPM
Evaporator |  |  |  | 10: 120; 12: ±1%; 14: rpm; 15: 120; 16: 1.2; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Evaporator-RPM
Evaporator |  |  |  | 10: 150; 12: ±1%; 14: rpm; 15: 150; 16: 1.5; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Evaporator-RPM
Evaporator |  |  |  | 10: 200; 12: ±1%; 14: rpm; 15: 200; 16: 2; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Evaporator-RPM
Evaporator |  |  |  | 10: 250; 12: ±1%; 14: rpm; 15: 250; 16: 2.5; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Evaporator-RPM
Evaporator |  |  |  | 10: 300; 12: ±1%; 14: rpm; 15: 300; 16: 3; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Evaporator-RPM
Evaporator |  |  |  | 21: #N/A |  | #N/A | #N/A | Evaporator-
Evaporator |  |  |  | 9: Temp (*) - °C; 10: 30; 11: Temp (*) - °C; 13: Nhiệt độ; 14: ℃; 15: 30; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Evaporator-DEG
Evaporator |  |  |  | 10: 40; 14: ℃; 15: 40; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Evaporator-DEG
Evaporator |  |  |  | 10: 50; 14: ℃; 15: 50; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Evaporator-DEG
Evaporator |  |  |  | 10: 70; 14: ℃; 15: 70; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Evaporator-DEG
Evaporator |  |  |  | 10: 80; 14: ℃; 15: 80; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Evaporator-DEG
Evaporator |  |  |  | 10: 90; 14: ℃; 15: 90; 19: DEG; 21: 0.2 |  | 0.1 | 0.2 | Evaporator-DEG
Card-winder |  |  |  | 5: Card Winder; 6: Card Winder; 9: Layer Ply; 10: 1; 11: Layer Ply; 13: Độ dày lớp; 14: mm; 21: 0 |  | 0 | 0 | Card-winder -
Card-winder |  |  |  | 10: 2; 21: 0 |  | 0 | 0 | Card-winder -
Card-winder |  |  |  | 10: 3; 21: 0 |  | 0 | 0 | Card-winder -
Card-winder |  |  |  | 10: 4; 21: 0 |  | 0 | 0 | Card-winder -
Card-winder |  |  |  | 10: 5; 21: 0 |  | 0 | 0 | Card-winder -
Card-winder |  |  |  | 10: 6; 21: 0 |  | 0 | 0 | Card-winder -
Card-winder |  |  |  | 10: 7; 21: 0 |  | 0 | 0 | Card-winder -
Card-winder |  |  |  | 10: 8; 21: 0 |  | 0 | 0 | Card-winder -
Card-winder |  |  |  | 10: 9; 21: 0 |  | 0 | 0 | Card-winder -
Card-winder |  |  |  | 9: Speed (*) - rpm; 10: 1; 11: Speed (*) - rpm; 13: Tốc độ; 14: rpm; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Card-winder -RPM
Card-winder |  |  |  | 10: 2; 14: rpm; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Card-winder -RPM
Card-winder |  |  |  | 10: 3; 14: rpm; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Card-winder -RPM
Card-winder |  |  |  | 10: 4; 14: rpm; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Card-winder -RPM
Card-winder |  |  |  | 10: 5; 14: rpm; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Card-winder -RPM
Card-winder |  |  |  | 10: 6; 14: rpm; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Card-winder -RPM
Card-winder |  |  |  | 10: 7; 14: rpm; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Card-winder -RPM
Card-winder |  |  |  | 10: 8; 14: rpm; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Card-winder -RPM
Card-winder |  |  |  | 10: 9; 14: rpm; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Card-winder -RPM
Card-winder |  |  |  | 9: Layer Distance - mm; 10: 1; 11: Layer Distance - mm; 13: Khoảng cách lớp; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Card-winder -DIM
Card-winder |  |  |  | 10: 2; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Card-winder -DIM
Card-winder |  |  |  | 10: 3; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Card-winder -DIM
Card-winder |  |  |  | 10: 4; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Card-winder -DIM
Card-winder |  |  |  | 10: 5; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Card-winder -DIM
Card-winder |  |  |  | 10: 6; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Card-winder -DIM
Card-winder |  |  |  | 10: 7; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Card-winder -DIM
Card-winder |  |  |  | 10: 8; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Card-winder -DIM
Card-winder |  |  |  | 10: 9; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Card-winder -DIM
Friction-yarn |  |  | TAV-QTHTD TAV-QTHKL | 5: Friction Tester; 6: Friction Tester; 9: Speed - m/min; 10: 1; 11: Speed - m/min; 13: Tốc độ; 14: m/min; 15: 1; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Friction-yarn-RPM
Friction-yarn |  |  |  | 10: 2; 14: m/min; 15: 2; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Friction-yarn-RPM
Friction-yarn |  |  |  | 10: 3; 14: m/min; 15: 3; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Friction-yarn-RPM
Friction-yarn |  |  |  | 10: 4; 14: m/min; 15: 4; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Friction-yarn-RPM
Friction-yarn |  |  |  | 10: 5; 14: m/min; 15: 5; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Friction-yarn-RPM
Friction-yarn |  |  |  | 10: 10; 14: m/min; 15: 10; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Friction-yarn-RPM
Friction-yarn |  |  |  | 10: 15; 14: m/min; 15: 15; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Friction-yarn-RPM
Friction-yarn |  |  |  | 10: 20; 14: m/min; 15: 20; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Friction-yarn-RPM
Friction-yarn |  |  |  | 10: 25; 14: m/min; 15: 25; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Friction-yarn-RPM
Friction-yarn |  |  |  | 10: 30; 14: m/min; 15: 30; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Friction-yarn-RPM
Friction-yarn |  |  |  | 10: 35; 14: m/min; 15: 35; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Friction-yarn-RPM
Friction-yarn |  |  |  | 10: 40; 14: m/min; 15: 40; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Friction-yarn-RPM
Friction-yarn |  |  |  | 11: *; 14: m/min; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Friction-yarn-RPM
Friction-yarn |  |  |  | 9: Weight - g; 10: 1010; 11: Weight - g; 13: Tải trọng; 14: g; 15: 1010; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Friction-yarn-MASS
Friction-yarn |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Friction-yarn-
Friction-yarn |  |  |  | 9: Force - F/cN Lực; 10: 100; 11: *; 14: N; 15: 100; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Friction-yarn-FORCES
Friction-yarn |  |  |  | 10: 200; 14: N; 15: 200; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Friction-yarn-FORCES
Friction-yarn |  |  |  | 10: 500; 14: N; 15: 500; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Friction-yarn-FORCES
Friction-yarn |  |  |  | 10: 1000; 14: N; 15: 1000; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Friction-yarn-FORCES
Friction-yarn |  |  |  | 10: 1010; 14: N; 15: 1010; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Friction-yarn-FORCES
Friction-yarn |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Friction-yarn-
Friction-yarn |  |  |  | 11: *; 21: #N/A |  | #N/A | #N/A | Friction-yarn-
Friction-yarn | bar: 11
Nickel |  | NICKEL REALEASE MACHINE | TAV-QTHTD TAV-QT-DIM | NICKEL RELEASE MACHINE |  | Nickel Release Machine |  | Speed (*) - rpm | 30 | ±2 | Speed (*) - rpm | Tốc độ | rpm | 30 | 2 | 28 | 32 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Nickel-RPM
Nickel |  |  |  | 9: Dimension - mm; 12: N/A; 11: Dimension - mm; 13: Độ dài; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Nickel-DIM
Nickel |  |  |  | 9: Timer - sec; 12: N/A; 11: Timer - sec; 13: Thời gian; 14: sec; 19: TIME; 21: 0.02 |  | 0.001 | 0.02 | Nickel-TIME
Nickel | bar: 11
Nickel | bar: 11
Rain |  |  |  | 5: RAIN PROOF TESTER; 7: Rain Proof Tester; 9: Distance Nozzle to Specimen - mm; 10: 305; 12: N/A; 11: Distance Nozzle to Specimen - mm; 13: Khoảng cách từ vòi phun đến mẫu; 14: mm; 15: 305; 19: DIM
Rain
Rain |  |  |  | 9: Spring clamp - mm; 10: 152x155; 12: N/A; 11: Spring clamp - mm; 13: Kích thước vòng giữ mẫu; 14: mm; 15: 152; 19: DIM
Rain |  |  |  | 12: N/A; 14: mm; 15: 155; 19: DIM
Rain |  |  |  | 9: Masonite board - mm; 10: 190x285; 12: N/A; 11: Masonite board - mm; 14: mm; 15: 190; 19: DIM
Rain |  |  |  | 12: N/A; 14: mm; 15: 285; 19: DIM
Rain |  |  |  | 9: Rain test nozzo - inch; 10: 2.05; 12: N/A; 11: Rain test nozzo - inch; 13: Kich thước vòi phun; 14: inch; 19: DIM
Rain |  |  |  | 10: 2-3/16; 12: N/A; 14: inch; 19: DIM
Rain |  |  |  | 10: 7/32; 12: N/A; 14: inch; 19: DIM
Rain |  |  |  | 9: Temperature (*) - °C; 10: 20; 12: N/A; 11: Temperature (*) - °C; 13: Nhiệt độ; 14: ℃; 15: 20; 19: DEG
Rain |  |  |  | 10: 27; 12: N/A; 14: ℃; 15: 27; 19: DEG
Rain |  |  |  | 9: Nozzo Hole - mm; 10: 0.99; 12: N/A; 11: Nozzo Hole - mm; 13: Đường kính lỗ vò phun; 14: mm; 15: 0.99; 19: DIM
Rain
Rain |  |  |  | 9: Volume - mm; 10: 0.6-0.8; 12: N/A; 11: Water column - mm; 13: Chiều dài ống thể tích; 14: mm; 19: DIM
Rain |  |  |  | 10: 0.8-1.0; 12: N/A; 14: mm; 19: DIM
Rain |  |  |  | 10: 1.0-1.2; 12: N/A; 14: mm; 19: DIM
Rain |  |  |  | 10: 1.2-1.4; 12: N/A; 14: mm; 19: DIM
Rod |  |  | TAV-QT-DIM | 5: Test rod; 7: Test rod; 9: Diameter - mm; 11: Diameter - mm; 13: Đường kính; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Rod-DIM
Rod
Rod
Rod
Flamability |  |  | TAV-QT-TIME | 5: Flamability Tester; 7: Flamability Tester; 9: Timer - sec; 10: 3; 12: N/A; 11: Timer - sec; 13: Thời gian; 14: sec; 15: 3; 19: TIME
Flamability |  |  |  | 10: 5; 12: N/A; 14: sec; 15: 5; 19: TIME
Flamability |  |  |  | 10: 10; 12: N/A; 14: sec; 15: 10; 19: TIME
Flamability |  |  |  | 10: 15; 12: N/A; 14: sec; 15: 15; 19: TIME
Flamability |  |  |  | 10: 20; 12: N/A; 14: sec; 15: 20; 19: TIME
Flamability |  |  |  | 10: 25; 12: N/A; 14: sec; 15: 25; 19: TIME
Flamability |  |  |  | 10: 30; 12: N/A; 14: sec; 15: 30; 19: TIME
Flamability |  |  |  | 12: N/A; 14: sec; 19: TIME
Flamability |  |  |  | 12: N/A; 14: sec; 19: TIME
Flamability |  |  |  | 12: N/A; 14: sec; 19: TIME
Flamability |  |  |  | 12: N/A; 14: sec; 19: TIME
Flamability |  |  |  | 12: N/A; 14: sec; 19: TIME
Flamability |  |  |  | 9: Angle - ⁰; 10: 45; 12: N/A; 14: sec; 19: ANGLE
Flamability |  |  |  | 9: Template - mm; 12: N/A; 11: Template - mm; 13: Kích thước; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 12: N/A; 14: mm; 19: DIM
Flamability |  |  |  | 9: Weight - g; 10: 30; 12: N/A; 11: Weight - g; 13: Tải trọng; 14: g; 19: MASS
Flamability |  |  |  | 12: N/A; 14: g; 19: MASS
Flamability |  |  |  | 12: N/A; 14: g; 19: MASS
Flamability |  |  |  | 12: N/A; 14: g; 19: MASS
Flamability |  |  |  | 12: N/A; 14: g; 19: MASS
Flamability |  |  |  | 12: N/A; 14: g; 19: MASS
Flamability |  |  |  | 12: N/A; 14: g; 19: MASS
Flamability |  |  |  | 12: N/A; 14: g; 19: MASS
Flamability |  |  |  | 12: N/A; 14: g; 19: MASS
Flamability |  |  |  | 12: N/A; 14: g; 19: MASS
Flamability |  |  |  | 12: N/A; 14: g; 19: MASS
Zipper |  |  | TAV-QTHTD | 5: ZIPPER RECIPROCATING; 7: Zipper Reciprocating; 9: Speed -
Zipper |  |  |  | 9: Stroke length -mm
Zipper |  |  |  | 9: Force
Zipper
Bite |  | Bite test | TAV-QTHKL TAV-QT-DIM | 5: Bite test clamp; 6: Bite test clamp; 7: Bite Test; 9: Weight; 17: 0; 18: 0
Bite |  |  |  | 17: 0; 18: 0
Bite | [7 empty] | Outer Diameter |  | N/A | Outer Diameter - mm | Đường kính ngoài | mm | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Bite-DIM
Bite | [7 empty] | Inner Diameter |  | N/A | Inner Diameter - mm | Đường kính trong | mm | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Bite-DIM
Bite | [7 empty] | Dimension |  | N/A | Dimension - mm | Độ dài | mm | N/A | N/A | #VALUE! | #VALUE! | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Bite-DIM
Bite
Bite
Button |  | Button Impact | TAV-QTHKL TAV-QT-DIM | 5: Button Impact; 6: Bite test clamp; 7: Button Impact; 9: Weight - g; 17: 0; 18: 0
Button |  |  |  | 9: Diameter - mm; 11: Diameter - mm; 13: Đường kính; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Button-DIM
Button |  |  |  | 9: Dimension - mm
Button
Button
Ball | ASTM | Fixture for Rattles, Small Ball, Squeeze toys, Teethers | TAV-QT-DIM | 5: Fixture for Rattles, Small Ball, Squeeze toys, Teethers; 7: Fixture for Rattles, Small Ball, Squeeze toys, Teethers; 9: Diameter - mm; 11: Diameter - mm; 13: Đường kính; 14: mm; 17: 0; 18: 0; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Ball-DIM
Ball
Ball
Ball
Ball
Spray-penetration | ISO 18695, AATCC 42-2000 |  | TAV-QT-SPR | Water Resistance Impact Penetration Tester | Water Resistance Impact Penetration Tester | Thiết bị thử độ thấm nước |  | Glass funnel diameter - mm | 152 | N/A | Glass funnel diameter - mm | Đường kính phễu | mm | 152 | 2 | 150 | 154 | DIM |  | 0.02 | 1200000 | Service & Verification | 0.01 | 0.02 | Spray-penetration-DIM |  |  | 152 - Glass funnel diameter - mm
Spray-penetration |  |  |  | 9: Angle of specimen placed - ⁰; 10: 45; 12: N/A; 11: Angle of specimen placed - ⁰; 13: Góc mẫu; 14: ⁰; 15: 45; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Spray-penetration-ANG
Spray-penetration | [7 empty] | Distance from nozzle to speciment - mm | 610 | ±10 | Distance from nozzle to speciment - mm | Khoảng cách từ vòi phun đến mẫu | mm | 610 | 10 | 600 | 620 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Spray-penetration-DIM |  |  | 610 - Distance from nozzle to speciment - mm
Spray-penetration | [7 empty] | Measuring glass volume - ml | 500 | ±10 | Measuring glass volume - ml | Thể tích cốc đo | g | 500 | 10 | 490 | 510 | VOL |  | 0.1 |  |  | 0.01 | 0.1 | Spray-penetration-VOL |  |  | 500 - Measuring glass volume - ml
Spray-penetration | [7 empty] | Spring clip 1 - mm | 152 | ±5 | Spring clip 1 - mm | Vòng kẹp 1 | mm | 152 | 5 | 147 | 157 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Spray-penetration-DIM |  |  | 152 - Spring clip 1 - mm
Spray-penetration | [7 empty] | Spring clip 2 - mm | 152 | ±5 | Spring clip 2 - mm | Vòng kẹp 2 | mm | 152 | 5 | 147 | 157 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Spray-penetration-DIM |  |  | 152 - Spring clip 2 - mm
Spray-penetration | [7 empty] | Total weight - g | 450 | ±5 | Total weight - g | Tải trọng | g | 450 | 5 | 445 | 455 | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Spray-penetration-MASS |  |  | 450 - Total weight - g
Spray-penetration |  |  |  | 9: Spray head nozzle; 11: Spray head nozzle; 13: Đầu phun; 21: #N/A |  | #N/A | #N/A | Spray-penetration- |  |  | Spray head nozzle
Spray-penetration |  |  |  | 9: Diameter 1 - mm; 10: 56; 12: N/A; 11: Diameter 1 - mm; 14: mm; 15: 56; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Spray-penetration-DIM
Spray-penetration |  |  |  | 9: Diameter 2 - mm; 10: 52; 12: N/A; 11: Diameter 2 - mm; 14: mm; 15: 52; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Spray-penetration-DIM
Spray-penetration |  |  |  | 9: Diameter 3 - mm; 10: 51; 12: N/A; 11: Diameter 3 - mm; 14: mm; 15: 51; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Spray-penetration-DIM
Spray-penetration |  |  |  | 9: Diameter 4 - mm; 10: 46; 12: N/A; 11: Diameter 4 - mm; 14: mm; 15: 46; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Spray-penetration-DIM
Spray-penetration |  |  |  | 9: Diameter 5 - mm; 10: 32.5; 12: N/A; 11: Diameter 5 - mm; 14: mm; 15: 32.5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Spray-penetration-DIM
Spray-penetration |  |  |  | 9: Diameter 6 - mm; 10: 18.3; 12: N/A; 11: Diameter 6 - mm; 14: mm; 15: 18.3; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Spray-penetration-DIM
Spray-penetration |  |  |  | 9: Outer holes - 12 - mm; 10: 0.99; 12: N/A; 11: Outer holes - 12 - mm; 14: mm; 15: 0.99; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Spray-penetration-DIM
Spray-penetration |  |  |  | 9: Inner holes - 13 - mm; 10: 0.99; 12: N/A; 11: Inner holes - 13 - mm; 14: mm; 15: 0.99; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Spray-penetration-DIM
Spray-penetration |  |  |  | 11: *; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Spray-penetration-DIM
Spray-penetration |  |  |  | 11: *; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Spray-penetration-DIM
Spray-penetration |  |  |  | 11: *; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Spray-penetration-DIM
Spray-penetration |  |  |  | 11: *; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Spray-penetration-DIM
Spray-penetration |  |  |  | 11: *; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Spray-penetration-DIM
Seam-fatigue |  |  | TAV-QTHKL TAV-QT-DIM | DYNAMIC SEAM FATIGUE TESTER | DYNAMIC SEAM FATIGUE TESTER |  |  | Rubber-faced wheel - mm | 127 | N/A | Rubber-faced wheel - mm | Bánh xe mặt cao su | mm | 127 |  |  |  | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Seam-fatigue-DIM
Seam-fatigue |  |  |  | 9: Rubber-faced wheel - inch; 10: 5; 12: N/A; 11: Rubber-faced wheel - inch; 13: Bánh xe mặt cao su; 14: inch; 15: 5; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Seam-fatigue-DIM
Seam-fatigue
Seam-fatigue |  |  |  | 9: Impact height - mm; 10: 150; 12: N/A; 11: Impact height - mm; 13: Chiều cao tác động; 14: mm; 15: 150; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Seam-fatigue-DIM
Seam-fatigue |  |  |  | 9: Impact height - inch; 10: 6; 12: N/A; 11: Impact height - inch; 13: Chiều cao tác động; 14: inch; 15: 6; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Seam-fatigue-DIM
Seam-fatigue
Seam-fatigue |  |  |  | 9: Mass (*) - kg; 10: 3.75; 12: N/A; 11: Mass (*) - kg; 13: Khối lượng; 14: kg; 15: 3.75; 19: MASS; 21: 0.14 |  | 0.0001 | 0.14 | Seam-fatigue-MASS
Seam-fatigue |  |  |  | 9: Mass (*) - lbs; 10: 8.25; 12: N/A; 11: Mass (*) - lbs; 13: Khối lượng; 14: lb; 15: 8.75; 19: MASS; 21: 0.14 |  | 0.0001 | 0.14 | Seam-fatigue-MASS
Seam-fatigue
Seam-fatigue | [7 empty] | Frequency - cycle/min | 25 | ±2 | Frequency - cycle/min | Tốc độ | cycle/min | 25 | 2 | 23 | 27 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Seam-fatigue-RPM
Crush |  |  | TAV-QTHTD TAV-QTHKN | 5: Edge-Wise Crush Machine; 6: Edge-Wise Crush Machine; 7: Máy thử nén cạnh; 9: Speed - …........ Tốc độ; 11: *; 13: Tốc độ; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Crush-RPM
Crush |  |  |  | 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Crush-RPM
Crush |  |  |  | 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Crush-RPM
Crush |  |  |  | 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Crush-RPM
Crush |  |  |  | 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Crush-RPM
Crush |  |  |  | 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Crush-RPM
Crush |  |  |  | 9: Force (*) - ….... Lực; 10: 1 -; 11: *; 13: Lực; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Crush-FORCE
Crush |  |  |  | 10: 2 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Crush-FORCE
Crush |  |  |  | 10: 3 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Crush-FORCE
Crush |  |  |  | 10: 4 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Crush-FORCE
Crush |  |  |  | 10: 5 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Crush-FORCE
Crush |  |  |  | 10: 6 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Crush-FORCE
Crush |  |  |  | 10: 7 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Crush-FORCE
Crush |  |  |  | 10: 8 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Crush-FORCE
Crush |  |  |  | 10: 9 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Crush-FORCE
Crush |  |  |  | 10: 10 -; 19: FORCE; 21: 0.14 |  | 0.0001 | 0.14 | Crush-FORCE
Bounce |  |  | TAV-QTHTD TAV-QTHKL | 5: BOUNCE BETTY MACHINE; 6: BOUNCE BETTY MACHINE; 9: Mass (*) - …..... Khối lượng; 10: 1 -; 11: *; 13: Khối lượng; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Bounce-MASS
Bounce |  |  |  | 10: 2 -; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Bounce-MASS
Bounce |  |  |  | 10: 3 -; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Bounce-MASS
Bounce |  |  |  | 10: 4 -; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Bounce-MASS
Bounce |  |  |  | 10: 5 -; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Bounce-MASS
Bounce |  |  |  | 10: 6 -; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Bounce-MASS
Bounce |  |  |  | 10: 7 -; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Bounce-MASS
Bounce |  |  |  | 10: 8 -; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Bounce-MASS
Bounce |  |  |  | 10: 9 -; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Bounce-MASS
Bounce |  |  |  | 10: 10 -; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Bounce-MASS
Bounce |  |  |  | 9: Frequency - cycle/min; 10: 1 -; 14: cycle/min; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Bounce-RPM
Bounce |  |  |  | 10: 2 -; 14: cycle/min; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Bounce-RPM
Bounce |  |  |  | 10: 3 -; 14: cycle/min; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Bounce-RPM
Bounce |  |  |  | 10: 4 -; 14: cycle/min; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Bounce-RPM
Bounce |  |  |  | 10: 5 -; 14: cycle/min; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Bounce-RPM
Bounce |  |  |  | 10: 6 -; 14: cycle/min; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Bounce-RPM
Bounce |  |  |  | 10: 7 -; 14: cycle/min; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Bounce-RPM
Bounce |  |  |  | 10: 8 -; 14: cycle/min; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Bounce-RPM
Bounce |  |  |  | 10: 9 -; 14: cycle/min; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Bounce-RPM
Bounce |  |  |  | 10: 10 -; 14: cycle/min; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Bounce-RPM
Verification | bar: 5 to 16; 19 to 20
Verification | bar: 11
Verification | bar: 11
Verification | bar: 11
Verification | bar: 11
Verification | bar: 11
Verification | bar: 11
Verification | bar: 11
Verification | bar: 11
Verification | bar: 11
Verification | bar: 11
Verification | bar: 11
Verification | bar: 11
Verification | bar: 11
Verification | bar: 11
Rub-jis | IS L0849-2013 ,GB/T 3920, JIS L1084-2009 |  | TAV-QTHKL TAV-QT-DIM | JIS Colour Fastness To Rubbing Tester | JIS Colour Fastness To Rubbing Tester | Máy thủ bền màu ma sát JIS |  | Counter - Cycle | 100 | N/A | Counter - Cycle | Bộ đếm | Cycle | 100 |  |  |  | TIM1 |  | 0.02 |  |  | 0.001 | 0.02 | Rub-jis-TIM1
Rub-jis
Rub-jis |  |  |  | 9: Speed - cpm; 10: 30; 12: N/A; 11: Speed - cpm; 14: cpm; 15: 30; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Rub-jis-RPM
Rub-jis
Rub-jis |  |  |  | 9: Abrasion head - mm; 10: 1 - 20x20; 12: N/A; 11: Abrasion head - mm; 13: Đầu mài; 14: mm; 15: 20; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Rub-jis-DIM
Rub-jis |  |  |  | 12: N/A; 14: mm; 15: 20; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Rub-jis-DIM
Rub-jis |  |  |  | 10: 2 - 20x20; 12: N/A; 14: mm; 15: 20; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Rub-jis-DIM
Rub-jis |  |  |  | 12: N/A; 14: mm; 15: 20; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Rub-jis-DIM
Rub-jis |  |  |  | 10: 3 - 20x20; 12: N/A; 14: mm; 15: 20; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Rub-jis-DIM
Rub-jis |  |  |  | 12: N/A; 14: mm; 15: 20; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Rub-jis-DIM
Rub-jis |  |  |  | 10: 4 - 20x20; 12: N/A; 14: mm; 15: 20; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Rub-jis-DIM
Rub-jis |  |  |  | 12: N/A; 14: mm; 15: 20; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Rub-jis-DIM
Rub-jis |  |  |  | 10: 5 - 20x20; 12: N/A; 14: mm; 15: 20; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Rub-jis-DIM
Rub-jis |  |  |  | 12: N/A; 14: mm; 15: 20; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Rub-jis-DIM
Rub-jis |  |  |  | 10: 6 - 20x20; 12: N/A; 14: mm; 15: 20; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Rub-jis-DIM
Rub-jis |  |  |  | 12: N/A; 14: mm; 15: 20; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Rub-jis-DIM
Rub-jis |  |  |  | 9: Down ward force - g; 10: 1 - 200; 12: N/A; 11: Down ward force - g; 13: Lực hướng xuống; 14: g; 15: 200; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Rub-jis-FORCES
Rub-jis |  |  |  | 10: 2 - 200; 12: N/A; 14: g; 15: 200; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Rub-jis-FORCES
Rub-jis |  |  |  | 10: 3 - 200; 12: N/A; 14: g; 15: 200; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Rub-jis-FORCES
Rub-jis |  |  |  | 10: 4 - 200; 12: N/A; 14: g; 15: 200; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Rub-jis-FORCES
Rub-jis |  |  |  | 10: 5 - 200; 12: N/A; 14: g; 15: 200; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Rub-jis-FORCES
Rub-jis |  |  |  | 10: 6 - 200; 12: N/A; 14: g; 15: 200; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Rub-jis-FORCES
Rub-jis |  |  |  | 9: Stroke length - mm; 10: 100; 12: N/A; 11: Stroke length - mm; 13: Độ dài hành trình; 14: mm; 15: 100; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Rub-jis-DIM
Rub-jis |  |  |  | 9: Others; 11: *
Rub-jis | bar: 9; 11
Rub-jis | bar: 9; 11
Rub-jis | bar: 9; 11
Rub-jis | bar: 9; 11
Rub-jis | bar: 9; 11
Heel-fatigue |  |  | TAV-QT-DIM TAV-QTHTD TAV-QTHKL | Heel fatigue Tester | Heel fatigue Tester | Máy kiểm tra độ mỏi gót chân |  | Diameter Pendulum - mm | 57 | ±1 | Diameter Pendulum - mm | Đường kính con lắc | mm | 57 | 1 | 56 | 58 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel-fatigue-DIM
Heel-fatigue
Heel-fatigue | [7 empty] | Thickness pendulum - mm | 20 | ±1 | Thickness pendulum - mm | Độ dày con lắc | mm | 20 | 1 | 19 | 21 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel-fatigue-DIM
Heel-fatigue
Heel-fatigue | [7 empty] | Diameter pendulum shaft - mm | 12.5 | ±1 | Diameter pendulum shaft - mm | Đường kính trục con lắc | mm | 12.5 | 1 | 11.5 | 13.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel-fatigue-DIM
Heel-fatigue
Heel-fatigue | [7 empty] | Thickness of striker head - mm | 6 | ±0.5 | Thickness of striker head - mm | Độ dày đầu gõ | mm | 6 | 0.5 | 5.5 | 6.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel-fatigue-DIM
Heel-fatigue
Heel-fatigue | [7 empty] | With of striker head - mm | 20 | ±0.5 | With of striker head - mm | Độ rộng đầu gõ | mm | 20 | 0.5 | 19.5 | 20.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel-fatigue-DIM
Heel-fatigue
Heel-fatigue | [7 empty] | Lenth of striker head - mm | 35 | ±2 | Lenth of striker head - mm | Độ dài đầu gõ | mm | 35 | 2 | 33 | 37 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel-fatigue-DIM
Heel-fatigue
Heel-fatigue | [7 empty] | Radius of the rounded striker - mm | 3 | ±0.5 | Radius of the rounded striker - mm | Bán kính vật gõ | mm | 3 | 0.5 | 2.5 | 3.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel-fatigue-DIM
Heel-fatigue
Heel-fatigue |  |  |  | 9: Check weight - g; 12: N/A; 11: Check weight - g; 13: Kiểm tra khối lượng; 14: g; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Heel-fatigue-MASS
Heel-fatigue
Heel-fatigue |  |  |  | 9: Speed (*) - rpm; 10: 60; 12: N/A; 11: Speed (*) - rpm; 13: Tốc độ; 14: rpm; 15: 60; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Heel-fatigue-RPM
Heel-fatigue
Heel-fatigue |  |  |  | 9: Counter - Cycle; 10: 100; 12: N/A; 11: Counter - Cycle; 13: Bộ đếm; 14: Cycle; 15: 100; 19: TIM1; 21: 0.02 |  | 0.001 | 0.02 | Heel-fatigue-TIM1
Heel-fatigue |  |  |  | 9: Others; 11: *
Heel-fatigue | bar: 9; 11
Heel-fatigue | bar: 9; 11
Heel-fatigue | bar: 9; 11
Heel-fatigue | bar: 9; 11
Heel-fatigue | bar: 9; 11
Heel-impact |  |  | TAV-QT-DIM | Heel Impact tester | Heel Impact tester | Máy kiểm tra tác động gót chân |  | Diameter Pendulum - mm | 108 | ±1 | Diameter Pendulum - mm | Đường kính con lắc | mm | 108 | 1 | 107 | 109 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel-impact-DIM
Heel-impact
Heel-impact | [7 empty] | Thickness pendulum - mm | 49 | ±2 | Thickness pendulum - mm | Độ dày con lắc | mm | 49 | 2 | 47 | 51 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel-impact-DIM
Heel-impact
Heel-impact | [7 empty] | Diameter pendulum shaft - mm | 25 | ±0.5 | Diameter pendulum shaft - mm | Đường kính trục con lắc | mm | 25 | 0.5 | 24.5 | 25.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel-impact-DIM
Heel-impact
Heel-impact | [7 empty] | Thickness of striker head - mm | 6 | ±0.5 | Thickness of striker head - mm | Độ dày đầu gõ | mm | 6 | 0.5 | 5.5 | 6.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel-impact-DIM
Heel-impact
Heel-impact | [7 empty] | Width of striker head - mm | 25 | ±0.5 | Width of striker head - mm | Độ rộng đầu gõ | mm | 25 | 0.5 | 24.5 | 25.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel-impact-DIM
Heel-impact
Heel-impact | [7 empty] | Lenth of striker head - mm | 35 | ±2 | Lenth of striker head - mm | Độ dài đầu gõ | mm | 35 | 2 | 33 | 37 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel-impact-DIM
Heel-impact | [22 empty] | #N/A | #N/A | Heel-impact-
Heel-impact | [7 empty] | Radius of the rounded striker - mm | 3 | ±0.5 | Radius of the rounded striker - mm | Bán kính vật gõ | mm | 3 | 0.5 | 2.5 | 3.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Heel-impact-DIM
Heel-impact | [22 empty] | #N/A | #N/A | Heel-impact-
Heel-impact |  |  |  | 9: Check weight - g; 11: Check weight - g; 13: Kiểm tra khối lượng; 14: g; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Heel-impact-MASS
Heel-impact |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Heel-impact-
Heel-impact |  |  |  | 9: 1…; 11: *; 21: #N/A |  | #N/A | #N/A | Heel-impact-
Heel-impact |  |  |  | 9: 2…; 11: *; 21: #N/A |  | #N/A | #N/A | Heel-impact-
Heel-impact |  |  |  | 9: 3…; 11: *; 21: #N/A |  | #N/A | #N/A | Heel-impact-
Heel-impact |  |  |  | 9: 4…; 11: *; 21: #N/A |  | #N/A | #N/A | Heel-impact-
Heel-impact |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Heel-impact-
Circular |  | Circular Locus Tester |  | Circular Locus Tester | Circular Locus Tester | Thiết bị kiểm tra định vị hình tròn | Diameter Area Calculation | Dimension of specimen - mm (Area - 100 cm²) | 112.8 | ±0.5 | Dimension of specimen - mm (Area - 100 cm²) | Diện tích | mm | 112.8 | 0.5 | 112.3 | 113.3 | DIM |  | #N/A |  |  | #N/A | #N/A | Circular -DIM
Circular | [7 empty] | Diameter | 40 | ±1 | Diameter | Đương kính | mm | 40 | 1 | 39 | 41 | DIM |  | #N/A |  |  | #N/A | #N/A | Circular -DIM
Circular | [7 empty] | Specimen motion speed (*) - rpm | 60 | ±1 | Specimen motion speed (*) - rpm | Tốc độ | rpm | 60 | 1 | 59 | 61 | RPM |  | #N/A |  |  | #N/A | #N/A | Circular -RPM
Circular | [7 empty] | Specimen pressure - cN | 490 | N/A | Specimen pressure - cN | Áp suất | cN | 490 |  | 490 | 490 | PRESS |  | #N/A |  |  | #N/A | #N/A | Circular -PRESS
Circular |  |  |  | 10: 590; 12: N/A; 14: cN; 15: 590; 17: 590; 18: 590; 19: PRESS; 21: #N/A |  | #N/A | #N/A | Circular -PRESS
Circular |  |  |  | 10: 780; 12: N/A; 14: cN; 15: 780; 17: 780; 18: 780; 19: PRESS; 21: #N/A |  | #N/A | #N/A | Circular -PRESS
Circular | [7 empty] | Counter - cycle/min | 60 | ±1 | Counter - cycle/min | Bộ đếm | Cycle | 60 |  | 60 | 60 | TIM1 |  | #N/A |  |  | #N/A | #N/A | Circular -TIM1
Circular |  |  |  | 9: Height of Brush; 12: N/A; 11: Height of Brush; 13: Chiều cao; 14: mm; 17: 0; 18: 0; 19: DIM; 21: #N/A |  | #N/A | #N/A | Circular -DIM
Rub-circular | SATRA TM8 |  | TAV-QT-DIM | Circular Rubbing Tester | Circular Rubbing Tester | Máy thử cọ xát tròn |  | External diameter - mm | 25 | ±1 | External diameter - mm | Đường kính ngoài |  | 25 | 1 | 24 | 26 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Rub-circular-DIM
Rub-circular
Rub-circular | [7 empty] | Inner diameter - mm | 3 | ±0.5 | Inner diameter - mm | Đường kính trong |  | 3 | 0.5 | 2.5 | 3.5 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Rub-circular-DIM
Rub-circular
Rub-circular | [7 empty] | Load (*) - N | 7.1 | ±0.2 | Load (*) - N | Tải trọng | N | 7.1 | 0.2 | 6.9 | 7.3 | FORCES |  | 0.14 |  |  | 0.0001 | 0.14 | Rub-circular-FORCES
Rub-circular |  |  |  | 10: 24.5; 12: ±0.5; 14: N; 15: 24.5; 16: 0.5; 17: 24; 18: 25; 19: FORCES; 21: 0.14 |  | 0.0001 | 0.14 | Rub-circular-FORCES
Rub-circular | [7 empty] | Revolution of rubbing axis - rpm | 150 | ±5 | Revolution of rubbing axis - rpm | Vòng quay trục cọ xát | rpm | 150 | 5 | 145 | 155 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Rub-circular-RPM
Rub-circular | [22 empty] | #N/A | #N/A | Rub-circular-
Rub-circular |  |  |  | 9: Counter - Cycle; 10: 100; 12: N/A; 11: Counter - Cycle; 13: Bộ đếm; 14: Cycle; 15: 100; 19: TIM1; 21: 0.02 |  | 0.001 | 0.02 | Rub-circular-TIM1
Rub-circular |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Rub-circular-
Rub-circular |  |  |  | 9: 1…; 11: *; 21: #N/A |  | #N/A | #N/A | Rub-circular-
Rub-circular |  |  |  | 9: 2…; 11: *; 21: #N/A |  | #N/A | #N/A | Rub-circular-
Rub-circular |  |  |  | 9: 3…; 11: *; 21: #N/A |  | #N/A | #N/A | Rub-circular-
Rub-circular |  |  |  | 9: 4…; 11: *; 21: #N/A |  | #N/A | #N/A | Rub-circular-
Rub-circular |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Rub-circular-
Padder |  |  | TAV-QT-DIM TAV-QTHKL TAV-QTHTD | 5: Motorized lab padder; 6: Motorized lab padder; 9: Rollers Dia. - mm; 10: 54; 12: N/A; 11: Rollers Dia. - mm; 13: Đường kính cuộn; 14: mm; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Padder-DIM
Padder
Padder |  |  |  | 9: Speed - mm/s; 10: 25; 12: N/A; 11: Speed - mm/s; 13: Tốc độ; 14: mm/s; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Padder-RPM
Padder
Padder |  |  |  | 9: Speed (*) - rpm; 10: 9; 12: N/A; 11: Speed (*) - rpm; 13: Tốc độ; 14: rpm; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Padder-RPM
Padder
Padder |  |  |  | 9: Weight; 11: *; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Padder-MASS
Padder |  |  |  | 11: *; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Padder-MASS
Padder |  |  |  | 11: *; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Padder-MASS
Padder |  |  |  | 11: *; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Padder-MASS
Padder |  |  |  | 11: *; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Padder-MASS
Padder |  |  |  | 11: *; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Padder-MASS
Padder |  |  |  | 11: *; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Padder-MASS
Padder |  |  |  | 11: *; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Padder-MASS
Padder |  |  |  | 11: *; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Padder-MASS
Padder |  |  |  | 11: *; 19: MASS; 21: 0.1 |  | 0.01 | 0.1 | Padder-MASS
Padding |  |  | TAV-QT-DIM TAV-QTHKL TAV-QTHTD | Laboratory Padding Mangle | Laboratory Padding Mangle | Máy ép ngấm hóa chất lên vải - Kiểu trục đứng |  | Rollers Dia. - mm | N/A | N/A | Rollers Dia. - mm | Đường kính trục lăn | mm |  |  |  |  | DIM |  | #REF! |  |  | #REF! | #REF! | Padding-DIM
Padding |  |  |  | 9: Length of rollers - mm; 10: N/A; 12: N/A; 11: Length of rollers - mm; 13: Chiều dài trục lăn; 14: mm; 19: DIM; 21: #REF! |  | #REF! | #REF! | Padding-DIM
Padding |  |  |  | 9: Speed (*) - rpm; 10: 20; 12: N/A; 11: Speed (*) - rpm; 13: Tốc độ; 14: rpm; 19: RPM; 21: #REF! |  | #REF! | #REF! | Padding-RPM
Padding |  |  |  | 10: 25; 12: N/A; 19: RPM; 21: #REF! |  | #REF! | #REF! | Padding-RPM
Padding |  |  |  | 10: 30; 12: N/A; 19: RPM; 21: #REF! |  | #REF! | #REF! | Padding-RPM
Padding |  |  |  | 10: 35; 12: N/A; 19: RPM; 21: #REF! |  | #REF! | #REF! | Padding-RPM
Padding |  |  |  | 10: 40; 12: N/A; 19: RPM; 21: #REF! |  | #REF! | #REF! | Padding-RPM
Padding |  |  |  | 10: 45; 12: N/A; 19: RPM; 21: #REF! |  | #REF! | #REF! | Padding-RPM
Padding |  |  |  | 10: 50; 12: N/A; 19: RPM; 21: #REF! |  | #REF! | #REF! | Padding-RPM
Padding |  |  |  | 9: Pessure air supply - mpa; 10: 0.2; 12: N/A; 11: Pessure air supply - mpa; 13: Áp suất đầu vào; 14: mpa; 19: PRESS; 21: #REF! |  | #REF! | #REF! | Padding-PRESS
Padding |  |  |  | 10: 0.4; 12: N/A; 19: PRESS; 21: #REF! |  | #REF! | #REF! | Padding-PRESS
Padding |  |  |  | 10: 0.6; 12: N/A; 19: PRESS; 21: #REF! |  | #REF! | #REF! | Padding-PRESS
Padding |  |  |  | 10: 0.8; 12: N/A; 19: PRESS; 21: #REF! |  | #REF! | #REF! | Padding-PRESS
Padding |  |  |  | 9: Pessure Left - mpa; 10: 0.2; 12: N/A; 11: Pessure Left - mpa; 13: Áp suất trái; 14: mpa; 19: PRESS; 21: #REF! |  | #REF! | #REF! | Padding-PRESS
Padding |  |  |  | 10: 0.4; 12: N/A; 19: PRESS; 21: #REF! |  | #REF! | #REF! | Padding-PRESS
Padding |  |  |  | 10: 0.6; 12: N/A; 19: PRESS; 21: #REF! |  | #REF! | #REF! | Padding-PRESS
Padding |  |  |  | 10: 0.8; 12: N/A; 19: PRESS; 21: #REF! |  | #REF! | #REF! | Padding-PRESS
Padding |  |  |  | 9: Pessure Right - mpa; 10: 0.2; 12: N/A; 11: Pessure Right - mpa; 13: Áp suất phải; 14: mpa; 19: PRESS; 21: #REF! |  | #REF! | #REF! | Padding-PRESS
Padding |  |  |  | 10: 0.4; 12: N/A; 19: PRESS; 21: #REF! |  | #REF! | #REF! | Padding-PRESS
Padding |  |  |  | 10: 0.6; 12: N/A; 19: PRESS; 21: #REF! |  | #REF! | #REF! | Padding-PRESS
Padding |  |  |  | 10: 0.8; 12: N/A; 19: PRESS; 21: #REF! |  | #REF! | #REF! | Padding-PRESS
Padding |  |  |  | 9: Others….; 11: Others….; 21: #REF! |  | #REF! | #REF! | Padding-
Padding |  |  |  | 9: …; 21: #REF! |  | #REF! | #REF! | Padding-
Padding |  |  |  | 9: …...; 21: #REF! |  | #REF! | #REF! | Padding-
Padding |  |  |  | 9: ….; 21: #REF! |  | #REF! | #REF! | Padding-
Padding |  |  |  | 21: #REF! |  | #REF! | #REF! | Padding-
Padding |  |  |  | 21: #REF! |  | #REF! | #REF! | Padding-
Padding |  |  |  | 21: #REF! |  | #REF! | #REF! | Padding-
Mace |  | Mace Snagging | TAV-QTHTD TAV-QT-PILL2 | Mace Snagging Tester | Mace Snagging Tester | Thiết bị thử độ xước móc | Rotational Speed Counter Pin height | Rotational Speed (*) - rpm | Station 1 - 60 | ±2 | Rotational Speed (*) - rpm | Tốc độ xoay | rpm | 60 | 2 | 58 | 62 | RPM |  | 0.1 |  |  | 0.1 | 0.1 | Mace-RPM
Mace |  |  |  | 10: Station 2 - 60; 12: ±2; 14: rpm; 15: 60; 16: 2; 17: 58; 18: 62; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Mace-RPM
Mace |  |  |  | 10: Station 3 - 60; 12: ±2; 14: rpm; 15: 60; 16: 2; 17: 58; 18: 62; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Mace-RPM
Mace |  |  |  | 10: Station 4 - 60; 12: ±2; 14: rpm; 15: 60; 16: 2; 17: 58; 18: 62; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Mace-RPM
Mace | [7 empty] | Counter - Cycle | 60 | ±1 | Counter - Cycle | Bộ đếm | Cycle | 60 | 1 | 59 | 61 | TIM1 |  | 0.02 |  |  | 0.001 | 0.02 | Mace-TIM1
Mace | [7 empty] | Pin height - mm | Station 1 - 9.52 | ±0.5 | Pin height - mm | Chiều cao đinh | mm | 9.52 | 0.5 | 9.02 | 10.02 | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Mace-DIM
Mace |  |  |  | 10: Station 2 - 9.52; 12: ±0.5; 14: mm; 15: 9.52; 16: 0.5; 17: 9.02; 18: 10.02; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Mace-DIM
Mace |  |  |  | 10: Station 3 - 9.52; 12: ±0.5; 14: mm; 15: 9.52; 16: 0.5; 17: 9.02; 18: 10.02; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Mace-DIM
Mace |  |  |  | 10: Station 4 - 9.52; 12: ±0.5; 14: mm; 15: 9.52; 16: 0.5; 17: 9.02; 18: 10.02; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Mace-DIM
Button |  |  | TAV-QTHKL TAV-QT-DIM | Button Impact | Button Impact |  |  | Weight - g | 840 | N/A | Weight - g | Tải trọng | g | 840 |  |  |  | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Button-MASS
Button
Button |  |  |  | 9: Dimension - mm; 10: 32; 12: N/A; 11: Dimension - mm; 13: Độ dài; 14: mm; 15: 32; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Button-DIM
Button |  |  |  | 10: 44; 12: N/A; 14: mm; 15: 44; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Button-DIM
Button |  |  |  | 10: 67; 12: N/A; 14: mm; 15: 67; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Button-DIM
Button |  |  |  | 10: 83; 12: N/A; 14: mm; 15: 83; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Button-DIM
Button |  |  |  | 10: 95; 12: N/A; 14: mm; 15: 95; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Button-DIM
Button |  |  |  | 10: 117; 12: N/A; 14: mm; 15: 117; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Button-DIM
Button |  |  |  | 10: 200; 12: N/A; 14: mm; 15: 200; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Button-DIM
Button |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A | Button-
Button |  |  |  | 9: 1…; 11: *; 21: #N/A |  | #N/A | #N/A | Button-
Button |  |  |  | 9: 2…; 11: *; 21: #N/A |  | #N/A | #N/A | Button-
Button |  |  |  | 9: 3…; 11: *; 21: #N/A |  | #N/A | #N/A | Button-
Button |  |  |  | 9: 4…; 11: *; 21: #N/A |  | #N/A | #N/A | Button-
Button |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A | Button-
Angle |  | 45˚ Fixed angle table | TAV-QTHTD TAV-QT-PILL2 | 45˚ Fixed angle table | 45˚ Fixed angle table | Tấm bảng nghiêng | Angle | Angle - ⁰ | 45 | N/A | Angle - ⁰ | Góc | ⁰ | N/A | N/A | #VALUE! | #VALUE! | ANG |  | 0.03 |  |  | 0.05 | 0.03 | Angle-ANG
Angle |  |  |  | 9: Angle - ⁰; 10: 90; 12: N/A; 11: Angle - ⁰; 14: ⁰; 15: N/A; 16: N/A; 19: ANG; 21: 0.03 |  | 0.05 | 0.03 | Angle-ANG
Down-proof | GB/T 12705.2, FTMS 191A |  |  | Fabric Down-proof Property Tester | Fabric Down-proof Property Tester |  |  | Box dimension - mm | 450 | N/A | Box dimension - mm | Kích thước hộp | mm | 450 |  |  |  | DIM |  | 0.02 |  |  | 0.01 | 0.02 | Down-proof-DIM
Down-proof |  |  |  | 10: 450; 12: N/A; 15: 450; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Down-proof-DIM
Down-proof |  |  |  | 10: 450; 12: N/A; 15: 450; 19: DIM; 21: 0.02 |  | 0.01 | 0.02 | Down-proof-DIM
Down-proof |  |  |  | 9: Speed (*) - rpm; 10: 45; 12: ±2; 13: Tốc độ; 15: 45; 16: 2; 17: 43; 18: 47; 19: RPM; 21: 0.1 |  | 0.1 | 0.1 | Down-proof-RPM
Down-proof | [23 empty] | #N/A
Down-proof |  |  |  | 9: Number of revolutions - cycle; 10: 100; 12: N/A; 11: Number of revolutions - cycle; 13: Bộ đếm; 14: Cycle; 15: 100; 19: TIM1; 21: 0.02 |  | 0.001 | 0.02 | Down-proof-TIM1
Down-proof | [22 empty] | #N/A | #N/A
Down-proof | [7 empty] | Weight - g | 140 | ±5 |  | Tải trọng | g | 140 | 5 | 135 | 145 | MASS |  | 0.1 |  |  | 0.01 | 0.1 | Down-proof-MASS
Down-proof |  |  |  | 9: Others; 11: *; 21: #N/A |  | #N/A | #N/A
Down-proof |  |  |  | 9: 1…; 11: *; 21: #N/A |  | #N/A | #N/A
Down-proof |  |  |  | 9: 2…; 11: *; 21: #N/A |  | #N/A | #N/A
Down-proof |  |  |  | 9: 3…; 11: *; 21: #N/A |  | #N/A | #N/A
Down-proof |  |  |  | 9: 4…; 11: *; 21: #N/A |  | #N/A | #N/A
Down-proof |  |  |  | 9: 5…; 11: *; 21: #N/A |  | #N/A | #N/A
Note on G2754: TAV-Trong: J-50C HANS SCHMIDT